TEMPERATURE SENSOR CT2B-121

CHARACTERISTICS

The temperature sensor type CT2B-121 is designed for temperature measurements in liquid and semi-solid substances and also for air and gases measurements. It may be used during the laboratory, field and industrial measurements. The measuring element is a platinum resistor Pt-1000 Class B. The sensor has acid-resistant stainless steel body which in most cases protects the measuring element against mechanical damage and against damage caused by some aggressive solutions .As standard the sensor is equipped with 1m silicone cable and Chinch connector. Silicone cable is resistant to temperatures up to 250 °C, it may be also in steel braid to protect it against mechanical damage. The end of the sensor may be ended flat or sharp, depending on customer's needs.

TECHNICAL DATA

Measuring Range			-70  to 400 °C

Measuring element			Pt-1000 Class B

Body diameter				3,0mm ± 0,5 mm

Body length (without cable socket) 	115 ± 5 mm

Minimal depth of immersing		10 mm

Reaction time t99 20 to 100°C	12 seconds

Body material				acid-resistant stainless steel

Steel type 0H18N9

Cable length 				1 m

Cable material 				silicone (resistan to 250 °C)

Handle material				Tarnamide (resistant to 120 °C)

Connector				Chinch

(  Sp. j.

  - Zabrze  . Witosa 10  POLAND

  - tel.  +48 32 / 2738106   fax  +48 32 / 2738114

www.elmetron.com.pl     e-mail: info@elmetron.com.pl