LABORATORY

pH / ION METER

CPI-505

USER’S MANUAL

USER’S MANUAL

LABORATORY

pH/ION METER

CPI-505

Before use please read the instruction carefully!

TABLE OF CONTENT

I.	INTRODUCTION	1

1.		EXPLOITATION NOTICES	3

2.		THE CHARACTERISTICS OF THE METER	4

3.		WHAT IS THE METER DESIGNED FOR	5

4.		OUTSIDE VIEW	6

5.		SWITCHING THE METER ON AND OFF	9

6.		PREPARATION TO WORK	10

6.1.	Choosing the kind of temperature compensation	10

6.2.	Changing the resolution of the measurements	11

6.3.	Changing the electrode number	12

6.4.	Readout of the date of last calibration	13

6.5.	Entering the calibration validity date	14

II.	pH MEASUREMENT	15

7.		PREPARATION OF THE pH ELECTRODE	17

8.		CALIBRATION	18

8.1.	Calibration modes	19

8.2.	Calibration with use of NIST sample solutions	20

8.3.	Calibration with introducing the pH value of the sample solutions	21

8.4.	Entering the calibration parameters	22

8.5.	pH electrode calibration	24

9.		CHECKING THE ELECTRODE CONDITION	27

10.	pH MEASUREMENT	28

10.1.	Measurement with automatic temperature compensation	28

10.2.	Measurement with manual temperature compensation	29

11.	NOTICES ABOUT THE TEMPERATURE COMPENSATION AND

INTERPRETATION OF THE pH MEASUREMENTS RESULTS	30

III.	ION SELECTIVE MEASUREMENTS	33

12.	NOTICES ABOUT ION SELECTIVE MEASUREMENTS	34

13.	Entering the parameters of the ion meter	37

13.1.	Choosing the unit	37

14.	ION SELECTIVE ELECTRODES CALIBRATION	38

14.1.	Entering the calibration parameters	38

15.	ION SELECTIVE MEASUREMENT	44

IV.	ORP (mV) AND TEMPERATURE  MEASUREMENT	45

16.	ORP (mV) MEASUREMENT	47

17.	TEMPERATURE MEASUREMENT	47

V.	OTHER	49

18.	CLOCK WITH DATE	51

18.1.	Clock	51

18.2.	Date	51

18.3.	Setting the time and date	52

18.4.	The LCD backlight mode	52

18.5.	Brightness control	53

18.6.	Readout of the software version number	54

19.	STORING AND READOUT OF THE RESULTS	55

19.1.	Parameters of storing and readout from the memory	55

19.2.	Storing the single measurements to the memory	57

19.3.	Storing the measuring series	57

19.4.	Reviewing the results	58

19.5.	Deleting the stored results	58

20.	CO-OPERATION WITH A PC	59

21.	TECHNICAL DATA	60

22.	EQUIPMENT	62

##### INTRODUCTION

# EXPLOITATION NOTICES

Dear User!

We present you a device distinguished by accuracy according to the technical data and by a high stability of the displayed results. We believe that the measurements will not cause you any trouble and that the meter will operate without any inconvenience. Wide range of additional functions requires careful reading of the manual in other case some of the features may stay unused or using the meter may cause you a troubles.

The employing of good-quality electrodes and their replacing after a suitable time ensures obtaining of high measuring parameters. We want to call your attention to the fact, that this equipment has a much shorter working life than the meter. Typical symptoms of an improper operation of the electrode are deterioration of final result stability, it’s flowing as well as a higher measuring error. Part of the users has problems, arising from employing electrodes not being preconditioned before the measurement or making the measurements without removing the shielding ring from the liquid junction or taking measurements with a plugged junction. To avoid this situation it is necessary to choose the proper kind of electrode for solutions which are going to be measured ex. special electrodes for the sewage, liquids with deposits, meat `cheese etc. Therefore, if you observe improper operation of the device, please take control measurements with another electrode or check the used electrode with another pH-meter. Generally the deterioration of the meter work is caused by the electrode and not by the meter.

During ion selective measurements a very often case is worsening of the results stability, this is connected only with the quality of the ion selective electrodes. It is important to keep in cleanness the measuring diaphragm of the ion selective electrode which may be contaminated during measurements in solutions containing compounds influencing the activity of electrode’s membrane. The methods of cleaning the electrodes are given by their manufacturer.

The essential feature of our  products is their low failure frequency. However if your meter will fail, our firm immediately performs its warranty repair.

We wish you a pleasant and trouble-free work with our meter.

# THE CHARACTERISTICS OF THE METER

The pH / Ion meter CPI-505 belongs to the newest generation of measuring devices which offer wide range of additional functions. The meter ensures high accuracy and repeatability of the readings. The newest generation electronic elements used in the meter made it’s memory independent to power supply. The internal clock is  powered from battery which is enough for many years. The meter is equipped with custom, backlight LCD display, which enables simultaneous observing of the measured function, temperature value and additional symbols which make working easier.

Main features of the CPI-505 are:

- high accuracy and stability of the reading;

- measurement of pH, ions, mV and temperature;

- automatic and manual temperature compensation;

- independent storing of calibration characteristics of three pH electrodes and ion selective electrodes;

- electrode calibration in 1 to 5 points;

- automatic recognition of pH buffers and standards;

- imposed values of standard solutions with possibility of changing their value;

- option of automatic introduction of temperature correction on the value of pH buffer solutions (NIST norm);

- information about the condition of the electrode;

- in ion meter measurements in pX, g/l, M/l or ppm with automatic conversion between units;

- in the ion meter possibility of introducing a freely chosen values of the sample solutions;

- storing the date and parameters of last calibration of three electrodes;

- storing the measurements results with time, date, individually or taken as series of measurements with set time interval;

- USB port;

- real clock with date.

# WHAT IS THE METER DESIGNED FOR

The pH / ion meter CPI-505 is  precise and easy to use meter designed for hydrogen ion concentration measurements in pH units, ion selective measurements in pX, g/l, M/l or ppm and Oxidation Reduction Potential (mV). There is possibility of measurements of mono-valent and bivalent ions, positive and negative. Measurement of the chosen ion is connected with using a combination ion selective electrode (ISE) or measuring and reference electrodes.

The meters may be also used for accurate temperature measurement of solutions and air in °C.

The meters are being used in food, chemical, pharmaceutical, energetic industries, in water treatment stations, laboratories, agriculture, universities, scientific laboratories etc.

CPI-505 are prepared to work with all types of combination pH electrodes and ion selective electrodes equipped with BNC-50 connector. There is also possibility of using a separate measuring and reference electrode. The meters co-operate with Pt-1000 temperature probe with a Chinch connector.

The meter may collect measurements taken as single or series of measurements with set time interval.

The built in USB port enables sending the data to the PC. There is possibility of sending current results of measurements or collected in the meter’s memory together with date and time of their collecting.

In case of necessity to collect series longer than the memory capacity results it is possible to use special PC software offered by our company.

# OUTSIDE VIEW

In the upper part of the meter there is a LCD display (Pic. 1), on which depending on the chosen function following symbols are displayed:

- result of the pH measurement in pH units;

- result of ion measurements in pX, g/l, M/l or ppm;

- result of the mV measurement in mV;

- time and date.

Choosing the function is done by specific buttons and is signalised by lighting LED diode placed on the button.

Simultaneously with the result a measured temperature value is displayed in oC. Symbols of the units are displayed by the results.

Pic. 1.

Beside the temperature value a  symbol for automatic temperature compensation is displayed or  for manual. CAL symbol on the left side of the display informs that the meter is in calibration mode. On the left side of the display a number of chosen electrode is showed (E1; E2; E3) which inform about the number of characteristic which will be taken into consideration during all calculations. Deleted characteristic, validity time of electrode calibration exceeded and (only in pH function) lost pH electrode efficiency from last calibration is signalised  by flashing symbol of electrode. During calibration between the upper and lower row of numbers number of the calibration point is displayed ( P1, P2, P3, P4, P5).

In the MODE all parameters introduced by the user are displayed.

The keyboard  (Pic. 2) placed under the display is used for switching the meter on and off, choosing the measuring function, calibration, entering the parameters, printing and storing the results in the memory.

-	switches the meter on and off;

-	chooses the pH measuring function;

-	chooses the mV measuring function;

-	chooses the Ion measuring function;

-	displays the time and date;

-	longer pressing of this button enters the calibration mode (CAL symbol displayed). Short pressing in this mode confirms the calibration result;

-	short pressing causes storing single results or measuring series, longer pressing enters the stored results readout mode;

-	chooses the entered parameter;

, 	-	buttons for entering the parameters.

On the back wall of the meter there are inputs with below given symbols:

pH/mV	-	BNC-50 connector for connecting combined or measuring pH electrode and ORP or ion selective electrodes;

Gnd	-	connector for reference electrode;

temp	-	Chinch connector for temperature probe;

USB	-	USB port for connecting the meter with the PC;

POWER	-	connector for power adapter.

Pic. 2.

# SWITCHING THE METER ON AND OFF

The meter is switched on by pressing the  button. The meter tests the memory and display on which all symbols are displayed (Pic. 3).

Pic. 3.

If the test was successful, after about 1.5 s the meter switches it self automatically to the measuring mode, in which it was switched off. If an  sign will be displayed it means that the meter has lost the factory settings and requires the service repair. If after the 1,5 s all symbols will be continuously displayed it informs that the calibration parameters of electrodes or cells  were lost.

After pressing the  button the meter will take standard characteristics:

- shift = 0 pH, characteristic slope = 100% for pH electrode;

And will enter the measuring mode. It will be necessary to calibrate the conductivity cell, pH electrodes and oxygen sensor.

The meter is switched off by pressing and longer holding of the  button.

# PREPARATION TO WORK

Before starting the work one should:

- join the power adapter plug to the Power input;

- to pH/mV input (BNC-50) join the ready to work combination pH electrode, ion selective electrode or redox electrode;

- in case of using the reference electrode it should be connected to the Gnd input;

- in case of using the temperature probe it should be connected with the chinch temperature input - temp;

- in case of work with the PC join the cable with the USB port;

- switch the meter on by pressing the  button.

## Choosing the kind of temperature compensation

The meter switches it self to the automatic temperature compensation mode automatically after joining the temperature probe, after disconnecting it the meter enters the manual temperature compensation mode. In ATC mode near the displayed temperature a  symbol appears. Manual temperature compensation is indicated by  symbol near the value entered by the user, it’s value may be changed using  or  keys.

## Changing the resolution of the measurements

The measurement results may be displayed with chosen resolution. To change it one should:

-	in the measuring mode press the   button, a  (resolution) sign will be displayed. (Pic. 4);

-	using keys  or  one may choose:

- (low) resolution of the measurement;

- (high) resolution of the measurement.

Pic. 4.

For the pH measurement:

- resolution of the measurement 0.01 pH;

- resolution of the measurement 0.001 pH.

In the ion meter function:

for measurement in pX::

- resolution of the measurement 0.01 pX;

- resolution of the measurement 0.001 pX.

for measurement in M/l, g/l and ppm:

- resolution of the measurement 0.1%;

- resolution of the measurement 0.01 %.

There is no possibility of change the resolution in mV measuring mode.

Return to the measuring mode by pressing the chosen function button.

## Changing the electrode number

If in the meter’s memory more than one electrode characteristic was stored, it is possible to replace the electrodes without the need of calibration. This option is quite useful in case of using different kinds of electrodes ex. for sewage, and clear water etc. It is necessary to connect the earlier calibrated electrode, marked with the number which responds to the number which was stored in the memory and choose this number.

To do so in the measuring mode:

-	press the   button till the moment of displaying in the upper row of the display the ,  or  symbol, than using the  ,  buttons choose the electrode number. Under this number the calibration results will be stored (Pic. 5). Above the electrode number following signs will be displayed:

-	under this number there is no characteristic stored and producers values are provided. In the measuring mode the electrode number will be flashing.

-	under this number there are values of last calibration stored.

- the last calibration showed that the electrode is loosing it’s efficiency and in short time it’s calibration won’t be possible. (only for electrode calibrated in pH function). In the measuring mode the electrode number will be flashing.

Additionally under the electrode number the points in which the electrode was calibrated will be showed.

Pic. 5.

In the ion meter function during changing the electrode number it is possible to check the ion number for which the electrode was calibrated. It is done by short pressing the . button, in the lower row a  symbol will be

displayed and in the upper row it’s number. After pressing the  button the meter returns to the redout of the electrode number mode. Pressing of any function button returns the meter to measuring mode.

## Readout of the date of last calibration

The meter remembers dates of calibration of all electrodes and sensors in every measuring function. Before starting the work one may check the date of last calibration.

To do so in the measuring mode:

- press the   button till the moment of displaying in the upper row of the display the electrode number (,  or  symbol);

- using the  ,  buttons choose the electrode number which we calibration date we want to check and shortly press the  button. The date of last calibration will be displayed in the form: month – day and below the year (Pic. 6). On the left side is the number of the electrode displayed.

Pic. 6.

The meter stores the date in the moment of calibration. If after the calibration the date will be changed the meter will signalise with error the exceeding of the calibration validity, so it is necessary to remember about setting the correct date before the calibration.

Return to the readout of the electrode number after pressing the  button, or to measuring mode by pressing any of the function buttons.

## Entering the calibration validity date

In each measuring function the meter stores the validity date of the calibration for each of the electrodes or sensors. Exceeding the set time the meter will signalise with blinking number of the electrode (symbols ,  or ), informing that the electrode calibration characteristic stored under this number is no longer valid and the electrode should be calibrated.

To set the next calibration validity time one should in the measuring mode of the chosen function:

- Press the  button till the moment of displaying in the upper row of LCD a  (time out) symbol. Below the earlier introduced number of days till the next calibration will be displayed (Pic. 7);

- With  ,  buttons enter the requested number of days till the next calibration. If after choosing 1 day time one will press the  button a  symbol will be displayed instead of the digits – this will block the function of calibration time control.

Pic. 7.

Return to the measuring mode by pressing any of the function buttons.

##### pH MEASUREMENT

# PREPARATION OF THE pH ELECTRODE

The electrode should be prepared to work according to the producer instructions. If the instructions weren’t given please follow the steps:

-	new electrode should be put into distilled water or in saturated KCl solution for about 3 hours:

-	Before starting the measurements protecting rings (if used in this kind of electrode) should be removed. The ring placed on the junction – lower part of the electrode should be moved up along the electrode’s body and the upper, which protects the KCl refilling hole, down along the body.

Removing the lower ring is essential, in other case the electrode won’t measure.

Upper ring should be removed during measurements of high temperature solutions or to protect the junction during measurements in solutions with deposits or oils. Sometimes instead of ring a cork is used;

-	during measurements in laboratory it is advisable to use an electrode holder;

-	after every measurement the electrode should be washed in distilled water;

-	excess of liquid on the electrode should be removed by gentle touching the glass with a tissue paper;

-	after work the electrode should be stored in saturated KCl solution what prolongs it’s life. The protecting rings should be moved on the junction and upper hole;

-	in case of long breaks between the measurements the electrode should be stored, after drying, in the package;

-	after taking the electrode of the package the eventual deposit should be removed using water;

-	before using the electrode it should be placed in distilled water for about 1 hour;

-	if the construction of the electrode enables refilling the electrolyte, it should be controlled and refilled periodically by the upper hole in the electrode’s body (usually as the electrolyte a saturated KCl solution is used).

CAUTION:	storing of the electrode in distilled water shortens it’s life time and may cause shifting of the zero point.

# CALIBRATION

Before starting the measurement with new electrode, after long-lasting using, or before making measurements which require higher accuracy the electrode connected with the meter should be calibrated. The results of measurements done without calibration will have a great error. The calibration is done in the buffer or sample solutions with known value. It is based on comparing pH value of the buffer solutions with the value displayed by the meter and next automatic introduction of correction to the meter’s memory and taking it into consideration during measurements. The calibration should be periodically repeated because during work the parameters of the electrode are changing what influences the accuracy. The frequency of this procedure depends on the demanded accuracy, number of the measurements carried out, conditions in which the electrode was used, temperature and value of the measured solutions. When the highest accuracy is required it is recommended to use certified sample solutions. In technical usage buffer solutions are used with lower accuracy but mostly with total values ex. 2.00 pH; 4.00 pH etc. In this manual to make it easier a buffer solution name was used for both buffer and sample solutions. For accurate measurements it is necessary to use fresh good quality sample solutions.

The first action taken before the calibration is entering the values of used pH buffer solutions to the meters memory. This should be done before the first calibration and repeated in case of changing the used buffers to other with different values.

During the calibration after immersing the pH electrode and temperature probe the meter automatically detects the value of the used buffer, if it was earlier stored in the meter’s memory.

There is possibility of calibration in minimum 1 buffer solution and maximum in 5. The more points of calibration is used the higher is the accuracy in the whole measuring range.

Calibration in one buffer solution can’t ensure high accuracy. If only one buffer solution is used it’s value should be close to the forecasted value of measured solution. If the required accuracy isn’t very high and the measurements will be done in the whole range the one point calibration should be done with buffer solution close to 7.00 pH. Thanks this the so called zero electrode shift will be eliminated. In other points a standard characteristic will be taken from the meter’s memory. This characteristic corresponds with theoretical efficiency of the pH electrode.

If measurements are done both in acids and alkalis and the measurements are not done at the ends of measuring range it is enough to calibrate the electrodes in 3 buffer solutions with values in range given in table 2 – calibration points 2, 3 and 4. In case of accurate measurements in the whole range it is recommended to calibrate the electrode in all 5 points given in the table. In CPI-505 the characteristic of the electrodes is approximated linearly between the calibration points.

Starting the calibration under one of the electrode numbers does not remove the set buffer values, but irreparably removes the electrode characteristic stored under this number.

There is no possibility of calibrating only one point and leaving the rest of data from the last calibration.

The memory in CPI-505 enables independent storing of 3 characteristics of pH electrodes. This feature is especially useful when fast replacing of the electrodes is necessary or when replacing a broken one. Those electrodes should be earlier calibrated and brought into memory under following symbols (,  , ).

The order of using the buffer solutions is freely chosen.

## Calibration modes

The user may choose two independent ways of action:

- use the values of pH sample solutions entered to the memory by the manufacturer. Those values are in conformity with the NIST norm. This type of calibration switches on the automatic correction introduction connected with the change of sample solution along with temperature changes.(description in chapter 8.2).

- enter the value of pH buffer solutions depending on the actually used buffers, in the range given for each point of calibration; (chapter 8.3).

Choosing the first option makes the calibration much easier. The  user doesn’t need to warm up or cool down the sample solutions to the temperature given by the user. This option includes some simplifications and when very high accuracy is required it shouldn’t be used. The meter stores independently 5 values of the buffers for each of the calibration modes described above.

## Calibration with use of NIST sample solutions

In this mode 5 constant sample solutions values, according to NIST, are used. In the meters memory a table with dependence between the temperature and pH values for this 5 sample solutions is stored. This dependence is showed in the table 1. During the calibration the temperature of sample solution is measured and on it’s basis automatically the right pH value responding to this temperature is taken into consideration. Values between the points given in the table are approximated linearly by the meter.

Some sample solutions made according to the appropriate norms may differ on the third decimal place. The meter enables setting of the appropriate buffer solution values in range  ±0.010 pH from the values given in the table for 20(C. The way of entering the buffer solution value was described in chapter 8.4. In case of differences larger than ±0.010 pH the user may not use the mode with use of NIST sample solutions but the mode which enables setting the buffers value and before the calibration set, in the meter’s memory, the accurate value of the used sample solution according to pt. 8.3.

The range of the temperatures, taken into consideration during introducing the correction, is 0 to 60 oC and can’t be exceeded during calibration.

The way of action during calibration way choice and entering sample solutions values other than those in table 1 was described in chapters 8.4.

Table 1.

| Temp. 0C | Kind of buffer solution |
| --- | --- |
|  | 1 oxalate | 2 phthalate | 3 phosphate | 4 di-sodium tetraborate | 5 calcium hydroxide |
| 0 | 1.666 | 4.003 | 6.984 | 9.464 | 13.423 |
| 5 | 1.668 | 3.999 | 6.951 | 9.395 | 13.207 |
| 10 | 1.670 | 3.998 | 6.923 | 9.332 | 13.003 |
| 15 | 1.672 | 3.999 | 6.900 | 9.276 | 12.810 |
| 20 | 1.675 | 4.002 | 6.881 | 9.225 | 12.627 |
| 25 | 1.679 | 4.008 | 6.865 | 9.180 | 12.454 |
| 30 | 1.683 | 4.015 | 6.853 | 9.139 | 12.289 |
| 35 | 1.688 | 4.024 | 6.844 | 9.102 | 12.133 |
| 40 | 1.694 | 4.030 | 6.838 | 9.063 | 11.984 |
| 45 | 1.700 | 4.047 | 6.834 | 9.038 | 11.841 |
| 50 | 1.707 | 4.060 | 6.833 | 9.011 | 11.705 |
| 55 | 1.715 | 4.075 | 6.834 | 8.985 | 11.574 |
| 60 | 1.723 | 4.091 | 6.836 | 8.962 | 11.449 |

## Calibration with introducing the pH value of the sample solutions

Values of the pH buffer solutions set by the producer in the meter’s memory are given in the Table 2, this values may be changed in the range given in the table.  If the values of used buffer solutions are different than the given above the user should change the set values to the actually used.

The meter does not allow to enter the pH values different than those given in the table 2.

Table 2.

| Calibration point | Factory value | Range of changes |
| --- | --- | --- |
| 1 | 2,000 | 0,800 ( 2,100 |
| 2 | 4,000 | 3,900 ( 4,100 |
| 3 | 7,000 | 6,800 ( 7,100 |
| 4 | 9,000 | 8,900 ( 10,200 |
| 5 | 12,000 | 11,800( 14,000 |

In case of using buffer solutions with values different than the given by factory, it is necessary to remember about introducing this values to the meter’s memory. Way of doing it was described in the point 8.4.

The range of introducing the pH values of the buffer solutions in the individual points of calibration is quite wide what enables to use buffer solutions with values which differ from those set by the manufacturer. In each case the buffer will be automatically detected by the meter.

The manufacturers of sample solutions often give the values of the solutions in few temperatures. This data may be used and the meter may be calibrated in other temperature than 20 oC, by introducing to the meters memory value of the buffer which responds to this temperature, this will enable eliminating the necessity of warming up or cooling the solutions. It is possible to introduce to the meter’s memory values of the pH sample solutions with 2 or 3 decimal places, depending on the chosen resolution.

The value of sample solutions introduced to the meter’s memory by the user will be stored till the moment of changing it to other value.

## Entering the calibration parameters

It is necessary to:

a.	Choose or block the function of automatic correction of the change of pH value of buffer solution with the temperature. (described in chapter 8.1).

To do so one should press the  button than a  symbol (points of calibration) will be displayed, (Pic. 8) next using  or  buttons choose in the lower row on LCD:

– automatic change of the stored pH value of the buffer solution by it’s temperature change according to table 1.

– setting the value of the buffer solution in the range given in table 2.

Pic. 8.

b.	After choosing the  or the   symbol one should start checking and eventually changing the stored values of calibration points. To do so one should press the  button. In the lower row a  symbol will be displayed, point one of calibration, and in the upper row value of the buffer. (Pic. 9) If one uses buffer solution with value a bit different than this displayed one, using  or  buttons bring the displayed value to the value of used buffer solution.

Range of changes for each of calibration points in the  mode is ±0.010 from the values given in the table 1 for 20(C, and the range of changes in the   mode is given in the table 2.

Pic. 9.

To pass to the second point of calibration press the  button in the lower row a , symbol will be displayed (point 2) and in the upper row the stored in this point value of buffer solution. When checking or changing the value in the next steps please follow the instructions given above.

After the end of introducing the pH values of buffer solutions using the  button return to the choosing the calibration points mode  / , or by pressing the  button enter the pH measuring mode and calibration, chapter 8.5.

During next calibrations if the values of the previously used buffer solutions haven’t changed the actions described in this point may be omitted.

## pH electrode calibration

It is necessary to:

a.	choose the resolution with which the pH sample solution value will be introduced, point 6.2.

b.	choose the electrode number under which the results of calibration will be stored, point 6.3.

c.	prepare the electrode to work according to the producer’s instructions and mark it with number responding to number chosen in the meter (,  , ).

- connect with the meter the prepared combination pH electrode and temperature probe, use the right connectors pH/mV and temp;

- If the function of automatic change of the pH sample solution value is not used it is necessary to measure the temperature of sample  solutions and bring it to the temperature given by the manufacturer to receive the pH value identical like this entered to the meter’s memory. There is no limitation in order of using the samples.

Under each of the electrode numbers it is possible to calibrate different electrode in sample solutions with different values (after entering those values to the meters memory).

If the pH sample solutions with values according to NIST are used the values should not be changed. The meter takes under consideration only the values of sample solutions detected during calibration.

### Calibration with automatic temperature compensation

After preparing the meter to the calibration it is necessary in pH function to:

a	press and hold the  button till the appearance of the CAL (Pic. 10), symbol on the display; the old parameters of calibration are deleted;

b	put the electrode and the temperature probe to the buffer solution, a P symbol with number of calibration point will be displayed, wait till stabilisation of the result. The result may be different than pH value of the buffer solution (Pic. 10).

Pic. 10.

After stabilisation of the result press the .button.

In this moment the result will blink what will inform about the storing the calibration result, in this same time in the upper row the corrected value of measurement will be displayed – it will be equal the value of used buffer solution (Pic. 11).

If the function of automatic correction of the temperature influence on pH value of the buffer solution is used, this value will be considered in the displayed value of the buffer solution. If by accident someone will use buffer solution different than those set for the calibration the meter won’t be able to detect this value and  sign will be displayed. In this case it is necessary to check the value of the buffer solution or the electrode which may be broken or clogged.

Pic. 11.

- one may finish the calibration at this moment by pressing the  button or continue the calibration in other buffer solutions, washing the electrode and temperature probe before every immersing in the buffer and acting according to the point b.

After calibrating one electrode it is possible to calibrate two others, choosing the two remaining symbols according to 6.3.

If after choosing the electrode number and entering the calibration mode one will exit it without making the calibration the earlier stored characteristic will be deleted and a standard characteristic will be adapted

### Calibration with manual temperature compensation

To start the calibration with manual temperature compensation one has to disconnect the temperature probe. Disconnecting the temperature probe switches the meter to manual compensation. On the display the entered value of the temperature is displayed, not the measured one. The ,  buttons will be unblocked and using them one may enter the value of the buffer solution temperature. This value is displayed in the lower row of numbers on the display.

Next thing is to connect the pH electrode with the meter and act according to points a ( c. of the previous section. It is important to remember that in case of manual temperature compensation the temperature value entered by the user is taken under consideration.

Caution:	simultaneous pressing of both  and  buttons sets the temperature to 20 0C.

# CHECKING THE ELECTRODE CONDITION

If after the calibration the electrode symbol is blinking (E1, E2, E3) it informs that the electrode has lost its efficiency and in short time its calibration won’t be possible. After entering the mode of electrode number changing under the number of electrode a  symbol will be displayed (Pic. 12).

It is necessary to prepare a new electrode.

Pic. 12.

# pH MEASUREMENT

Before starting the measurement one has to prepare the meter for work (chapter 6) and pH electrode (chapter 7). Good condition of the electrode is the main condition of correct measurements. If the electrode was calibrated and is connected with the meter it is advisable to check whether the number of the electrode is this same as the number of the characteristic chosen from the memory. If not, the number should be changed according to section  6.3 and the resolution of measurement should be chosen according to the chapter 6.2.

## Measurement with automatic temperature compensation

During measurements with automatic temperature compensation, the meter co-operates with the temperature probe and measures the temperature of the solution simultaneously with the pH measurement and takes it into consideration during compensation.

In case of measurement with automatic temperature compensation one should:

- turn the meter on by pressing the  button;

- with the  button choose the pH measuring function;

- connect the temperature probe and the combination pH electrode to the right connectors on the meter pH/mV and temp, symbol  will be displayed;

- if the electrode wasn’t calibrated or was already used for some time it is advisable to calibrate it. (chapter 8);

- insert the electrode and the temperature probe to the measured solution. During measurements in vessels one should not touch the bottom and the walls with the electrode. Best way is to use an electrode stand;

- After the stabilisation read the result

Accurate laboratory measurements require using of electromagnetic stirrer.

NOTICE: exceeding of the measuring range is indicated by blinking of the digits on the display. Exceeding the automatic temperature compensation range is signalised by blinking digits and  symbol.

## Measurement with manual temperature compensation

Disconnecting the temperature probe from the meter switches the meter to the manual temperature compensation mode (symbol  is displayed).

Measurement with manual temperature compensation is similar to the measurement with ATC, the difference is that using ,  buttons one enters the temperature measured with other thermometer. This value is displayed under the pH value and is taken into consideration during the compensation.

The manual compensation may be used during stable conditions ex. during pH measurements in the laboratory, especially when a thermostat is used, or when the temperature probe was damaged.

During the measurement with manual temperature compensation one should:

- turn the meter on using the  button;

- using the  button choose the pH function;

- insert the pH electrode to the vessel with the measured solution, if the electrode wasn’t calibrated or already was used for a long time the calibration should be done (chapter 8). During measurements in a vessel one shouldn’t touch the bottom and the walls with the electrode. The best way is to use an electrode stand;

- using a laboratory thermometer measure the temperature;

- with    buttons enter the value of measured temperature;

- wait till the stabilisation of the value and read the result.

Caution: simultaneous pressing of  and   buttons sets the temperature value to 20 0C.

# NOTICES ABOUT THE TEMPERATURE COMPENSATION AND INTERPRETATION OF THE pH MEASUREMENTS RESULTS

The CPI-505 pH / ion meter has a manual and automatic temperature compensation, what enables eliminating errors resulting from changes of the electrode characteristics under the influence of temperature changes. To explain the role of the temperature compensation it is important to remind that pH meter is an mV meter which displays voltage counted to pH unit. In constant temperature for one pH unit a constant mV value is changed. In temperature of 20 0C it is 58,168 mV. The value of mV on one pH unit changes together with the temperature, what is taken into consideration in the formula for “k coefficient” of the pH electrode.

k=0.198422 T

Taking this change into consideration during measurement is called temperature compensation, which is connected with change of the electrode efficiency and not with the pH value change of the measured solution caused by temperature change.

Changes of the pH values of some solutions with the temperature are very small and others big ex. pure water.

When comparing solutions, which change their values together with the temperature change, it should be done in this same temperature.

Sometimes it happens that results of measurement in this same solution in stable temperature are different. When analysing such a situation it is important to take following factors into consideration:

- Differences may occur because of  not highest quality electrode;

- The result was treated as stabilised to fast (medium class electrode needs about 40 seconds to full stabilisation);

- The measured solution may not be homogeneous and lack of magnetic stirrer doesn’t allow for obtaining similar results;

- During measurements in sewage some chemical reactions, which change the result, may occur;

- The source of additional errors, which occur during the calibration, is the not accurate value of the buffer solutions.

- Very small differences may be caused by the meters accuracy. The 
CPI-505 has accuracy of 0.002 pH, 1 digit, what practically means that in extreme situation results of 2 measurements may differ for 0.005 pH and this will be an acceptable error, because 1 measurement will be made with 0.002 pH error and second with -0.002 pH. 1 digit information explains difference caused by rounding up of the result on the last visible place on LCD (discretisation error).

If after making a two point calibration in pH buffers 7.00 pH and 4.00 pH (acidic conditions), the results are checked in 9.00 pH (alkali conditions),

in some cases the result may be 8.90 pH or 9.10 pH. This may occur when the electrode has unsymmetrical characteristic. Making at least a 3 point calibration with alkali, neutral and acidic buffers may prevent from such errors. Sometimes the measurement results are unstable, the quality of the electrode has the crucial influence. In most cases the reason for slow drifting of the result, it’s unstability or prolonging the time of stabilisation is the clogged junction, broken electrode or contaminated membrane.

Often this situation  happens when wrong kind of electrode was chosen for the kind of measured solution.

Leaving the electrode for few hours in the distilled water or placing it in water with detergent may eliminate this symptoms, especially if the measurements were done in solutions with deposits, fats or oils. The electrode which for a long time wasn’t used may have the junction clogged by KCl crystals, what may be removed by placing the electrode in distilled water and if this doesn’t help one may try to immerse the electrode in the thiourea in hydrochloric acid. Heavily contaminated electrode may be cleaned in chloroform and deposits of iron in 2N HCl. The electrode life may be prolonged by storing it in saturated KCl solution. Depending on the kind of measured solution or substance proper kind of electrode should be chosen. They differ one from another with shape, membranes look, kind of junction and body. Using electrode not proper for the kind of measured solution may cause it’s damaging and make the measurement impossible.

##### ION SELECTIVE MEASUREMENTS

# NOTICES ABOUT ION SELECTIVE MEASUREMENTS

The ion selective electrode changes it’s potential depending on the concentration (activity) of the individual ions in the solution. The main rule of the measurement is based on the linear dependence of the electrode potential from the logarithm of ion activity in the solution and is expressed by Nernst’s equation:

E = Eo + 2.303 RT/nF  log (a j)

where:	E	-	Electro Motoric Force of the measuring cell which consists of the ion selective and reference electrode in the measured solution (V)

Eo	-	normal ion selective electrode potential, depends mainly on the internal electrolyte activity of the ion selective electrode and kind of reference electrode (V),

ai	-	activity of the measured ion,

n	-	valency of the measured ion,

R	-	gas constant (8.31 J / K mol ),

T	-	temperature in K,

F	-	Faraday's constant (96487 As/val)

The measurement of ion activity is based on measuring of the potential of ion selective electrode in relation with the reference electrode. In most cases this electrodes are in separate housings. The reference electrode should have an appropriate external electrolyte. This electrodes have a construction which enables replacing the electrolyte to the right one for the specific ion selective electrode. The electrode manufacturers give information what electrolyte should be used depending on the measured ions.

| Measured ION | External electrolyte |
| --- | --- |
| F-, Cl-, Br-, J-, S2-, CN-, Ag+, Cu2+, Cd2+ | 1,0 M KNO3 |
| Ca2+ | 1,0 M KNO3 lub 0,1 M KCl |
| NH3 | 1,0 M KNO3 lub 0,1 M CH3COOLi |
| K+, Li+ | 0,1 M NH4NO3 |
| Na+ | 0,1 M NH4NO3 lub 0,1 M CH3COOLi |
| NO3- | saturated K2SO4 |

Table 3. The most often used electrolytes.

The ion selective electrodes differ one from another with their characteristics. The typical shape of the characteristic of the ion selective electrode is showed on the picture 13.

Pic. 13.

The best results may be achieved when the measurements are done in ranges that correspond with the linear parts of the characteristic. The essential electrode calibration on the sample solutions in that case may be done with larger difference of concentration of this solutions. The measurements made in ranges where the characteristic bends require choosing the sample solution which differ very little, because this lowers the electrode error. The measurements which are done in the ranges which correspond with lower part of the characteristic will have bigger and bigger errors, because of the lowering of electrode sensitivity.  For the bi-valent ions the electrode sensitivity is twice lower than for mono-valent ions. Presence of other ions may have influence on the measurement result. The ions which interfere the electrode are given by it’s producer, that’s why by accurate measurements knowledge of the ion composition in the measured solution is important and it is necessary to act according to the hints which enable lowering the measurement error.

The condition of receiving accurate results is calibration and measurements in solutions with this same temperature. The best results are achieved in temperature close to 20 oC. High temperatures decrease the sensitivity and life time of the electrode. Large differences between the temperatures of calibration solutions and the measured samples, although using the temperature compensation, introduce error connected with lack of precisely determined isopotential point of the electrode.

Essential influence on the ion measurements has the pH value.

During accurate measurements the pH value of the measured sample should be identical with the pH value of the sample solutions. For each electrode a pH range is given in which it is working without interference.

Exact complying with the electrode producer’s hints is the basic condition of success during measurements.

Preparation of the measured solution has a great influence on the final result. The fundamental meaning by the measurements of high concentration solutions has the phenomenon of ion activity changes together with change of their concentration, what may be described with the formula:

ai = fi • ci

where:	ai	-	ion activity in the solution,

fi	-	ion activity coefficient,

ci	-	ion concentration in the solution in  mol/l or g/l.

Stabilisation of the ion activity in the high concentration solutions is the condition of proper measurement. Adding the stabilising factor, recommended by the electrode producer, enables linearisation of the electrode characteristic for this concentrations, increases it’s sensitivity and enables for stable measurements.

Ageing of the electrode has an influence on the lowering the electrode sensitivity, prolonging the reaction time, shrinking the measuring range and increasing of the membrane resistivity.

# Entering the parameters of the ion meter

# Before the calibration it is necessary to make the activities described in the chapter 6. Additionally it is necessary to choose the unit in which we are going to make the calibration and measurement.

## Choosing the unit

The result may be displayed in: g/l. M/l ppm or pX. To choose the unit:

-	in the ion meter mode press and hold the , button till displaying in the upper row of LCD a   (unit) sign;

- with  ,  buttons choose the unit displayed on the right side 
(Pic. 14).

Pic. 14.

- enter the measurement mode by pressing the  button.

# ION SELECTIVE ELECTRODES CALIBRATION

Because of the not linear characteristic of the ion selective electrodes, the values of sample solutions should be close to the forecasted value of the measured solution and the ratio of the concentration shouldn’t be greater than 1:100. Usually in practice the prepared sample solutions have 1:10 ratio and by 3 point calibration 1:10. It is advised to prepare the sample solutions and measured solutions in this same volume. Differences in volume may have some influence on the result. The most accurate results are achieved in this same temperature in which the calibration was done. During calibration it is possible to choose the m mol weight of one of twenty ions stored in the memory of introduce mol weight and valency of ion which is not stored in the memory. It enables changing the unit in which the measurement is done for the calibrated ion (ex from g/l to M/l) and reading the result in new unit.

## Entering the calibration parameters

The meter remembers the values of calibration points independently for each of 3 ion selective electrodes and after changing the number of the electrode  restores values on which it was calibrated recently. After changing the ion the values of calibration points entered in g/l or ppm will change. Therefore It is very important to remember about the need of introducing the value of calibration points after changing the ion.

Caution:	the changes will have influence on counting just after next calibration.

It is necessary to:

a.	Enter the ion and values of the sample solutions used for calibration.

To do so it is necessary to press the  , button, in the upper row a   (points of calibration) symbol will be displayed and in the lower row a  (user settable value of the sample solution pic. 15);

Pic. 15.

b.	Press the  button. In the upper row the  symbol will be displayed  and in the lower one the number of the ion. According to the table placed on the meter (which is shortened form of table given below) with the buttons  ,  choose the number of ion according to used electrode:

Table 4.

| no | ion | weight |  | ion | weight |
| --- | --- | --- | --- | --- | --- |

| 1 | H+ | 1.007 | 11 | F - | 18.998 |
| --- | --- | --- | --- | --- | --- |

| 2 | Ag+ | 107.868 | 12 | I - | 126.904 |
| --- | --- | --- | --- | --- | --- |

| 3 | Br- | 79.904 | 13 | K+ | 39.098 |
| --- | --- | --- | --- | --- | --- |

| 4 | BF-4 | 86.805 | 14 | Li+ | 6.941 |
| --- | --- | --- | --- | --- | --- |

| 5 | Ca2+ | 40.078 | 15 | Na+ | 22.990 |
| --- | --- | --- | --- | --- | --- |

| 6 | Cd2+ | 112.411 | 16 | NH4+ | 18.038 |
| --- | --- | --- | --- | --- | --- |

| 7 | Cl - | 35.453 | 17 | NO3- | 62.005 |
| --- | --- | --- | --- | --- | --- |

| 8 | ClO4- | 99.450 | 18 | Pb2+ | 207.200 |
| --- | --- | --- | --- | --- | --- |

| 9 | CN - | 26.018 | 19 | S2- | 32.066 |
| --- | --- | --- | --- | --- | --- |

| 10 | Cu 2+ | 63.546 | 20 | SCN- | 58.083 |
| --- | --- | --- | --- | --- | --- |

In case of measurement of ion which is not listed in the table one should:

-	press and hold the  button till displaying in place of ion number a  (user) symbol;

-	short press the   button, in the lower row a , symbol will be displayed and in the upper the last entered mol weight (Pic. 16);

-	with the  ,  buttons enter the mol weight of the measured ion;

Pic. 16.

-	short press the  , button in lower row a , symbol will be displayed and in upper the last entered valency; (Pic. 17);

-	with the  ,  buttons choose the right valency of the ion 
(, , , );

By short pressing of the   button return to the ion choosing screen, or enter the measuring mode by pressing the chosen function button.

The entered mol weight and valency will be taken by the meter into consideration after the calibration.

Pic. 17.

c.	Press the   button. In the lower row a  – (point 1) symbol will be displayed – first sample solution and in the upper one value of the sample solution (pic. 18). With  ,  buttons enter the value of sample solution used in this point.

Pic. 18.

To move to the second point of calibration introduction it is necessary to press the   button, in the lower row a  (point 2) symbol will be displayed and in the upper the stored value of sample solution for this point. It is possible to control or change the values of sample solutions in the rest of points as it was described above. In every moment one may finish entering the values of sample solution by pressing the  button to return to   mode or  button to enter the measuring mode.

The values of sample solutions entered by user are stored till their changing to other.

During next calibrations, if the values of previously used sample solutions were not changed it is not necessary to make the actions described in this point.

Caution: the values of calibration points are stored in pX and than counted for displaying in the chosen unit. That’s why during entering the small values in units different than pX the changes of values will be visible after pressing the  or  buttons a few times. Pressing and holding  or  button (repetition) eliminates this inconvenience.

14.2. Actions during calibration

- choose  the electrode number under which the calibration will be stored (point 6.3;)

- choose the unit in which the calibration will be done (point 13.1);

prepare the electrode to work according to the producer’s instructions and mark it with number responding to number chosen in the meter (, , );

- connect with the meter the prepared electrodes: ion selective and reference and temperature probe, use the right connectors pH/mV/Ion and temp;

- prepare the sample solutions according to the instructions of ion selective electrode producer.

Under each number next electrode in different sample solutions may be calibrated (after entering those values to the meter’s memory).

14.3. Calibration of the ion selective electrode

After preparing the meter to the calibration as it was described above it is necessary in ion function to:

a.	press and hold the  button till the appearance of the CAL (pic. 19) symbol on the display; the old parameters of calibration are deleted. Between the upper and lower row of the display a P1 (point 1 of calibration) symbol will be displayed;

Pic. 19.

b.	put the electrodes and the temperature probe to the buffer solution, wait till stabilisation of the result. The result may be different than the value of the sample solution. After stabilisation of the result press the  button. In this moment the result will blink what will inform about the storing the calibration result, in this same time in the upper row the corrected value of measurement will be displayed – it will be equal to the value of used sample solution. the P1 symbol will be changed to P2 (second calibration point);

c.	continue the calibration in other buffer solutions, washing the electrodes and temperature probe before every immersing in the sample solution and acting according to the point b.

The attempt of calibration of next point in the same solution causes appearing  symbol and it will stay on this point. The accidental exchange of solutions in relation to points of calibration causes the appearing the  symbol and exit from the calibration mode with deleted electrode characteristic.

Calibration may be finished in every moment by pressing any function button.

After calibrating one electrode it is possible to calibrate two others, choosing the two remaining symbols according to point 6.3.

If after choosing the electrode number and entering the calibration mode one will exit it without making the calibration the earlier stored characteristic will be deleted and a standard characteristic for the chosen ion will be adapted.

# ION SELECTIVE MEASUREMENT

The ion selective measurement is made in mV (linear scale) with taking the temperature under consideration. After choosing the unit M/l or g/l the results in mV are converted and displayed in M/l or g/l (logarithmic scale).

During accurate measurements it is recommended to make the calibration and measurements in this same temperature. The meter with connected temperature probe may be used simultaneously for temperature measurements during the ion selective measurements.

The best measurement accuracy may be achieved if the measurements are done right after calibration. The calibration is essential.

# Small changes of the result close to the measured value depend on the quality of used electrode.

To make the measurement it is necessary to:

- correct chosen ion electrodes (measuring and reference) connect to the right connectors (pH/mV/Ion and Gnd);

- connect the temperature probe (temp);

- choose the number of electrode under which it was calibrated (point 6.3.);

- choose the unit according to point 13.1;

- set the required resolution according to point 6.2;

- insert the electrodes and the temperature probe to the measured solution;

- bring the temperature of the measured solution to the sample solution temperature;

- make the measurement and after the stabilisation read the result.

Pic. 20.

##### ORP (mV) AND TEMPERATURE
 MEASUREMENT

# ORP (mV) MEASUREMENT

pH  meter CPI-505 is an accurate mV meter (Oxidation Reduction Potential). The measurement may be done with special redox electrode or during titration. The result may be checked after choosing the mV mode with  button 
(Pic.  21).

Pic. 21.

# TEMPERATURE MEASUREMENT

The temperature measurement is done as follows:

-	by pressing the  button switch the meter on;

-	to the Chinch connector connect the temperature probe, on the display a  symbol will be displayed;

-	put the temperature probe to the measured solution;

-	wait till the stabilisation of the value and read the result in the lower row of numbers.

The meter co-operates with the Pt-1000 probe. Depending on its class the accuracy of the measurement changes.

CAUTION: break in the circuit of the temperature probe switches the meter to the manual temperature compensation mode. It is signalised by changing the  symbol to  symbol. On the display the value of the temperature entered by the user is displayed.

Blinking -50°C value during measurement in positive temperatures informs about short circuit in the temperature probe.

##### OTHER

# CLOCK WITH DATE

After choosing the time mode with  button the meter will display the actual time. By pressing the  button one may display the date, backlight mode, auto switch off time and number of the software version.

## Clock

The time is displayed in two rows on the display (Pic. 22). In the upper one the hour and minutes are displayed and in the lower one seconds. The way of changing the hours is described in point 18.3.

Pic. 22.

## Date

The date is displayed as follows: Month – Day – Year. (Pic. 23).  In the upper row the month and day are displayed and in the lower the year.

Pic. 23.

## Setting the time and date

The setting mode of the actually displayed parameter (hour or date) one enters by pressing and holding the . button. The position which we are going to change starts blinking, the value is changed with  or . buttons. The pulsating position is chosen by short pressing of the   button. There is no way to set the seconds, they are deleted after the setting mode is left. Press the  button to leave the setting mode.

## The LCD backlight mode

In the time mode press the  button till displaying in the upper row of LCD a  sign (Pic. 24). In the lower row an ,  or  parameter will be displayed.

- the backlight is switched off;

- mode of automatic switching on of the backlight for 60 seconds    after each pressing of any button;

- the backlight is always on.

Pic. 24.

With  or  buttons one may choose the backlight mode.

Return to the time mode after pressing the  button.

## Brightness control

Controlling the brightness of the backlight is especially important during the work In different conditions of lighting. During the process of choosing the backlight mode (p. 18.4) it is possible to control the brightness by short pressing the  button. In the place of  or  parameter the value of brightness in % will be displayed. With  or  buttons it is possible to set the brightens value from 10 ÷ 100%, with 10% step.

Pic. 25.

Return to the backlight mode be pressing the  button, and to the measuring mode by pressing any function button.

## Readout of the software version number

In the  time function press the  button till displaying a screen, as on the picture below (Pic. 26). In the upper row the software version number is displayed

Pic. 26.

Return to the time mode after pressing the  button.

# STORING AND READOUT OF THE RESULTS

The meter enables storing of 950 results of the actually measured function. The results are stored in EEPROM memory, which is non-volatile, so the data isn’t lost even after complete lack of power. Before starting the work it is necessary to choose the parameters of storing or readout of the stored results.

## Parameters of storing and readout from the memory

The parameters are changed in the readout mode. It may be entered from every measuring function by pressing and holding the  button, until on the display the number of last stored result will be displayed. This number is displayed on turns with the stored result.

Before starting the storing one has to choose the kind of results collecting: on request or automatically as a series, and also the way of displaying the result.

Next pressing of the  button shows screens with following functions which may be changed:

a.	  - taking or printing series or single results.

Pic. 27.

With  or  button choose in lower row   or  symbols  
(Pic. 27).  Choosing  activates automatic storing of the results and  single, after every pressing of  button.

b.	  - time interval during series taking (Pic. 28).

Pic. 28.

Value of the time interval is displayed in the lower row of digits, and the informative symbol   in the upper row.

Buttons  and  are used to set the chosen time in minutes and seconds. The shortest time is 1 second and the longest 60 minutes. Holding the button makes changing quicker. (repetition).

In case of choosing the  parameter to  the position  is not displayed.

c		 - Way of displaying the stored results

- successively number of sample, result, time and date of storing the result.

- successively number of sample and result

Changing with  or  buttons.

Return to the results readout display after pressing  button.

Exit from the readout mode after pressing of the button of chosen measuring function.

## Storing the single measurements to the memory

If according to previous section storing of single results was chosen, every pressing of  button stores the measured result. The results are stored as the next ones after the last stored. If someone was checking the earlier stored results and didn’t return to the last one the results won’t be deleted and the value will be stored from the first empty position. If the user wants to store the result from the chosen number first he has to delete the results (as described in the point 19.5) and next start storing the results by pressing the button.

During storing of the results the number of result will be displayed for a moment.

If after pressing the  button instead of number an  sign will be displayed it informs that the maximal number of results was stored.

## Storing the measuring series

There is possibility to store series of measurements in the meter’s memory. To do so one has to :

-	choose the serial collecting the measurements (p 19.1.a);

- enter the time interval (p 19.1.b);

delete the stored results starting from the chosen one (p 19.5);

-	 by the proper button choose the function which results one wants to store;

-	with  button start taking the series. The measurements will be stored starting from the first free number.

The red LED in the  button will begin flashing informing that the meter is during storing of series of measurements. Before each  storing of the result the number of measurement will be displayed for a moment. The taking of the series may be stopped by pressing   button, any function key or by filling the memory. If the series results are not filling the memory totally one may start to take the next series till the end of memory.

## Reviewing the results

Reviewing of the stored results is started from the measuring mode, by pressing and holding the  button until the number of last stored result on turns with it’s value will be showed.

Every pressing of the  or  button shows the next or previous number and result with time and date if  function was chosen (p 19.1.c).

In this mode the buttons  ,  work with repetition and after longer holding the numbers change very quickly till stopping on the highest or lowest number.

The reviewing mode may be left by pressing any function button.

## Deleting the stored results

To delete the stored results one has to:

-	press and hold the  button;

-	with buttons  ,  set the number of measurement from which we want to start deleting the memory;

-	press and hold the  button what will delete the stored results from the chosen one till the end of the memory. On the display in the place of results there will be  sign displayed what confirms the deleting;

-	to exit the reviewing mode press any function button.

If one wants to clear all the memory the deleting should be started from the first number.

# CO-OPERATION WITH A PC

Connecting the meter with a PC enables storing data directly on the computer, what eliminates the limitation of the number of stored results. The PC should be equipped with USB port. For the transmission a special software prepared by our company may be ordered. The software is delivered on a CD. After inserting the CD to the drive the installation program will start automatically. It is necessary to follow the given instructions.

In the upper wall of the meter a USB port is placed, it enables connecting with a PC using a cable. After connecting the meter and the PC should be switched on and the transmission software should be run. We have two options available:

-	“Collect series” is used for collecting results of current measurement. After choosing this option we will see a window with the result of current measurement. Collected and stored will be only elements which are marked in the field “Send”. It is necessary to set the number of measurements which are to be stored and time breaks between storing it. On the basis of this data the software will count the time of collecting the whole series. The series is stored in temporary file . In case of lack of power the so fare collected data will be stored in a file “NoNamexx”. The collecting is started by pressing the “Collect” button;

-	“Download data from memory” enables sending the chosen part or whole of the data stored in the meter’s memory to a file. In option “Collect” we mark the data we want to be sent. The transfer is started by pressing the button “Download”

Caution:	the meter and PC should be switched on after connecting the cable.

# TECHNICAL DATA

pH MEASUREMENT:

| Range | Resolution | Accuracy (±1 digit) |
| --- | --- | --- |

| -2.000 16.000 pH | 0.001 / 0.01 pH | ±0.002 pH |
| --- | --- | --- |

INPUT IMPEDANCE:						1012

TEMPERATURE COMPENSATION:			manual/automatic

RANGE OF COMPENSATION:				-5.0 ÷ 110.0 oC

pH ELECTRODE CALIBRATION: 	 			automatic,

in 1 ( 5 points

RANGE OF RECOGNITION AND ENTERING OF THE pH BUFFER SOLUTIONS

| Calibration point | Range |
| --- | --- |
| 1 | 0,800 ( 2,100 |
| 2 | 3,900 ( 4,100 |
| 3 | 6,800 ( 7,100 |
| 4 | 8,900 ( 10,200 |
| 5 | 11,500( 14,000 |

AUTOMATIC CHANGE OF THE pH BUFFER VALUE TOGETHER WITH THE TEMPERATURE CHANGE, FOR SAMPLES CONSISTENT WITH NIST (table page 19) IN RANGE   			0 ÷ 60 oC

THERMAL STABILITY OF ZERO:		0.0005 pH/ oC

# ION SELECTIVE MEASUREMENT:

| Range | Resolution | Accuracy (±1 digit) |
| --- | --- | --- |

| -2.000 16.000 pX | 0.001 / 0.01 pX | ±0.002 pX |
| --- | --- | --- |
| 0 ÷ 100 M/l | 0.01 / 0.1 % | ±0.25 % |
| 0 ÷ 1 000 g/l | 0.01 / 0.1 % | ±0.25 % |
| 0 ÷ 1 000 000 ppm | 0.01 / 0.1 % | ±0.25 % |

INPUT IMPEDANCE:						1012

TEMPERATURE COMPENSATION:			none

CALIBRATION OF THE ELECTRODE:		in 1 ( 5 points

mV MEASUREMENT:

| Ranges | Resolution | Accuracy (±1 digit) |
| --- | --- | --- |

| -1000 1000 mV | 0.1 mV | ±0.1 mV |
| --- | --- | --- |

INPUT IMPEDANCE:				1012

# TEMPERATURE MEASUREMENT:

| Range | Resolution | Accuracy* (±1 digit) |
| --- | --- | --- |

| - 50.0 ÷ 199.9 oC | 0.1 oC | 0.1 oC |
| --- | --- | --- |

*	accuracy of the meter. Final accuracy of the measurement depends on the accuracy of the used PT-1000 probe

TEMPERATURE PROBE:				platinum resistor Pt-1000

ACCURACY OF THE PROBE IN RANGE 	0 ÷ 100 0C:

FOR PT1000B RESISTOR:				±0.8 0C

FOR PT10001/3B RESISTOR:			±0.27 0C

# OTHER:

# MEMORY OF THE RESULTS:				950 results

# WORKING TEMPERATURE:		0 ÷ 40 oC

POWER:		power adapter 12V/100mA.

POWER CONSUMPTION:		60 mW

SCREEN:		LCD 55 x 45 mm

DIMENSIONS:		200 x 180 x 50 mm

WEIGHT:		620g

# EQUIPMENT

The standard equipment for the CPI-505 meter is:

      - Combination glass electrode EPS-1

      - Temperature probe  Pt-1000B (standard);

      - Power adapter 12V/100mA.

      - Software for collecting the data on the PC

      - USB cable

      - User's manual with warranty.

The additional equipment available for this meter is:

1.	pH electrodes for special purposes (sewage, wastes, meet etc.);

- Temperature probe Pt-1000 1/3B with higher accuracy;

- Ion selective electrodes;

- ORP measuring electrodes;

- Reference electrodes

# W A R R A N T Y

The ELMETRON company ensures a 24 months warranty for the 
 CPI-505 pH meter number:.......................................................................................

In case of damage the producer will repair the meter within 14 days from the day of delivery.

The warranty doesn’t cover the damages caused by usage not in conformity with the user’s manual, using wrong power adapter, mechanical damages and damages caused by repairs made by unauthorised persons.

The pH electrode and conductivity cell, have the warranty of the producer.

NOTICE: Before sending the meter to us please contact the firm by phone or email.

When sending the meter, the used electrode, conductivity cell, oxygen sensor, temperature probe and power adapter should be also included.

Date of production....................................................................................

Date of sale..............................................................................................

Date of warranty expiry.............................................................................

(  Sp. j.

41-814  Zabrze   Witosa 10   Poland

tel.  +48 32 2738106   fax  +48 32 2738114

www.elmetron.com.pl

e-mail: info@elmetron.com.pl