(  Sp. j.

41-814  Zabrze   ul. Witosa 10   POLAND

tel.  +48 32 2738106    fax  +48 32 2738114

www.elmetron.com.pl

e-mail: info@elmetron.com.pl

LABORATORY

pH METER

CP- 501

CP- 502

USER’S MANUAL

USER’S MANUAL

LABORATORY

pH - METER

CP-501, CP-502

Before use please read the instruction carefully!

### W A R R A N T Y

The “ELMETRON” company ensures a 24 months warranty for the pH meter CP-501, CP-502* number:

..................................................

In case of damage the producer will repair the meter within 14 days from the day of delivery.

The warranty doesn’t cover the damages caused by usage not in conformity with the users manual, using wrong power adapter, mechanical damages and damages caused by repairs made by unauthorised persons.

The pH electrode, conductivity cell, oxygen sensor have a separate one year warranty of the producer.

NOTICE: Before sending the meter to us please contact the firm by phone or email.

When sending the meter, the used electrode, conductivity cell, oxygen sensor, temperature probe and power adapter and warranty with date of purchase should be also included.

Date of production.........................................

Date of sale...................................................

Date of warranty expiry..................................

* cross out the unnecessary

- 4 -

1. EXPLOITATION NOTICES

Dear User!

We present you a device distinguished by accuracy according to the technical data and by a high stability of the displayed results. We believe that the measurements will not cause you any trouble and that the meter will operate without any inconvenience. Wide range of additional functions requires careful reading of the manual in other case some of the features may stay unused or using the meter may cause you a troubles.

The employing of good-quality electrodes cells and probes and their replacing after a suitable time ensures obtaining of high measuring parameters. We want to call your attention to the fact, that this equipment has a much shorter working life than the meter. Typical symptoms of an improper operation of the electrode are deterioration of final result stability, it’s flowing as well as a higher measuring error. Part of the users has problems, arising from employing electrodes not being preconditioned before the measurement or making the measurements without removing the shielding ring from the liquid junction or taking measurements with a plugged junction. To avoid this situation it is necessary to choose the proper kind of electrode for solutions which are going to be measured ex. special electrodes for the sewage, liquids with deposits, meat cheese etc. Therefore, if you observe improper operation of the device, please take control measurements with another electrode or check the used electrode with another pH-meter. Generally the deterioration of the meter work is caused by the electrode and not by the meter.

The essential feature of our  products is their low failure frequency. However if your pH-meter will fail, our firm immediately performs its warranty repair.

During ion selective measurements a very often case is worsening of the results stability, this is connected only with the quality of the ion selective electrodes.

We wish you a pleasant and trouble-free work with our meter.

TABLE OF CONTENT

I. Introduction

1.	Exploitation notices	  4

2.	The characteristic of the meter	  5

3.	What is the meter designed for	  6

4.	Outside view	  7

5.	Switching the meter on and off	10

6.	Preparation to work	11

6.1.	Choosing the kind of temperature compensation	11

6.2.	Changing the resolution of the measurement	12

6.3.	Changing the electrode number	13

##### II. pH measurement

7.	Preparation of the pH electrode	16

8.	Calibration	17

8.1. 	Calibration mode	18

8.2. 	Calibration with use of NIST sample solutions	19

8.3. 	Introducing the pH values of buffer solutions	20

8.4.  	Entering the calibration parameters	21

8.5.	pH electrode calibration 	23

8.5.1. Calibration with automatic temperature compensation	24

8.5.2. Calibration with manual temperature compensation	26

9.	Checking the electrode’s condition	26

10.	pH measurement	27

10.1.	Measurement with automatic temperature compensation	27

10.2 	Measurement with manual temperature compensation	28

- Notices about the temperature compensation and

interpretation of the pH results	29

##### III. ORP (mV) and temperature measurement

12.	ORP (voltage) measurement	32

13.	Temperature measurement	32

##### IV. Other

14.	Clock with date, 	34

14.1.	Time display	34

14.2.	Date display	34

14.3.	Readout of the meter’s software version 	35

14.4.	Setting the time and date 	36

15.	Storage and readout of the results	37

15.1.	Storing or printing	37

15.2.	Parameters of storing and reading from the memory	37

15.3.	Storing the single measurements in the memory 	39

15.4.	Storing the measuring series	39

15.5.	Reviewing of the results	40

15.6.	Deleting the stored results	40

- 2 -

16.	Printouts on the printer	41

16.1.	Choosing the direction of sending the data	41

16.2.	Printout of the result – single or serial	41

16.3.	Printout of the results stored in the memory 	42

16.4.	Internal printer maintenance in CP-502	42

17.	Cooperation with a PC 	43

18.	Technical data	44

19.	Equipment	46

- 3 -

##### I. INTRODUCTION

- 8 -

The keyboard has the following keys:

-	switches the meter on and off;

-	chooses the pH measuring function;

-	chooses the mV measuring function;

-	displays the time and date;

-	longer pressing of this button enters the calibration mode (CAL symbol displayed). Short pressing in this mode confirms the calibration result;

-	short pressing causes storing or printing of single result or measuring series, longer pressing enters the stored results readout mode.

-	chooses the entered parameter;

, 	-	buttons for entering the parameters.

-	enables pulling out and slipping in of the paper by the printer (only in CP-502).

On the back wall of the meter there are inputs with below given symbols:

pH/mV	-	BNC-50 connector for connecting combined or measuring pH electrode and ORP or ion selective electrodes;

Gnd	-	connector for reference electrode;

temp	-	Chinch connector for temperature probe;

POWER	-	connector for power adapter.

On the right side of the meter there are RS-232 and Centronics outputs.

- 5 -

2. THE CHARACTERISTICS OF THE METER

The pH meters CP-501 and CP-502 belong to the newest generation of measuring devices which offer wide range of additional functions. Technical data, extra features and way of maintenance are identical in both models. The CP-502 has additionally build in thermal printer which enables printing the actual and stored in the memory results.

The meters ensures high accuracy and repeatability of the readings. The newest generation electronic elements used in the meter made it’s memory independent to power supply. The internal clock is  powered from battery which is enough for many years. The meter is equipped with custom LCD display, which enables simultaneous observing of the measured function, temperature value and additional symbols which make working easier. Waterproof housing makes working in difficult conditions possible.

Main features of the CP-501 and CP-502 are:

- high accuracy and stability of the reading;

- measurement of pH, mV and temperature;

- automatic and manual temperature compensation;

- electrode calibration in 1 to 5 points;

- automatic recognition of pH buffers and standards;

- imposed values of standard solutions with possibility of changing their value;

- option of automatic introduction of temperature correction on the value of pH buffer solutions (NIST norm);

- information about the condition of the pH electrode;

- independent storing of three characteristics of pH electrodes;

- internal datalogger for 200 (optionally 450 or 950) measurements with time date and temperature;

- taking series of measurements with set time interval;

- RS-232 output

- possibility of printing the measurement results or values stored in the memory on standard printer ;

- real clock and date.

- 6 -

3. WHAT IS THE METER DESIGNED FOR

The pH meters CP-501 and CP-502 are  precise and easy to use meters designed for hydrogen ion concentration measurements in pH units and Oxidation Reduction Potential (mV).

The meters may be also used for accurate temperature measurement of solutions and air in °C.

The meters are being used in food, chemical, pharmaceutical, energetical industries, in water treatment stations, laboratories, agriculture, universities, scientific laboratories etc.

CP-501 and CP-502 are prepared to work with all types of combination pH electrodes equipped with BNC-50 connector. There is also possibility of using a separate measuring and reference electrode. The meters co-operate with Pt-1000 temperature probe with Chinch connector.

The meter may collect measurements taken as single or series of measurements with set time interval. Built in RS-232 interface enables connecting the meter wit a PC for sending the data. Centronics output enables connecting the meter with a standard printer. There is possibility of printing the data collected in the memory or current results of the measurement.

CP-502 enables printouts with use of the built in thermal printer (60 mm width).

In case of necessity to collect series longer than the capacity of memory it is possible to use special PC software offered by our company.

- 7 -

4. . OUTSIDE VIEW

On the left side of the meter there is a LCD display (pic. 1) on which depending on the chosen function following symbols are displayed:

- result of the pH measurement in pH units;

- result of the mV measurement in mV;

- time and date.

Choosing the function is done by specific buttons and is signalised by lighting LED diode placed on the button.

Simultaneously with the result a measured temperature value is displayed in oC. Symbols of the units are displayed by the results.

Pic. 1.

Beside the temperature value a  symbol for automatic temperature compensation is displayed or  for manual. CAL symbol on the left side of the display informs that the meter is in calibration mode.

On the left side of the display a number of chosen electrode is showed (E1; E2; E3) which inform about the number of characteristic which will be taken into consideration during all calculations. On the right side of the display a battery symbol which informs about the battery condition is displayed. During calibration between the upper and lower row of numbers number of the calibration point is displayed ( P1, P2, P3, P4, P5).

In the MODE all parameters introduced by the user are displayed and also the value of atmospheric pressure measurement (in oxygen measurement function). The keyboard  (pic. 2) placed under the display is used for switching the meter on and off, choosing the measuring function, calibration, entering the parameters, printing and storing the results in the memory.

- 12 -

6.2 Changing the resolution of the measurements

The measurement results may be displayed with chosen resolution. To change it one should:

- in the measuring mode press the  button, a  (resolution) sign will be displayed. (Pic. 4)

- Using keys  or  one may choose:

- (low) resolution of the measurement;

- (high) resolution of the measurement.

Pic. 4.

For the pH measurement:

- resolution of the measurement 0.01 pH;

- resolution of the measurement 0.001 pH.

Return to the measuring mode by pressing the chosen function button

- 9 -

Pic. 2.

- 10 -

5. SWITCHING THE METER ON AND OFF

The meter is switched on by pressing the  button. The meter tests the memory and display on which all symbols are displayed (pic.3).

Pic. 3.

If the test was successful, after about 1.5 s the meter switches it self automatically to the measuring mode, in which it was switched off. If an  sign will be displayed it means that the meter has lost the factory settings and requires the service repair. If after the 1,5 s all symbols will be continuously displayed it informs that the calibration parameters of electrodes or cells  were lost.

After pressing the  button the meter will take standard characteristics:

- shift = 0 pH, characteristic slope = 100% for pH electrode;

And will enter the measuring mode. It will be necessary to calibrate the conductivity cell and  pH electrodes.

The meter is switched off by pressing and longer holding of the  button

- 11 -

6. PREPARATION TO WORK

Before starting the work one should:

-	join the power adapter plug to the Power input;

-	to pH/mV input (BNC-50) join the ready to work combination pH electrode, ion selective electrode or redox electrode;

- in case of using the temperature probe it should be connected with the chinch temperature input - temp;

- In case of work with PC join the 5XX-PC cable with RS-232 input

- In case of printing results on the printer connect the printer cable with Centronics output.

- switch the meter on by pressing the  button.

6.1. Choosing the kind of temperature compensation

The meter switches it self to the automatic temperature compensation mode automatically after joining the temperature probe, after disconnecting it the meter enters the manual temperature compensation mode. In ATC mode near the displayed temperature a  symbol appears. Manual temperature compensation is indicated by  symbol near the value entered by the user, it’s value may be changed using  or  keys.

- 16 -

7. PREPARATION OF THE pH ELECTRODE

The electrode should be prepared to work according to the producer instructions. If the instructions weren’t given please follow the steps:

- new electrode should be put into distilled water or in saturated KCl solution for about 5 hours:

- Before starting the measurements protecting rings (if used in this kind of electrode) should be removed. The ring placed on the junction – lower part of the electrode should be moved up along the electrode’s body and the upper, which protects the KCl refilling hole, down along the body.

Removing the lower ring is essential, in other case the electrode won’t measure.

Upper ring should be removed during measurements of high temperature solutions or to protect the junction during measurements in solutions with deposits or oils. Sometimes instead of ring a cork is used;

- during measurements in laboratory it is advisable to use an electrode holder;

- after every measurement the electrode should be washed in distilled water;

- excess of liquid on the electrode should be removed by gentle touching the glass with a tissue paper;

- after work the electrode should be stored in one of the above given solutions. The protecting rings should be moved on the junction and upper hole;

- in case of long breaks between the measurements the electrode should be stored, after drying, in the packaging;

- after taking the electrode of the package the eventual deposit should be removed using water;

- before using the electrode it should be placed in distilled water for about 2 hours;

- if the construction of the electrode enables refilling the electrolyte, it should be controlled and refilled periodically by the upper hole in the electrode’s body (usually as the electrolyte a KCl solution is used).

CAUTION:	storing of the electrode in distilled water shortens it’s life time and may cause shifting of the zero point.

.- 13 -

6.3. Changing the electrode number

If in the meter’s memory more than one electrode characteristic was stored, it is possible to replace the electrodes without the need of calibration. This option is quite useful in case of using different kinds of electrodes ex. for sewage, and clear water etc. It is necessary to connect the earlier calibrated electrode, marked with the number which responds to the number which was stored in the memory and choose this number.

To do so in the pH or ion measuring mode:

Press the   button, than using  or  buttons choose the electrode number (symbols , , ) under this number the calibration results will be stored (pic. 5). Above the electrode number following signs will be displayed:

-	under this number there is no characteristic stored and producers values are provided.

-	under this number there are values of last calibration stored.

-	the last calibration showed that the electrode is loosing it’s efficiency and in short time it’s calibration won’t be possible. (only for electrode calibrated in pH function)

Pic. 5.

Pressing of any function button returns the meter to measuring mode.

- 14 –

- 15 -

### II. pH measurement

- 20 -

Table 3.

| Temp. 0C | Kind of buffer solution |
| --- | --- |
|  | 1 oxalate | 2 phthalate | 3 phosphate | 4 di-sodium tetraborate | 5 calcium hydroxide |
| 0 | 1.666 | 4.003 | 6.984 | 9.464 | 13.423 |
| 5 | 1.668 | 3.999 | 6.951 | 9.395 | 13.207 |
| 10 | 1.670 | 3.998 | 6.923 | 9.332 | 13.003 |
| 15 | 1.672 | 3.999 | 6.900 | 9.276 | 12.810 |
| 20 | 1.675 | 4.002 | 6.881 | 9.225 | 12.627 |
| 25 | 1.679 | 4.008 | 6.865 | 9.180 | 12.454 |
| 30 | 1.683 | 4.015 | 6.853 | 9.139 | 12.289 |
| 35 | 1.688 | 4.024 | 6.844 | 9.102 | 12.133 |
| 40 | 1.694 | 4.030 | 6.838 | 9.063 | 11.984 |
| 45 | 1.700 | 4.047 | 6.834 | 9.038 | 11.841 |
| 50 | 1.707 | 4.060 | 6.833 | 9.011 | 11.705 |
| 55 | 1.715 | 4.075 | 6.834 | 8.985 | 11.574 |
| 60 | 1.723 | 4.091 | 6.836 | 8.962 | 11.449 |

8.3. Introducing the pH values of buffer solutions.

Values of the pH buffer solutions set by the producer in the meter’s memory are given in the table 2. If the values of used buffer solutions are different than the given above the user should change the set values to the actually used.

Table 2.

| Point of Calibration | Resolution 0,001 | Resolution 0,01 |
| --- | --- | --- |
| 1 | 2,000 | 2,00 |
| 2 | 4,000 | 4,00 |
| 3 | 7,000 | 7,00 |
| 4 | 9,000 | 9,00 |
| 5 | 12,000 | 12,00 |

The range of possible changes is limited for each calibration point.

The meter doesn’t allow for introducing pH values in ranges other than those given in the table 3.

- 17 -

8. CALIBRATION

Before starting the measurement with new electrode, after long-lasting using, or before making measurements which require higher accuracy the electrode connected with the meter should be calibrated. Results of measurements done without calibration will have a great error. The calibration is done on the buffer or sample solutions. It depends on comparing pH value of the buffer solutions with the value displayed by the meter and next automatic introduction of correction which is taken into consideration during measurements. The calibration should be periodically repeated because during work the parameters of the electrode are changing what influences the accuracy. The frequency of this procedure depends on the demanded accuracy, number of the measurements carried out, conditions in which the electrode was used, temperature and value of the measured solutions. When the highest accuracy is required it is recommended to use sample solutions with certificates. In technical usage buffer solutions are used with lower accuracy but mostly with total values ex. 2.00 pH; 4.00 pH etc. In this manual to make it easier a buffer solution name was used for both buffer and sample solutions. For accurate measurements it is necessary to use fresh good quality buffers.

The first action taken before the calibration is entering the values of used pH buffer solutions to the meters memory. This should be done before the first calibration and repeated in case of changing the used buffers to other values.

During the calibration after immersing the pH electrode and temperature probe the meter automatically detects the value of the used buffer, if it was earlier stored in the meter’s memory.

There is possibility of calibration in minimum 1 buffer solution and maximum in 5. The more points of calibration is used the higher is the accuracy in the whole measuring range.

Calibration in one buffer solution can’t ensure high accuracy. If only one buffer solution is used it’s value should be close to the forecasted value of measured solution. If the required accuracy isn’t very high and the measurements will be done in the whole range the one point calibration should be done with buffer solution close to 7.00 pH. Thanks this the so called zero electrode shift will be eliminated. In other points a standard characteristic will be taken from the meter’s memory. This characteristic corresponds with theoretical efficiency of the pH electrode.

- 18 -

If measurements are done both in acids and alkalis and the measurements are not done at the ends of measuring range it is enough to calibrate the electrodes in 3 buffer solutions with values in range given in table 2 – calibration points 2, 3 and 4. In case of accurate measurements in the whole range it is recommended to calibrate the electrode in all 5 points given in the table. In CP-501 and CP-502 characteristic of the electrodes is approximated linearly between the calibration points.

Starting the calibration under one of the electrode numbers does not remove the set buffer values, but irreparably removes the electrode characteristic stored under this number.

There is no possibility of calibrating only one point and leaving the rest of data from the last calibration.

Memory in CP-501 and CP-502 enables independent storing of 3 characteristics of pH electrodes. This feature is especially useful when fast replacing of the electrodes is necessary or when it was broken. Those electrodes should be earlier calibrated and brought into memory under following symbols , , .

The order of using the buffer solutions is freely chosen.

8.1. Calibration mode

The user may choose two independent ways of action:

- use the values of pH sample solutions entered to the memory by the manufacturer. Those values are in conformity with the NIST norm. This type of calibration switches on the automatic correction introduction connected with the change of sample solution along with temperature changes.

- enter the value of pH buffer solutions depending on the actually used buffers, in the range given for each point of calibration;

Choosing the first option makes the calibration much easier. The  user doesn’t need to warm up or cool down the sample solutions to the temperature given by the user. This option includes some simplifications and when very high accuracy is required it shouldn’t be used. The meter stores independently 5 values of the buffers for each of the calibration modes described above.

- 19 -

8.2. Calibration with use of NIST sample solutions

In this mode 5 constant sample solutions values, according to NIST, are used. In the meters memory a table with dependence between the temperature and pH values for this 5 sample solutions is stored. This dependence is showed in the table 1. During the calibration the temperature of sample solution is measured and on it’s basis automatically the right pH value responding to this temperature is taken into consideration. Values between the points given in the table are approximated linearly by the meter.

Some sample solutions made according to the appropriate norms may differ on the third decimal place. The meter enables setting of the appropriate buffer solution values differing  +/_ 0.010 pH from the values given in the table for 20(C. The way of entering the buffer solution value was described in chapter 8.4.  In case of differences larger than +/_ 0.010 pH the user may not use the mode with use of NIST sample solutions but the mode which enables setting the buffers value and before the calibration set, in the meter’s memory, the accurate value of the used sample solution according to pt. 8.3.

The range of the temperatures, taken into consideration during introducing the correction, is 0 to 60 oC and can’t be exceeded.

The way of action during calibration way choice and entering sample solutions values other than those in table 1 was described in chapters 8.4.

- 24 -

e. If the function of automatic change of the pH sample solution value is not used it is necessary to measure the temperature of sample  solutions and bring it to the temperature given by the producer to receive the pH value identical like this entered to the meter’s memory. There is no limitation in order of using the samples.

Under each of the electrode numbers it is possible to calibrate different electrode in sample solutions with different values (after entering those values to the meters memory).

If the pH sample solutions with values according to NIST are used the values should not be changed.

The meter takes under consideration only the values of sample solutions detected during calibration.

8.5.1. Calibration with automatic temperature compensation

After preparing the meter to the calibration it is necessary in pH function to:

- press and hold the  button till the appearance of the CAL (pic. 7) symbol on the display; the old parameters of calibration are deleted;

- put the electrode and the temperature probe to the buffer solution, a P symbol with number of calibration point will be displayed, wait till stabilisation of the result. The result may be different than pH value of the buffer solution (pic. 8).

Pic. 8.

- 21 -

There is possibility to introduce the values of buffer solutions with two or three decimal places, depending on the chosen resolution.

Table 3.

| Calibration point | Range |
| --- | --- |
| 1 | 0,800 ( 2,100 |
| 2 | 3,900 ( 4,100 |
| 3 | 6,800 ( 7,100 |
| 4 | 8,900 ( 10,200 |
| 5 | 11,800( 14,000 |

The values of pH buffer solutions entered by the user are stored in the meter’s memory till their changing to others. Introducing of the sample solutions values should be repeated in case of using buffers with values different than those set in the meter’s memory.

The range of introducing the pH values of the buffer solutions in the individual points of calibration is quite wide what enables to use buffer solutions with values which differ from those set by the producer. For example buffers with value 2,00 pH; 7,00 pH; 9,00 pH; and 12,00 pH may be used. In every case the introduced buffer solution will be automatically recognised by the meter.

Together with the temperature change value of the pH buffer solution changes. The producers often give the values of the solutions in few temperatures. This data may be used and the meter may be calibrated in other temperature than 20 oC, by introducing to the meters memory value of the buffer which responds to this temperature

8.4.  Entering the calibration parameters

It is necessary to:

a.	Choose or block the function of automatic correction of the change of pH value of buffer solution with the temperature. (as described on p. 18)

To do so one should press the  button than a  symbol (points of calibration) will be displayed, (pic. 6) next using  or  buttons choose:

– automatic change of the stored pH value of the buffer solution by it’s temperature change according to table 3.

– setting the value of the buffer solution in the range given in table 2.

- 22 -

Pic. 6.

b.	After choosing the  or the    symbol one should start checking and eventually changing the stored values of calibration points. To do so one should press the  button. In the lower row a   symbol will be displayed, point one of calibration, and in the upper row value of the buffer. (pic. 7) If one uses buffer solution with value a bit different than this displayed one, using  or  buttons bring the displayed value to the value of used buffer solution.

Range of changes for each of calibration points in the  mode is of +/_ 0.010 from the values given in the table 1 for 20(C, and the range of changes in the  mode is given in the table 3.

Pic. 7.

- 23 -

To pass to the second point of calibration press the  button in the lower row a , symbol will be displayed (point 2) and in the upper row the stored in this point value of buffer solution. When checking or changing the value in the next steps please follow the instructions given above.

After the end of introducing the pH values of buffer solutions using the  button return to the choosing the calibration points mode ( / ,) or by pressing the  enter the pH measuring mode and calibration, chapter 8.4.

During next calibrations if the values of the previously used buffer solutions haven’t changed the actions described in this point may be omitted.

8.5.  pH electrode calibration

It is necessary to:

a.	choose the resolution with which the pH sample solution value will be introduced, point 6.2.

b.	choose the electrode number under which the results of calibration will be stored, point 6.3.

c.	prepare the electrode to work according to the producer’s instructions and mark it with number responding to number chosen in the meter (, , ).

d.	connect with the meter the prepared combination pH electrode and temperature probe, use the right connectors pH/mV and temp;

- 28 -

10.2. Measurements with manual temperature compensation

Disconnecting the temperature probe from the meter switches the meter to the manual temperature compensation mode (symbol  is displayed).

Measurement with manual temperature compensation is similar to the measurement with ATC, the difference is that using ,  buttons one enters the temperature measured with other thermometer. This value is displayed under the pH value and is taken into consideration during the compensation.

The manual compensation may be used during stable conditions ex. during pH measurements in the laboratory, especially when a thermostat is used, or when the temperature probe was damaged.

During the measurement with manual temperature compensation one should:

- insert the pH electrode to the vessel with the measured solution, if the electrode wasn’t calibrated or already was used for a long time the calibration should be done (chapter 8). During measurements in a vessel one shouldn’t touch the bottom and the walls with the electrode. The best way is to use an electrode stand;

- using a laboratory thermometer measure the temperature;

- turn the meter on using the  button;

- using the  button choose the pH function;

- with    buttons enter the value of measured temperature;

- wait till the stabilisation of the value and read the result.

Caution: simultaneous pressing of  and   buttons sets the temperature value to 20 0C.

- 25 -

After stabilisation of the result press the .button.

In this moment the result will blink what will inform about the storing the calibration result, in this same time in the upper row the corrected value of measurement will be displayed – it will be equal the value of used buffer solution (pic. 9).

If the function of automatic correction of the temperature influence on pH value of the buffer solution is used, this value will be considered in the displayed value of the buffer solution. If by accident someone will use buffer solution different than those set for the calibration the meter won’t be able to detect this value and  sign will be displayed. In this case it is necessary to check the value of the buffer solution or the electrode which may be broken or clogged.

Pic. 9.

- one may finish the calibration at this moment by pressing the  button or continue the calibration in other buffer solutions, washing the electrode and temperature probe before every immersing in the buffer and acting according to the point b.

After calibrating one electrode it is possible to calibrate two others, choosing the two remaining symbols according to point 6.3. page 13.

If after choosing the electrode number and entering the calibration mode one will exit it without making the calibration the earlier stored characteristic will be deleted and a standard characteristic will be adapted

- 26 -

8.5.2 Calibration with manual temperature compensation

To start the calibration with manual temperature compensation one has to disconnect the temperature probe. Disconnecting the temperature probe switches the meter to manual compensation. On the display the entered value of the temperature is displayed, not the measured one. The  ,  buttons will be unblocked and using them one may enter the value of the buffer solution temperature. This value is displayed in the lower row of numbers on the display.

Next thing is to join the pH electrode with the meter and act according to points a ( c. of the previous section. It is important to remember that in case of manual temperature compensation the temperature value entered by the user is taken under consideration.

Caution:	simultaneous pressing of both  and  buttons sets the temperature to 20 0C.

9. CHECKING THE ELECTRODE CONDITION

If after the calibration the electrode symbol is blinking (E1, E2, E3) it informs that the electrode has lost its efficiency and in short time its calibration won’t be possible. After entering the mode of electrode number changing under the number of electrode a  symbol will be displayed (pic. 10).

It is necessary to prepare a new electrode.

Pic. 10.

- 27 -

- pH MEASUREMENT

Before starting the measurement one has to prepare the meter for work (chapter 6) and pH electrode (chapter 7). Good condition of the electrode is the main condition of correct measurements. If the electrode was calibrated and is joined with the meter it is advisable to check whether the number of the electrode is this same as the number of the characteristic chosen from the memory. If not, the number should be changed according to section 6.3. and the resolution of measurement should be chosen according to the chapter 6.2.

  - Measurement with automatic temperature compensation

During measurements with automatic temperature compensation, the meter co-operates with the temperature probe and measures the temperature of the solution simultaneously with the pH measurement and takes it into consideration during compensation.

In case of measurement with automatic temperature compensation one should:

- join the temperature probe and the combination pH electrode to the right connectors on the meter pH/mV and temp, symbol  will be displayed;

- if the electrode wasn’t calibrated or was already used for some time it is advisable to calibrate it. (chapter 8);

- insert the electrode and the temperature probe to the measured solution. During measurements in vessels one should not touch the bottom and the walls with the electrode. Best way is to use an electrode stand;

- turn the meter on by pressing the  button;

- using  button choose the pH measurement function.

- after the stabilisation read the result

Accurate laboratory measurements require using of electromagnetic stirrer.

NOTICE: exceeding of the measuring range is indicated by blinking of the digits on the display. Exceeding the automatic temperature compensation range is signalised by blinking   symbol.

- 32 –

12. ORP (VOLTAGE) MEASUREMENT

The pH meters CP-501, CP-502 are accurate mV meters (Oxidation Reduction Potential). The measurement may be done with special redox electrode or during titration. The result may be checked after choosing the mV mode with  button and setting the required resolution point 6.2.

Pic. 11.

Short pressing of the  button causes taking the actually measured value as zero and entering the relative voltage measurement (in left upper corner of the LCD a REL symbol is displayed). Next pressing of the  button returns the meter to standard mV measurement (REL symbol disappears).

13. TEMPERATURE MEASUREMENT

The temperature measurement is done as follows:

-	to the Chinch connector connect the temperature probe;

-	by pressing the  button switch the meter on;

-	put the temperature probe to the measured solution;

-	wait till the stabilisation of the value and read  the lower row of numbers.

The meter co-operates with the PT-1000 probe. Depending on it’s class the accuracy of the measurement changes.

NOTICE: break in the circuit of the temperature probe switches the meter to the manual temperature compensation mode. It is signalised by changing the  symbol to  symbol. On the display in place of the measured temperature value the value of the temperature entered by the user is displayed.

Blinking -50°C value during measurement in positive temperatures informs about short circuit in the temperature probe..

- 29 -

11. NOTICES ABOUT THE TEMPERATURE COMPENSATION AND INTERPRETATION OF THE pH RESULTS

The pH meters CP-501 and CP-502 have a manual and automatic temperature compensation, which enables eliminating errors resulting from changes of the electrode characteristics under the influence of temperature changes. To explain the role of the temperature compensation it is important to remind that pH meter is an mV meter which displays voltage counted to pH unit. In constant temperature for one pH unit a constant mV value is changed. In temperature of 20 0C it is 58,168 mV. The value of mV on one pH unit changes together with the temperature, what is taken into consideration in the formula for “k coefficient” of the pH electrode.

k=0.198422 T

Taking this change into consideration during measurement is called temperature compensation, which is connected with change of the electrode efficiency and not with the change of the measured solution caused by temperature change.

Changes of the pH values of some solutions with the temperature are very small and others big ex. pure water.

When comparing solutions, which change their values together with the temperature change, it should be done in this same temperature.

Sometimes it happens that results of measurement in this same solution in stable temperature are different. When analysing such a situation it is important to take following factors into consideration:

-	differences may occur because of  poor quality electrode;

-	The result was treated as stabilised to fast (medium class electrode needs about 40 seconds to full stabilisation);

-	the measured solution may not be homogeneous and lack of magnetic stirrer doesn’t allow for obtaining similar results;

- During measurements in sewage some chemical reactions, which change the result, may occur;

Very small differences may be caused by the meters accuracy. The CP-501 and CP-502 have accuracy  0.002 pH, 1 digit, what practically means that in extreme situation results of 2 measurements may differ for 0.005 pH and this will be an acceptable error, because 1 measurement will be made with 0.002 pH error and second with -0.002 pH. 1 digit information explains difference caused by rounding up of the result on the last visible place on LCD (discretisation error).

If after making a two point calibration in pH buffers 7.00 pH and 4.00 pH (acidic conditions), the results are checked in 9.00 pH (alkali conditions),

- 30 -

in some cases the result may be 8.90 pH or 9.10 pH. This may occur when the electrode has unsymmetrical characteristic. Making a 3 point calibration with alkali, neutral and acidic buffers may prevent from such errors. Sometimes the measurement results are unstable, the quality of the electrode has the crucial influence. In most cases the reason for slow drifting of the result, it’s unstability or prolonging the time of stabilisation is the clogged junction, broken electrode or contaminated membrane.

Often it happens when wrong kind of electrode was chosen for the kind of measured solution.

Leaving the electrode for few hours in the distilled water or placing it in water with detergent may eliminate this symptoms, especially if the measurements were done in solutions with deposits, fats or oils. The electrode which for a long time wasn’t used may have the junction clogged by KCl crystals, what may be removed by placing the electrode in distilled water. Heavily contaminated electrode may be cleaned in chloroform and deposits of iron in 2N HCl. The electrode life may be prolonged by storing it in KCl solution. Depending on the kind of measured solution or substance proper kind of electrode should be chosen. They differ one from another with shape, membranes look, kind of junction and body. Using electrode not proper for the kind of measured solution may cause it’s damaging and make the measurement impossible.

- 31 -

III. ORP (mV) and temperature measurement

- 36 -

14.4.	Setting the time and date

The setting mode of the actually displayed parameter (time or date) one enters by pressing and holding the  button. The position which we are going to make the change starts blinking, the value is changed with  or  buttons. The pulsating position is chosen by short pressing of the  button.

The setting mode is left by pressing the  button.

The clock is powered by lithium battery which will last for about 10 years. When afterturning the meter on the clock digits will blink it informs that the time settings were lost and this requires setting the date and time. If the situation is repeated after each turning the meter on, it is necessary to send it to our service for replacing the battery.

- 33 -

V. Other

- 34 -

14. CLOCK WITH DATE

After choosing the time mode with  button the meter will display the actual time. By pressing the  button one may display the date, auto off time and battery condition.

# 14.1. Time display

The hour is displayed in two rows. In the upper one the hour and minutes are displayed and in the lower one seconds (Pic. 12). The way of changing the hours is described in the point 28.4.

Pic.12.

# 14.2. Date display

The date is displayed as follows: Month – Day – Year. (pic. 13) In the upper row the month and day are displayed and in the lower the year

Pic. 13.

- 35 -

14.3. Readout of the meter’s software version

In the  time function press the  button till displaying a screen, as on the picture 14 for the CP-501 or on the picture 15 for the CP-502. In the upper row the software version is displayed.:

Pic. 14.

In case of CP-502 additionally in the lower row a P.int sign is displayed what informs thatthe internal printer is installed.

Pic. 15.

Return to the time mode is done by pressing the  button.

- 40 -

15.5. Reviewing of the results

Reviewing of the stored results is started from the measuring mode, by pressing and holding the  button until the number of last stored result on turns with it’s value will be showed.

Every pressing of the  or  button shows the next or previous number and result with time and date if  function was chosen (p. 15.2.e.).

In this mode the buttons ,  work with repetition and after longer holding the numbers change very quickly till stopping on the highest or lowest number.

The reviewing mode may be left by pressing any of the function buttons.

15.6. Deleting the stored results

To delete the stored results one has to:

- press and hold the  button;

- with buttons  or  set the number of measurement from which we want do delete the memory;

-	press the  button what will delete the stored results from the chosen one till the end of the memory. On the display in the place of results there will be --- sign displayed what confirms the deleting.

-	The reviewing mode may be left by pressing any of the function buttons.

In case of using all numbers of memory more results won’t be stored. To store new results it is necessary to delete the old ones acting as it was described above.

If one wants to clear all the memory the deleting should be started from the first number.

- 37 -

15. STORAGE AND READOUT OF THE RESULTS

15.1. Storing or printing

The meter enables storing, readout or printing results of actually measured function. The results are stored in EEPROM memory, which is non-volatile, so the data isn’t lost even after complete lack of power. There is also possibility to printout on request the result of measurement which is being done at this moment. For printing an Interface EI-401 (Elmetron) is required. Before starting the work it is necessary to choose the storing or printing function and determine the way of displaying the result.

15.2.	Parameters of storing and reading from the memory

The parameters are changed in the readout mode. It may be entered from every measuring function by pressing and holding the  button, until on the display the number of last stored result will be displayed. This number is displayed on turns with the stored result.

Before starting the storing one has to choose the kind of results collecting: on request or as series, and also the way of displaying the result.

Next pressing of the  button shows screens with following functions which may be changed:

- - taking or printing series or single results (pic. 16)

Pic. 16.

- 38 -

With  or  button choose in lower row   or  symbols (pic. 16).

Choosing  activates automatic storing of the results and  single, after every pressing of . button

b.	 - time interval during series taking (pic.17).

Pic. 17.

Value of the time interval is displayed in the lower row of digits, and the informative symbol   in the upper row.

Buttons  and  are used to set the chosen time in minutes and seconds. The shortest time is 1 second and the longest 60 minutes. Holding the buttons make changing quicker. (repetition).

In case of choosing the  parameter to  setting the interval isn’t necessary.

c.  - printing - yes or no.

With   buttons one has to choose  when the results should be sent to printing output which is determined  parameters, or , when they should be stored in the memory.

- - output – printing direction. Using the   buttons one may choose in the lower row of the display the following symbol:

- output for the PC by RS-232

- output for the external printer by CENTRONICS plug

- output for the internal printer (only CPC-502 model)

- - Way of displaying the stored results.

- successively number of sample, result, time and date of storing the result.

- successively number of sample and result

Changing with    buttons.

Return to the results readout display after pressing  button.

Exit from the readout mode after pressing any of the function buttons.

- 39 -

15.3. Storing the single measurements in the memory

If according to previous section storing of single results was chosen and that the results will be collected in the memory and not printed or sent to a PC every pressing of  button stores the measured result. The results are stored as the next ones after the last stored. If someone was checking the earlier stored results and didn’t return to the last one the results won’t be deleted and the value will be stored after the last measurement. If the user wants to store the result from the chosen number first he has to delete the results (as described in the point 15.6) and next start storing the results by pressing the  button.

During storing the results the number of result will be displayed for a moment.

If after pressing the  button instead of number an  sign will be displayed it informs that the maximal number of results was stored.

15.4. Storing the measuring series

There is possibility to store series of measurements. To do so one has to :

-	choose the way of collecting the measurements (p. 15.2a);

- enter the time interval (p. 15.2b)

- delete the stored results starting from the chosen one (p. 15.6);

- with the proper button choose the function which results one wants to store

- with  button start taking the series. The measurements will be stored starting from the first free number.

Taking the series is signalised by blinking LED diode in the  button. Before each  storing of the result the number of measurement will be displayed for a moment. Taking of the series may be stopped by pressing  button, any of the function buttons or by filling the memory. If taking the series was stopped earlier than the memory was filled one may start to take the next series till the end of memory.

- 44 -

18. TECHNICAL DATA

# pH MEASUREMENT:

| range | resolution | Accuracy (±1digit) |
| --- | --- | --- |

| -2.000 16.000 pH | 0.001 / 0.01 pH | ±0.002 pH |
| --- | --- | --- |

INPUT IMPEDANCE:				1012

TEMPERATURE COMPENSATION:		manual/automatic

RANGE OF COMPENSATION:			-5.0 ÷ 110.0 oC

CALIBRATION OF THE pH ELECTRODE:		automatic, in 1 ( 5 points

RANGE OF RECOGNITION AND ENTERING OF THE pH BUFFER SOLUTIONS

| Calibration point | Range |
| --- | --- |
| 1 | 0,800 ( 2,100 |
| 2 | 3,900 ( 4,100 |
| 3 | 6,800 ( 7,100 |
| 4 | 8,900 ( 10,200 |
| 5 | 11,800( 14,000 |

AUTOMATIC CHANGE OF THE pH BUFFER VALUE TOGETHER WITH THE TEMPERATURE CHANGE, FOR SAMPLES CONSISTENT WITH NIST, TABLE (page 19) IN RANGE:			 0 ÷ 60 oC

THERMAL STABILITY OF ZERO:				0.0005 pH/ oC

# mV (ORP) MEASUREMENT:

| range | resolution | Accuracy (±1 digit) |
| --- | --- | --- |

| -1000 1000 mV | 0.1 mV | ±0.1 mV |
| --- | --- | --- |

INPUT IMPEDANCE:				1012

- 41 -

16. PRINTOUTS ON THE PRINTERS

In the right wall of the meter there is a Centronics output which enables joining the meter with printer using standard cable. In the CP-502 the results may be also printed on the internal printer.

There is possibility to print the measured value, or the results stored in the memory.

16.1. Choosing the direction of sending the data

Before sending the current results or those stored to the printer or PC it is necessary, in the readout mode press the  button till appearing in the upper row of LCD an  (output) symbol, and next with  ,  buttons choose in the lower row one of the symbols:

-	output for PC by RS-232,

-	output for external printer by Centronics;

-	output for internal printer (only CP-502).

16.2. Printout of the result – single or serial

To print the result of the current measurement one should:

-	connect the printer with the meter;

-	switch the meter and the printer on;

- choose the direction of sending the data (chapter 16.1);

- set the printout parameter  to  (chapter 15.2.c);

- choose the printout form (chapter 15.2.e) and way of printing – single or serial printouts (chapter 15.2.a);

-	if the single kind of taking measurements was chosen, every pressing of the  button will cause printing the actually measured value temperature and number;

- if taking the series of measurements was chosen, after pressing the  button printing of actually measured value, temperature and number of measurement with chosen time interval will be started.

If the parameter  was set to  together with the value of actually measured result and  temperature the date and hour will be printed.

Printing of the series may be stopped by pressing the  button or any of the function buttons.

There are no limits in number of printed results.

- 42 -

16.3. Printout of the results stored in the memory

To print the results stored in the memory:

-	join the printer with the meter using the cable (optionally);

- switch the meter and printer on;

- choose the direction of sending the data (chapter 16.1);

- choose the printout format (point 15.2.e) and way of printing – serial or single (point 15.2.a)

-	enter the results viewing mode (point 15.5), with buttons ,  set the number of measurement from which the printing should be started. Pressing the .starts printing.

The results of measurements with numbers will be printed.

In case of setting the  parameter to  all results from the set one to the end of memory will be printed, in case of setting the  parameter to  only the result stored under the  chosen number will be printed.

In case of setting the  parameter to , together with the result of measured function and temperature the hour and date will be printed.

16.4. Internal printer maintenance in CP-502

The printer maintenance basically is based on installing the paper. To do this one should:

- change the direction of sending the data to  (point 16.1);

- move up the lever on the right side of the meter;

- into the slot on the back of the meter insert an evenly cut paper simultaneously pressing the  button, till the end of paper will appear in the upper slot of the meter;

- move down the lever, the LED diode on the  button will light what confirms the readiness of printer to work;

- Through the roll of paper put the handle and the bonded part place in the slot on the back of the meter and press it down.

The paper will move out, by one line, after each pressing of the  button. Holding the  button causes continuous moving out of the paper. The excess of the paper may by tear off on the notched edge of the slot.

In case of not using the printer for a longer time or using standard PC printer it is necessary to move up the lever and leave it in this position. The LED on the  button will go out, what confirms switching the printer to stand by mode. End of the paper and improper lever position is signalised by blinking LED on the  button.

- 43 -

17. CO-OPERATION WITH THE PC

Connecting the meter with a PC enables storing data directly on the computer, what eliminates the limitation of the memory. The PC should be equipped with serial RS-232 connector (typically COM2) configured for 9600 b/s, 8 bit, 1 even bit, 1 stop bit, lack of steering the transmission.

On the picture 18 there is the way of joining the meter with the computer showed. On the right side of the picture there is the way of joining the cables with the 9 or 25 pin plug with relevant pin numbers.

For the transmission a special software prepared by our company may be ordered or residential Windows program - HyperTerminal (in menu Accessories / Communication). If it isn’t there it should be installed from the Windows installation CD. The program should be configured with option „direct connection with port....COMx”.

In the right wall of the meter there is a RS connector which enables joining the meter with a PC by a cable.

Plug                                                                     Plug

9 pin                                                                     9 pin    25 pin

Pic. 18.

After connecting one should switch the meter and PC on and enter the results readout mode by longer pressing of the  button, next press the  button till displaying in the upper row of LCD a  (output) sign and with   ,  buttons chose in the lower row of LCD a  symbol output for PC by  RS-232C,

After pressing the  button transmission of the current result and temperature will start. The mode and format of transmitted data depends on the set ,  and  parameters.

Caution:	the meter and PC should be switched on after connecting the cable to both of them.

- 45 -

# TEMPERATURE MEASUREMENT:

| Range | Resolution | Accuracy* (±1 digit) |
| --- | --- | --- |

| - 50.0 ÷ 199.9 oC | 0.1 oC | 0.1 oC |
| --- | --- | --- |

* accuracy of the meter. Final accuracy of the measurement depends on the accuracy

of the used PT-1000 probe

TEMPERATURE PROBE:	platinum resistor Pt-1000

ACCURACY OF THE PROBE IN RANGE:	0 ÷ 100 0C:

FOR PT1000B RESISTOR:		±0.8 0C

FOR PT10001/3B RESISTOR:		±0.27 0C

# INTERNAL PRINTER  (only CP-502):

# MECHANISM:	LTP1245M

PAPER WIDTH :	60 mm

RESOLUTION:	384 pts/line

PRINTING DENSITY:	23 symbols/line

MEMORY OF THE RESULTS

# Standard:		200 results

Optionally:		450  or  950 results

(when ordering the meter)

# OTHER:

# WORKING TEMPERATURE:	0 - 40 oC

POWER:	stabilised power adapter

CP-501 - 12V/100mA, CP-502 - 6V/2A.

POWER CONSUMPTION:	60 mW

SCREEN:	LCD 55 x 45 mm

DIMENSIONS:	200 x 180 x 50 mm

WEIGHT:					CP-501 – 650g, CP-502 - 720 g

- 46 -

19. EQUIPMENT

The standard equipment of the meter is:

1.	Temperature probe  Pt-1000B (standard);

2.	Combination pH glass electrode

3.	Power adapter: CP-501  12V/100mA, CP-502 6V/2A;

4.	User’s manual with warranty.

The additional equipment available for this meter is:

1.	Combined pH electrode for measurements in sewage or liquids with deposits;

2.	Ion selective electrode chosen from the price list;

3.	Reference electrode for ion selective measurements;

4.	5XX-PC cable;

5.	Software for collecting large number of data on the PC;

6.	Temperature probe Pt-1000 1/3Bwith higher accuracy;

- ORP measuring electrodes;

- Electrodes for ORP (redox) measurement.