(  Sp. j.

41-814  Zabrze  ul. Witosa 10  POLAND

tel.  +48 32 / 2738106   fax  +48 32 / 2738114

www.elmetron.com.pl

e-mail: info@elmetron.com.pl

LABORATORY

pH / CONDUCTIVITY / SALINITY

METER

CPC- 501

CPC-502

USER’S MANUAL

USER’S MANUAL

LABORATORY

pH / CONDUCTIVITY / SALINITY METER

CPC-501, CPC-502

Before use please read the instruction carefully!

### W A R R A N T Y

The “ELMETRON” company ensures a 24 months warranty for the pH / conductivity / salinity  meter CPC-501, CPC-502* number:

..................................................

In case of damage the producer will repair the meter within 14 days from the day of delivery.

The warranty doesn’t cover the damages caused by usage not in conformity with the users manual, using wrong power adapter, mechanical damages and damages caused by repairs made by unauthorised persons.

The pH electrode, conductivity cell, oxygen sensor have one year warranty of the producer.

NOTICE: Before sending the meter to us please contact the firm by phone or email.

When sending the meter, the used electrode, conductivity cell, oxygen sensor, temperature probe and power adapter and warranty with date of purchase should be also included.

Date of production.........................................

Date of sale...................................................

Date of warranty expiry..................................

* cross out the unnecessary

- 4 -

1. EXPLOITATION NOTICES

Dear User!

We present you a device distinguished by accuracy according to the technical data and by a high stability of the displayed results. We believe that the measurements will not cause you any trouble and that the meter will operate without any inconvenience. Wide range of additional functions requires careful reading of the manual in other case some of the features may stay unused or using the meter may cause you a troubles.

The employing of good-quality electrodes cells and probes and their replacing after a suitable time ensures obtaining of high measuring parameters. We want to call your attention to the fact, that this equipment has a much shorter working life than the meter. Typical symptoms of an improper operation of the electrode are deterioration of final result stability, it’s flowing as well as a higher measuring error. Part of the users has problems, arising from employing electrodes not being preconditioned before the measurement or making the measurements without removing the shielding ring from the liquid junction or taking measurements with a plugged junction. To avoid this situation it is necessary to choose the proper kind of electrode for solutions which are going to be measured ex. special electrodes for the sewage, liquids with deposits, meat cheese etc. Therefore, if you observe improper operation of the device, please take control measurements with another electrode or check the used electrode with another pH-meter. Generally the deterioration of the meter work is caused by the electrode and not by the meter.

In case of conductivity measurements it is important to choose the cell with right constant K value for the measuring range. Wrong selection may cause larger error occur, similar situation occurs during measurements with automatic temperature compensation with incorrectly introduced ( coefficient.

Accuracy of the dissolved oxygen measurements depends on the sensor calibration and regular maintenance which consist in replacing the membranes, electrolyte and cleaning the electrodes. Lack of correct maintaining of the sensor after some time will make measurements impossible. Please turn your attention to the fact that stabile measurement is possible only with simulated or natural measured water flow.

The essential feature of our  products is their low failure frequency. However if your pH-meter will fail, our firm immediately performs its warranty repair.

We wish you a pleasant and trouble-free work with our meter.

TABLE OF CONTENT

I. Introduction

1.	Exploitation notices	  4

2.	The characteristic of the meter	  5

3.	What is the meter designed for	  6

4.	Outside view	  7

5.	Switching the meter on and off	10

6.	Preparation to work	11

6.1.	Choosing the kind of temperature compensation	11

6.2.	Changing the resolution of the measurement	12

6.3.	Changing the electrode number	13

##### II. pH measurement

7.	Preparation of the pH electrode	16

8.	Calibration of the pH electrode	17

8.1.	Calibration mode

8.2.	Calibration with use of NIST sample solutions	19

8.3.	Calibration with introduced values of buffer solutions	20

8.4.	Entering the calibration parameters	21

8.5.	pH electrode calibration 	23

9.	Checking the electrode’s condition	26

10.	pH measurement	27

10.1.	Measurement with automatic temperature compensation	27

10.2 	Measurement with manual temperature compensation	28

- Notices about the temperature compensation and

interpretation of the results	29

##### III. Conductivity and salinity measurement

- Basic information about the conductivity measurement	32

13.	Preparation to work 	33

13.1.	Choosing the unit	33

13.2.	Entering the WTDS coefficient	35

14.	Choice and maintenance of the conductivity cell 	36

14.1	The cell choice 	36

14.2	The conductivity cell maintenance 	37

15.	Calibration	38

15.1	Calibration without the sample solution	38

15.2	Calibration with use of sample solution 	39

16.	Entering the α coefficient value and reference temperature	41

  - Entering the α coefficient value	41

  - Entering the reference temperature value	42

17.	The simplified way of determining the α coefficient	43

18.	The conductivity measurement	45

18.1.	The measurement without temperature compensation	45

18.2.	The measurement with automatic  temperature compensation 	46

18.3.	The measurement with manual  temperature compensation 	47

- 2 -

19.	Salinity and TDS measurement	48

19.1.	Salinity measurement with conversion to NaCl or KCl content	49

19.2.	Determining the WTDS  coefficient	49

19.3.	The measurement of salinity with conversion to TDS	50

##### IV. ORP (mV) and temperature measurement

20.	ORP (voltage) measurement	64

21.	Temperature measurement	64

##### V. Other

22.	Clock with date, 	66

22.1.	Time display	66

22.2.	Date display	66

22.3.	Setting the time and date 	66

23.	Storage and readout of the results	67

23.1.	Storing or printing	67

23.2.	Parameters of storing and reading from the memory	67

23.3.	Storing the single measurements in the memory 	69

23.4.	Storing the measuring series	69

23.5.	Reviewing of the results	70

23.6.	Deleting the stored results	70

24.	Printouts on the printer	71

24.1.	Choosing the direction of sending the data	71

24.2.	Printout of the result – single or serial	71

24.3.	Printout of the results stored in the memory 	72

24.4.	Internal printer maintenance in CPC-502	72

25.	Cooperation with a PC 	73

26.	Technical data	74

27.	Equipment	77

- 3 -

##### I. Introduction

- 8 -

The keyboard has the following keys:

-	switches the meter on and off;

-	chooses the pH measuring function;

-	chooses the mV measuring function;

-	displays the time and date;

-	longer pressing of this button enters the calibration mode (CAL symbol displayed). Short pressing in this mode confirms the calibration result;

-	short pressing causes storing or printing of single result or measuring series, longer pressing enters the stored results readout mode.

-	chooses the entered parameter;

, 	-	buttons for entering the parameters.

-	enables pulling out and slipping in of the paper by the printer (only in CPC-502).

On the back wall of the meter there are inputs with below given symbols:

pH/mV	-	BNC-50 connector for connecting combined or measuring pH electrode and ORP or ion selective electrodes;

Gnd	-	connector for reference electrode;

temp	-	Chinch connector for temperature probe;

POWER	-	connector for power adapter.

On the right side of the meter there are RS-232 and Centronics outputs.

- 5 -

2. THE CHARACTERISTICS OF THE METER

The pH meters CPC-501 and CPC-502 belong to the newest generation of measuring devices which offer wide range of additional functions. Technical data, extra features and way of maintenance are identical in both models. The CPC-502 has additionally build in thermal printer which enables printing the actual and stored in the memory results.

The meters ensures high accuracy and repeatability of the readings. The newest generation electronic elements used in the meter made it’s memory independent to power supply. The internal clock is  powered from battery which is enough for many years. The meter is equipped with custom LCD display, which enables simultaneous observing of the measured function, temperature value and additional symbols which make working easier. Waterproof housing makes working in difficult conditions possible.

Main features of the CPC-501 and CPC-502 are:

- high accuracy and stability of the reading;

- measurement of pH, mV and temperature;

- automatic and manual temperature compensation;

- electrode calibration in 1 to 5 points;

- automatic recognition of pH buffers and standards;

- imposed values of standard solutions with possibility of changing their value;

- option of automatic introduction of temperature correction on the value of pH buffer solutions (NIST norm);

- information about the condition of the pH electrode;

- independent storing of three characteristics of pH electrodes;

- internal datalogger for 200 (optionally 450 or 950) measurements with time date and temperature;

- taking series of measurements with set time interval;

- RS-232 output

- possibility of printing the measurement results or values stored in the memory on standard printer ;

- real clock and date.

- 6 -

3. WHAT IS THE METER DESIGNED FOR

The pH meters CPC-501 and CPC-502 are  precise and easy to use meters designed for hydrogen ion concentration measurements in pH units and Oxidation Reduction Potential (mV).

The meters may be also used for accurate temperature measurement of solutions and air in °C.

The meters are being used in food, chemical, pharmaceutical, energetical industries, in water treatment stations, laboratories, agriculture, universities, scientific laboratories etc.

CPC-501 and CPC-502 are prepared to work with all types of combination pH electrodes equipped with BNC-50 connector. There is also possibility of using a separate measuring and reference electrode. The meters co-operate with Pt-1000 temperature probe with Chinch connector.

The meter may collect measurements taken as single or series of measurements with set time interval. Built in RS-232 interface enables connecting the meter wit a PC for sending the data. Centronics output enables connecting the meter with a standard printer. There is possibility of printing the data collected in the memory or current results of the measurement.

CPC-502 enables printouts with use of the built in thermal printer (60 mm width).

In case of necessity to collect series longer than the memory capacity results it is possible to use special PC software offered by our company.

- 7 -

4. . OUTSIDE VIEW

On the left side of the meter there is a LCD display (pic. 1) on which depending on the chosen function following symbols are displayed:

- result of the pH measurement in pH units;

- result of the mV measurement in mV;

- time and date.

Choosing the function is done by specific buttons and is signalised by lighting LED diode placed on the button.

Simultaneously with the result a measured temperature value is displayed in oC. Symbols of the units are displayed by the results.

Pic. 1.

Beside the temperature value a  symbol for automatic temperature compensation is displayed or  for manual. CAL symbol on the left side of the display informs that the meter is in calibration mode.

On the left side of the display a number of chosen electrode is showed (E1; E2; E3) which inform about the number of characteristic which will be taken into consideration during all calculations. On the right side of the display a battery symbol which informs about the battery condition is displayed. During calibration between the upper and lower row of numbers number of the calibration point is displayed ( P1, P2, P3, P4, P5).

In the MODE all parameters introduced by the user are displayed and also the value of atmospheric pressure measurement (in oxygen measurement function). The keyboard  (pic. 2) placed under the display is used for switching the meter on and off, choosing the measuring function, calibration, entering the parameters, printing and storing the results in the memory.

- 12 -

6.2 Changing the resolution of the measurements

The measurement results may be displayed with chosen resolution. To change it one should:

- in the measuring mode press the  button, a  (resolution) sign will be displayed. (Pic. 4)

- Using keys  or  one may choose:

- (low) resolution of the measurement;

- (high) resolution of the measurement.

Pic. 4

For the pH measurement:

- resolution of the measurement 0.01 pH;

- resolution of the measurement 0.001 pH.

Return to the measuring mode by pressing the chosen function button

- 9 -

Pic. 2.

- 10 -

5. SWITCHING THE METER ON AND OFF

The meter is switched on by pressing the  button. The meter tests the memory and display on which all symbols are displayed (pic.3).

Pic. 3.

If the test was successful, after about 1.5 s the meter switches it self automatically to the measuring mode, in which it was switched off. If an  sign will be displayed it means that the meter has lost the factory settings and requires the service repair. If after the 1,5 s all symbols will be continuously displayed it informs that the calibration parameters of electrodes or cells  were lost.

After pressing the  button the meter will take standard characteristics:

- shift = 0 pH, characteristic slope = 100% for pH electrode;

And will enter the measuring mode. It will be necessary to calibrate the conductivity cell and  pH electrodes.

The meter is switched off by pressing and longer holding of the  button

- 11 -

6. PREPARATION TO WORK

Before starting the work one should:

-	join the power adapter plug to the Power input;

-	to pH/mV input (BNC-50) join the ready to work combination pH electrode, ion selective electrode or redox electrode;

- in case of using the temperature probe it should be connected with the chinch temperature input - temp;

- In case of work with PC join the 5XX-PC cable with RS-232 input

- In case of printing results on the printer connect the printer cable with Centronics output.

- switch the meter on by pressing the  button.

6.1. Choosing the kind of temperature compensation

The meter switches it self to the automatic temperature compensation mode automatically after joining the temperature probe, after disconnecting it the meter enters the manual temperature compensation mode. In ATC mode near the displayed temperature a  symbol appears. Manual temperature compensation is indicated by  symbol near the value entered by the user, it’s value may be changed using  or  keys.

- 16 -

7. PREPARATION OF THE pH ELECTRODE

The electrode should be prepared to work according to the producer instructions. If the instructions weren’t given please follow the steps:

- new electrode should be put into distilled water or in saturated KCl solution for about 5 hours:

- Before starting the measurements protecting rings (if used in this kind of electrode) should be removed. The ring placed on the junction – lower part of the electrode should be moved up along the electrode’s body and the upper, which protects the KCl refilling hole, down along the body.

Removing the lower ring is essential, in other case the electrode won’t measure.

Upper ring should be removed during measurements of high temperature solutions or to protect the junction during measurements in solutions with deposits or oils. Sometimes instead of ring a cork is used;

- during measurements in laboratory it is advisable to use an electrode holder;

- after every measurement the electrode should be washed in distilled water;

- excess of liquid on the electrode should be removed by gentle touching the glass with a tissue paper;

- after work the electrode should be stored in one of the above given solutions. The protecting rings should be moved on the junction and upper hole;

- in case of long breaks between the measurements the electrode should be stored, after drying, in the packaging;

- after taking the electrode of the package the eventual deposit should be removed using water;

- before using the electrode it should be placed in distilled water for about 2 hours;

- if the construction of the electrode enables refilling the electrolyte, it should be controlled and refilled periodically by the upper hole in the electrode’s body (usually as the electrolyte a KCl solution is used).

CAUTION:	storing of the electrode in distilled water shortens it’s life time and may cause shifting of the zero point.

.- 13 -

6.3. Changing the electrode number

If in the meter’s memory more than one electrode characteristic was stored, it is possible to replace the electrodes without the need of calibration. This option is quite useful in case of using different kinds of electrodes ex. for sewage, and clear water etc. It is necessary to connect the earlier calibrated electrode, marked with the number which responds to the number which was stored in the memory and choose this number.

To do so in the pH or ion measuring mode:

Press the   button, than using  or  buttons choose the electrode number (symbols , , ) under this number the calibration results will be stored (pic. 5). Above the electrode number following signs will be displayed:

-	under this number there is no characteristic stored and producers values are provided.

-	under this number there are values of last calibration stored.

-	the last calibration showed that the electrode is loosing it’s efficiency and in short time it’s calibration won’t be possible. (only for electrode calibrated in pH function)

Pic. 5.

Pressing of any function button returns the meter to measuring mode.

- 14 –

- 15 -

### II. pH measurement

- 20 -

Table 3.

| Temp. 0C | Kind of buffer solution |
| --- | --- |
|  | 1 oxalate | 2 phthalate | 3 phosphate | 4 di-sodium tetraborate | 5 calcium hydroxide |
| 0 | 1.666 | 4.003 | 6.984 | 9.464 | 13.423 |
| 5 | 1.668 | 3.999 | 6.951 | 9.395 | 13.207 |
| 10 | 1.670 | 3.998 | 6.923 | 9.332 | 13.003 |
| 15 | 1.672 | 3.999 | 6.900 | 9.276 | 12.810 |
| 20 | 1.675 | 4.002 | 6.881 | 9.225 | 12.627 |
| 25 | 1.679 | 4.008 | 6.865 | 9.180 | 12.454 |
| 30 | 1.683 | 4.015 | 6.853 | 9.139 | 12.289 |
| 35 | 1.688 | 4.024 | 6.844 | 9.102 | 12.133 |
| 40 | 1.694 | 4.030 | 6.838 | 9.063 | 11.984 |
| 45 | 1.700 | 4.047 | 6.834 | 9.038 | 11.841 |
| 50 | 1.707 | 4.060 | 6.833 | 9.011 | 11.705 |
| 55 | 1.715 | 4.075 | 6.834 | 8.985 | 11.574 |
| 60 | 1.723 | 4.091 | 6.836 | 8.962 | 11.449 |

8.3. Introducing the pH values of buffer solutions.

Values of the pH buffer solutions set by the producer in the meter’s memory are given in the table 2. If the values of used buffer solutions are different than the given above the user should change the set values to the actually used.

Table 2.

| Point of Calibration | Resolution 0,001 | Resolution 0,01 |
| --- | --- | --- |
| 1 | 2,000 | 2,00 |
| 2 | 4,000 | 4,00 |
| 3 | 7,000 | 7,00 |
| 4 | 9,000 | 9,00 |
| 5 | 12,000 | 12,00 |

The range of possible changes is limited for each calibration point.

The meter doesn’t allow for introducing pH values in ranges other than those given in the table 3.

- 17 -

8. CALIBRATION

Before starting the measurement with new electrode, after long-lasting using, or before making measurements which require higher accuracy the electrode connected with the meter should be calibrated. Results of measurements done without calibration will have a great error. The calibration is done on the buffer or sample solutions. It depends on comparing pH value of the buffer solutions with the value displayed by the meter and next automatic introduction of correction which is taken into consideration during measurements. The calibration should be periodically repeated because during work the parameters of the electrode are changing what influences the accuracy. The frequency of this procedure depends on the demanded accuracy, number of the measurements carried out, conditions in which the electrode was used, temperature and value of the measured solutions. When the highest accuracy is required it is recommended to use sample solutions with certificates. In technical usage buffer solutions are used with lower accuracy but mostly with total values ex. 2.00 pH; 4.00 pH etc. In this manual to make it easier a buffer solution name was used for both buffer and sample solutions. For accurate measurements it is necessary to use fresh good quality buffers.

The first action taken before the calibration is entering the values of used pH buffer solutions to the meters memory. This should be done before the first calibration and repeated in case of changing the used buffers to other values.

During the calibration after immersing the pH electrode and temperature probe the meter automatically detects the value of the used buffer, if it was earlier stored in the meter’s memory.

There is possibility of calibration in minimum 1 buffer solution and maximum in 5. The more points of calibration is used the higher is the accuracy in the whole measuring range.

Calibration in one buffer solution can’t ensure high accuracy. If only one buffer solution is used it’s value should be close to the forecasted value of measured solution. If the required accuracy isn’t very high and the measurements will be done in the whole range the one point calibration should be done with buffer solution close to 7.00 pH. Thanks this the so called zero electrode shift will be eliminated. In other points a standard characteristic will be taken from the meter’s memory.

- 18 -

This characteristic corresponds with theoretical efficiency of the pH electrode.

If measurements are done both in acids and alkalis and the measurements are not done at the ends of measuring range it is enough to calibrate the electrodes in 3 buffer solutions with values in range given in table 2 – calibration points 2, 3 and 4. In case of accurate measurements in the whole range it is recommended to calibrate the electrode in all 5 points given in the table. In CPC-501 and CPC-502 characteristic of the electrodes is approximated linearly between the calibration points.

Starting the calibration under one of the electrode numbers does not remove the set buffer values, but irreparably removes the electrode characteristic stored under this number.

There is no possibility of calibrating only one point and leaving the rest of data from the last calibration.

Memory in CPC-501 and CPC-502 enables independent storing of 3 characteristics of pH electrodes. This feature is especially useful when fast replacing of the electrodes is necessary or when it was broken. Those electrodes should be earlier calibrated and brought into memory under following symbols , , .

The order of using the buffer solutions is freely chosen.

8.1. Calibration mode

The user may choose two independent ways of action:

- use the values of pH sample solutions entered to the memory by the manufacturer. Those values are in conformity with the NIST norm. This type of calibration switches on the automatic correction introduction connected with the change of sample solution along with temperature changes.

- enter the value of pH buffer solutions depending on the actually used buffers, in the range given for each point of calibration;

Choosing the first option makes the calibration much easier. The  user doesn’t need to warm up or cool down the sample solutions to the temperature given by the user. This option includes some simplifications and when very high accuracy is required it shouldn’t be used. The meter stores independently 5 values of the buffers for each of the calibration modes described above.

- 19 -

8.2. Calibration with use of NIST sample solutions

In this mode 5 constant sample solutions values, according to NIST, are used. In the meters memory a table with dependence between the temperature and pH values for this 5 sample solutions is stored. This dependence is showed in the table 1. During the calibration the temperature of sample solution is measured and on it’s basis automatically the right pH value responding to this temperature is taken into consideration. Values between the points given in the table are approximated linearly by the meter.

The highest accuracy may be obtained only when the pH values of the buffer solutions are identical as in the table.

If there are some differences between the values given in the table and values given by the producer of buffer solutions, one should judge weather the error which will occur during calibration with automatic correction of the value won’t be to big. If yes one should resign from using this option.

Some sample solutions made according to the appropriate norms may differ on the third decimal place. In case of accurate measurements the user may not use the mode with use of NIST sample solutions but the mode which enables setting the buffers value and before the calibration set, in the meter’s memory, the accurate value of the used sample solution according to pt. 8.2.

The range of the temperatures, taken into consideration during introducing the correction, is 0 to 60 oC and can’t be exceeded.

The way of action during calibration way choice and entering sample solutions values other than those in table 1 was described in chapters 8.2. and 8.3.

- 24 -

e. If the function of automatic change of the pH sample solution value is not used it is necessary to measure the temperature of sample  solutions and bring it to the temperature given by the producer to receive the pH value identical like this entered to the meter’s memory. There is no limitation in order of using the samples.

Under each of the electrode numbers it is possible to calibrate different electrode in sample solutions with different values (after entering those values to the meters memory).

If the pH sample solutions with values according to NIST are used the values should not be changed.

The meter takes under consideration only the values of sample solutions detected during calibration.

8.5.1. Calibration with automatic temperature compensation

After preparing the meter to the calibration it is necessary in pH function to:

- press and hold the  button till the appearance of the CAL (pic. 7) symbol on the display; the old parameters of calibration are deleted;

- put the electrode and the temperature probe to the buffer solution, a P symbol with number of calibration point will be displayed, wait till stabilisation of the result. The result may be different than pH value of the buffer solution (pic. 8).

Pic. 8.

- 21 -

There is possibility to introduce the values of buffer solutions with two or three decimal places, depending on the chosen resolution.

Table 3.

| Calibration point | Range |
| --- | --- |
| 1 | 0,800 ( 2,100 |
| 2 | 3,900 ( 4,100 |
| 3 | 6,800 ( 7,100 |
| 4 | 8,900 ( 10,200 |
| 5 | 11,800( 14,000 |

The values of pH buffer solutions entered by the user are stored in the meter’s memory till their changing to others. Introducing of the sample solutions values should be repeated in case of using buffers with values different than those set in the meter’s memory.

The range of introducing the pH values of the buffer solutions in the individual points of calibration is quite wide what enables to use buffer solutions with values which differ from those set by the producer. For example buffers with value 2,00 pH; 7,00 pH; 9,00 pH; and 12,00 pH may be used. In every case the introduced buffer solution will be automatically recognised by the meter.

Together with the temperature change value of the pH buffer solution changes. The producers often give the values of the solutions in few temperatures. This data may be used and the meter may be calibrated in other temperature than 20 oC, by introducing to the meters memory value of the buffer which responds to this temperature

8.4.  Entering the calibration parameters

It is necessary to:

a.	Choose or block the function of automatic correction of the change of pH value of buffer solution with the temperature. (as described on p. 18)

To do so one should press the  button than a  symbol (points of calibration) will be displayed, (pic. 6) next using  or  buttons choose:

– automatic change of the stored pH value of the buffer solution by it’s temperature change according to table 3.

– setting the value of the buffer solution in the range given in table 2.

- 22 -

Pic. 6.

After choosing the  symbol one should press the  button and start the calibration in buffer solutions according to chapter 8.4.

After choosing the  act according to the point b.

b.	After choosing the  (user setting) symbol one should start checking and eventually changing the stored values of calibration points. To do so one should press the  button. In the lower row a   symbol will be displayed, point one of calibration, and in the upper row value of the buffer. (pic. 7) If one uses buffer solution with value a bit different than this displayed one, using  or  buttons bring the displayed value to the value of used buffer solution.

Range of changes for each calibration points is given in the table 3.

Pic. 7.

- 23 -

To pass to the second point of calibration press the  button in the lower row a , symbol will be displayed (point 2) and in the upper row the stored in this point value of buffer solution. When checking or changing the value in the next steps please follow the instructions given above.

After the end of introducing the pH values of buffer solutions using the  button return to the choosing the calibration points mode ( / ,) or by pressing the  enter the pH measuring mode and calibration, chapter 8.4.

During next calibrations if the values of the previously used buffer solutions haven’t changed the actions described in this point may be omitted.

8.5.  pH electrode calibration

It is necessary to:

a.	choose the resolution with which the pH sample solution value will be introduced, point 6.2.

b.	choose the electrode number under which the results of calibration will be stored, point 6.3.

c.	prepare the electrode to work according to the producer’s instructions and mark it with number responding to number chosen in the meter (, , ).

d.	connect with the meter the prepared combination pH electrode and temperature probe, use the right connectors pH/mV and temp;

- 28 -

10.2. Measurements with manual temperature compensation

Disconnecting the temperature probe from the meter switches the meter to the manual temperature compensation mode (symbol  is displayed).

Measurement with manual temperature compensation is similar to the measurement with ATC, the difference is that using ,  buttons one enters the temperature measured with other thermometer. This value is displayed under the pH value and is taken into consideration during the compensation.

The manual compensation may be used during stable conditions ex. during pH measurements in the laboratory, especially when a thermostat is used, or when the temperature probe was damaged.

During the measurement with manual temperature compensation one should:

- insert the pH electrode to the vessel with the measured solution, if the electrode wasn’t calibrated or already was used for a long time the calibration should be done (chapter 8). During measurements in a vessel one shouldn’t touch the bottom and the walls with the electrode. The best way is to use an electrode stand;

- using a laboratory thermometer measure the temperature;

- turn the meter on using the  button;

- using the  button choose the pH function;

- with    buttons enter the value of measured temperature;

- wait till the stabilisation of the value and read the result.

Caution: simultaneous pressing of  and   buttons sets the temperature value to 20 0C.

- 25 -

After stabilisation of the result press the .button.

In this moment the result will blink what will inform about the storing the calibration result, in this same time in the upper row the corrected value of measurement will be displayed – it will be equal the value of used buffer solution (pic. 9).

If the function of automatic correction of the temperature influence on pH value of the buffer solution is used, this value will be considered in the displayed value of the buffer solution. If by accident someone will use buffer solution different than those set for the calibration the meter won’t be able to detect this value and  sign will be displayed. In this case it is necessary to check the value of the buffer solution or the electrode which may be broken or clogged.

Pic. 9.

- one may finish the calibration at this moment by pressing the  button or continue the calibration in other buffer solutions, washing the electrode and temperature probe before every immersing in the buffer and acting according to the point b.

After calibrating one electrode it is possible to calibrate two others, choosing the two remaining symbols according to point 6.3. page 13.

If after choosing the electrode number and entering the calibration mode one will exit it without making the calibration the earlier stored characteristic will be deleted and a standard characteristic will be adapted

- 26 -

8.5.2 Calibration with manual temperature compensation

To start the calibration with manual temperature compensation one has to disconnect the temperature probe. Disconnecting the temperature probe switches the meter to manual compensation. On the display the entered value of the temperature is displayed, not the measured one. The  ,  buttons will be unblocked and using them one may enter the value of the buffer solution temperature. This value is displayed in the lower row of numbers on the display.

Next thing is to join the pH electrode with the meter and act according to points a ( c. of the previous section. It is important to remember that in case of manual temperature compensation the temperature value entered by the user is taken under consideration.

Caution:	simultaneous pressing of both  and  buttons sets the temperature to 20 0C.

9. CHECKING THE ELECTRODE CONDITION

If after the calibration the electrode symbol is blinking (E1, E2, E3) it informs that the electrode has lost its efficiency and in short time its calibration won’t be possible. After entering the mode of electrode number changing under the number of electrode a  symbol will be displayed (pic. 10).

It is necessary to prepare a new electrode.

Pic. 10

- 27 -

- pH MEASUREMENT

Before starting the measurement one has to prepare the meter for work (chapter 6) and pH electrode (chapter 7). Good condition of the electrode is the main condition of correct measurements. If the electrode was calibrated and is joined with the meter it is advisable to check whether the number of the electrode is this same as the number of the characteristic chosen from the memory. If not, the number should be changed according to section 6.3. and the resolution of measurement should be chosen according to the chapter 6.2.

  - Measurement with automatic temperature compensation

During measurements with automatic temperature compensation, the meter co-operates with the temperature probe and measures the temperature of the solution simultaneously with the pH measurement and takes it into consideration during compensation.

In case of measurement with automatic temperature compensation one should:

- join the temperature probe and the combination pH electrode to the right connectors on the meter pH/mV and temp, symbol  will be displayed;

- if the electrode wasn’t calibrated or was already used for some time it is advisable to calibrate it. (chapter 8);

- insert the electrode and the temperature probe to the measured solution. During measurements in vessels one should not touch the bottom and the walls with the electrode. Best way is to use an electrode stand;

- Turn the meter on by pressing the  button;

- using  button choose the pH measurement function.

- After the stabilisation read the result

Accurate laboratory measurements require using of electromagnetic stirrer.

NOTICE: exceeding of the measuring range is indicated by blinking of the digits on the display. Exceeding the automatic temperature compensation range is signalised by blinking   symbol.

- 32 –

# 12. BASIC INFORMATION ABOUT THE CONDUCTIVITY MEASUREMENT

The conductivity measurement is based on applying a voltage with proper frequency between two electrodes and measuring it. In the CPC-501 and CPC-502 the frequency depends on the measuring range and may vary from 100 Hz up to 10 kHz. Depending on the kind of the measured liquid, it’s concentration and the temperature the voltage flow is easier or harder. The conductivity result indirectly informs about the salts concentration in the measured liquid. The greater is the salt concentration the bigger is the conductivity (KCl, NaCl) This dependence doesn’t concern all of the solutions. In some cases after exceeding some salinity value the conductivity starts decreasing. The temperature has also a great influence on the results of a conductivity measurement which increases together with the temperature growth. The measured conductivity may be displayed as salinity in g/l of NaCl or KCl assuming that the measured liquid includes homogeneous salt.

The electrodes surface and the distance between them have decisive influence on the so called constant K value of the cell. This value has a great influence on the accuracy of the measurement. Depending on the measured conductivity value cells with constant K = 0,1 cm-.1 up to  10 cm-.1 are used. During the measurement the meter multiplies the measured value by the constant K introduced to the meters memory and displays the result in units of conductivity (μS/cm or mS/cm). By the result a shortened symbol of the unit is displayed (μS or mS). The conductivity is changing together with temperature and salts concentration. To make the comparing of the results possible the measured value is counted by the meter to the value which responds to measurement in the set reference temperature. Usually it is 25°C. Measurement in the reference temperature is most accurate. In other temperatures there is so called temperature compensation used, this means that there are some counting made by the meter which take the actual measured temperature and ( coefficient under the consideration. This coefficient describes how much will the result change when the temperature will change for 1 0C. The  coefficient in this meter may be changed in range 0 ÷ 10.00% / oC. The value of this coefficient is introduced to the memory by the user. For NaCl in temperatures close to 25 0C it is 2%/ 0C. Ex. in case of measurements in 30 0C the result changes 5 x 2%=10%. The counting is done automatically and takes the introduced by the user value of the  coefficient into consideration. The value of this coefficient may be found in the literature or may be approximated by the user.

The conductivity, despite the electronic precision, should always be treated as burdened with some error, which depends on the error of conductivity cell (it’s linearity), temperature and first of all depending on the ( coefficient during measurements in temperatures different than the set reference temperature.

- 29 -

11. NOTICES ABOUT THE TEMPERATURE COMPENSATION AND INTERPRETATION OF pH  MEASUREMENT RESULTS

The pH meters CPC-501 and CPC-502 have a manual and automatic temperature compensation, which enables eliminating errors resulting from changes of the electrode characteristics under the influence of temperature changes. To explain the role of the temperature compensation it is important to remind that pH meter is an mV meter which displays voltage counted to pH unit. In constant temperature for one pH unit a constant mV value is changed. In temperature of 20 0C it is 58,168 mV. The value of mV on one pH unit changes together with the temperature, what is taken into consideration in the formula for “k coefficient” of the pH electrode.

k=0.198422 T

Taking this change into consideration during measurement is called temperature compensation, which is connected with change of the electrode efficiency and not with the change of the measured solution caused by temperature change.

Changes of the pH values of some solutions with the temperature are very small and others big ex. pure water.

When comparing solutions, which change their values together with the temperature change, it should be done in this same temperature.

Sometimes it happens that results of measurement in this same solution in stable temperature are different. When analysing such a situation it is important to take following factors into consideration:

-	differences may occur because of  poor quality electrode;

-	The result was treated as stabilised to fast (medium class electrode needs about 40 seconds to full stabilisation);

-	the measured solution may not be homogeneous and lack of magnetic stirrer doesn’t allow for obtaining similar results;

- During measurements in sewage some chemical reactions, which change the result, may occur;

Very small differences may be caused by the meters accuracy. The CPC-501 and CPC-502 have accuracy  0.002 pH, 1 digit, what practically means that in extreme situation results of 2 measurements may differ for 0.005 pH and this will be an acceptable error, because 1 measurement will be made with 0.002 pH error and second with -0.002 pH. 1 digit information explains difference caused by rounding up of the result on the last visible place on LCD (discretisation error).

If after making a two point calibration in pH buffers 7.00 pH and 4.00 pH (acidic conditions), the results are checked in 9.00 pH (alkali conditions),

- 30 -

in some cases the result may be 8.90 pH or 9.10 pH. This may occur when the electrode has unsymmetrical characteristic. Making a 3 point calibration with alkali, neutral and acidic buffers may prevent from such errors. Sometimes the measurement results are unstable, the quality of the electrode has the crucial influence. In most cases the reason for slow drifting of the result, it’s unstability or prolonging the time of stabilisation is the clogged junction, broken electrode or contaminated membrane.

Often it happens when wrong kind of electrode was chosen for the kind of measured solution.

Leaving the electrode for few hours in the distilled water or placing it in water with detergent may eliminate this symptoms, especially if the measurements were done in solutions with deposits, fats or oils. The electrode which for a long time wasn’t used may have the junction clogged by KCl crystals, what may be removed by placing the electrode in distilled water. Heavily contaminated electrode may be cleaned in chloroform and deposits of iron in 2N HCl. The electrode life may be prolonged by storing it in KCl solution. Depending on the kind of measured solution or substance proper kind of electrode should be chosen. They differ one from another with shape, membranes look, kind of junction and body. Using electrode not proper for the kind of measured solution may cause it’s damaging and make the measurement impossible.

- 31 -

III. Conductivity and salinity measurement

- 36 -

14. CHOICE AND MAINTENANCE OF THE CONDUCTIVITY CELL

14.1. The cell choice

The conductivity measuring range in CPC-501 and CPC-502 meters is 0 – 1999 mS/cm. The meter co-operates with conductivity cells with constant K = 0.010 ÷ 19.99 cm-1 and BNC-50 connector. Depending on the required measuring range it is necessary to choose the right cell with constant K which enables receiving correct results. Beyond the range the cell looses it’s linearity and the results have greater error. Accurate measurements in the whole range are possible with use of 3 different conductivity cells.

Depending on the expected measuring range one may choose the correct cell using the chart below.

Pic. 16.	Dependence between the measuring range and constant K of the used conductivity cells.

# The cell with constant K ≈ 0,1 cm –1  should be used for measurements of ultra pure and redistilled water. Those waters after contact with air are very quickly changing their conductivity, so their pouring into vessels for measurement with dip cell (immersing) may cause errors. For accurate measurements it is necessary to use flow through probe with build in temperature probe, what enables measurement of water which flows directly from the container.

# - 33 -

# 13. ENTERING THE CONDUCTIVITY METER PARAMETERS

# Before the calibration and measurements it is necessary to make all activities described in the chapter 6. Additionally according to the chapter below it is necessary to choose the unit in which we will make the calibration and measurement.

# 13.1. Choosing the unit

The result of the measurement may be displayed in units of conductivity and salinity. The salinity may be counted to NaCl, KCl or TDS content.

The result may be displayed in % of weight concentration or in g/l. To choose the unit one has to:

-	in the conductivity measuring mode press the  button till displaying in the lower row on LCD a   (unit) symbol;

-	with buttons  ,  choose in the upper row on LCD sign:

-	measurement in units of conductivity (pic. 11);

Pic. 11.

- 34 -

-	measurement result counted to NaCl in g/l  (pic. 12);

Pic. 12.

-	measurement result counted to KCl in g/l (pic. 13);

# Pic. 13.

-	measurement result counted to TDS in g/l (pic. 14).

Pic. 14.

# - 35 -

In case of salinity measurement (,  or ), short pressing of the  button chooses displaying the result in % of weight concentration or g/l. By the symbol with name of the salt a % or g/l symbol will be displayed.

-	by pressing the . enter the measuring mode.

The result of measurement in % of weight concentration may be counted to value in ppm according to dependence:

1% of weight concentration (C) = 10 000 ppm = 10 ppt

Resolution of measurement in % of weight concentration is 0.001% or 10 ppm.

13.2. Entering the WTDS coefficient

In case of measurement of salinity with conversion to TDS it is necessary to introduce the WTDS coefficient.

One should follow the steps:

-	in the conductivity measuring mode few times press the  button till moment of displaying in the upper row on LCD a   symbol – TDS coefficient (pic. 15);

-	with  ,  buttons enter the correct value of the TDS coefficient;

-	enter the measuring mode by pressing the  button.

# Pic. 15.

The way of determining the WTDS coefficient is described in section 19.2.

- 40 -

-	wait till the stabilisation of the value and press the  button. Pulsating of the results informs about storing in the memory. If  symbol will be displayed it is necessary to check the introduced value of the sample solution.

-	exit the calibration mode by pressing the   button.

Pic. 19.

The meter is calibrated and ready for the measurement.

15.2.3. Calibration with manual temperature compensation

To calibrate the meter it is necessary to:

-	turn the meter on with the  button;

-	choose the measurement of conductivity with  button;

-	disconnect the temperature probe and simultaneously press the ,  buttons. In lower row of LCD the introduced reference temperature value will appear;

-	introduce the value of the sample solution (chapter 15.2.1);

-	immerse the conductivity cell in the sample solution and hold it at least 1cm away from the bottom and walls of the vessel. The measuring cell should be completely filled with the sample solution and shouldn’t include any air bubbles, the electrode’s surface should be evenly moistened*;

-	with lab thermometer measure the temperature of the sample solution and bring it to the introduced reference temperature value;

-	press and hold the  button till CAL symbol (pic. 19) will appear on LCD.

-	wait till stabilisation of the result and press the  button. Pulsating of the result informs about storing it. If  sign will be displayed it is necessary to check the entered value of sample solution.

-	exit the calibration mode by pressing the  button.

The meter is calibrated and ready to work.

* -	air bubbles may be removed by moving the immersed cell. To make the moistening of the electrodes easier it is recommended to immerse the cell in distilled water with washing-up liquid, and than wash it in distilled water.

- 37 -

14.2. The conductivity cell maintenance

To receive stable results it is recommended to store the cell for few hours before the measurement in water, especially this is required in case of distilled water measurements.

The conductivity probe maintenance mainly consists of accurate washing the inside of the measuring cell with distilled water. It is forbidden to clean the platinum electrodes mechanically, because the platinum layer may be rub of, what will result in decreasing the accuracy, may cause lowering the stability and change the constant K.

Measurements of liquids with oils and heavy sediments content may cause platinum contamination, make the measurement impossible and irreparably damage the electrodes. In case of fat content in the measured liquids it is possible to clean the electrodes by immersing the cell in acetone, chloroform, fourhydrofurane or detergent.

Some norms propose universal liquid for cleaning the cells. This is mixture of equal parts of isopropyl alcohol, ethyl ether and hydrocholic acid, diluted with water in 1:1 ratio.

The conductivity probe with broken measuring cell can’t be used for further measurements because the constant K is greatly changed, the result is unstable and the influence of placing the cell in measuring vessel greatly changes the result.

- 38 -

15. CALIBRATION

The calibration is done to prepare the meter for cooperation with the used conductivity probe, this is the essential condition for obtaining accurate results. The calibration may be done without the sample solution (by entering to the meter’s memory known value of the constant K of cell) or with use of standard solution with known conductivity. Constant K of the conductivity cell may change as reaction to contamination of the electrodes, that’s why periodically it is necessary to check the meters indications on the sample solution and eventually repeat the calibration. When using three cells it is possible to store their characteristics under the  , ,  symbols.

# 15.1. Calibration without the sample solution

The meter has the possibility of calibration without the use of sample solution. To make this calibration it is necessary to know the constant K of the conductivity probe. This value may be given by the cell producer or may be determined using the CPC-501 and CPC-502 meters after making the calibration in the sample solution.

To make the calibration one should:

-	in the measuring mode press the  button till displaying the screen with the value of constant K (pic. 17).

- with  ,  buttons (symbols , , ) choose the number of cell;

-	press the  button on the display a CAL symbol will be displayed;

-	with  ,  buttons enter the value of constant K;

-	press the  button and return to the mode of choosing the number of cell or enter the measuring mode by pressing the  button.

Pic. 17.

- 39 -

15.2. Calibration with use of sample solution

The meter enables one point calibration in the freely chosen sample solution. To decrease the error it is recommended to use solutions with value close to the forecasted value of measurement. It is required to use high quality sample solutions. The calibration is done in the actually chosen unit (subchapter 13.1).

# 15.2.1. Entering the value of sample solution

To enter the value of sample solution one should:

-	choose the unit according to point 13.1;

-	in the conductivity measuring mode press the  button till appearing a  symbol in the lower row on LCD (pic. 18);

-	with  ,  buttons enter, in the upper row on LCD, the value of sample solution;

-	enter the measuring mode by pressing the  button.

Pic. 18.

15.2.2. Calibration with automatic temperature compensation

What to do:

-	enter the value of sample solution  (point 15.2.1);

-	connect the conductivity cell and temperature probe;

-	immerse both probes in the sample solution, hold the conductivity cell at least 1 cm away from the bottom and walls of the vessel. The measuring cell should be completely filled with the measured solution, there shouldn’t be any air bubbles and the electrodes should be evenly moistened*;

-	measure the temperature of solution and bring it to 25 oC;

-	press and hold the  button till appearing a CAL symbol on the display (pic. 19).

* -	air bubbles may be removed by moving the immersed cell. To make the moistening of the electrodes easier it is recommended to immerse the cell in distilled water with washing-up liquid, and than wash it in distilled water.

- 44 -

# To determine the coefficient one should:

# 1.	The measured solution bring to the reference temperature TR and measure it’s conductivity (GTR).

# 2.	Change the solution temperature TX to value in which the measurement will be done.

3.	Turn the meter to manual temperature compensation by disconnecting the temperature probe from the meter.

# 4.	Enter with the keyboard the value of reference temperature TR .

# 5.	Again measure the conductivity of the solution.  This value will be different than in TR temperature (GTx).

# Determine the  coefficient using the  formula:

=

# where:	TR		reference temperature value in 0C

# Tx	-	value of the changed temperature in 0C

# GTR	-	conductivity measured in ref. Temperature TR.

# GTx	-	conductivity measured in temperature Tx

In this way the   coefficient was determined for the reference temperature TR and measurement temperature Tx.

During measurements in temperatures different than reference temperature it is necessary to introduce the calculated  coefficient value for the measurement temperature.

# Caution: In case when the TR reference temperature is equal to 25°C the above mentioned formula will be:

=

# where:	Tx	-	value of the changed temperature in 0C

# G25	-	conductivity measured in 25 0C.

# GTx	-	conductivity measured in temperature Tx

- 41 -

16. ENTERING THE  COEFFICIENT VALUE

AND THE REFERENCE TEMPERATURE

16.1 Entering the   coefficient value.

The  coefficient range in CPC-501 and CPC-502 is 0.00 ÷ 10.00 % / oC with accuracy of regulation 0.01 % / oC. For the measurements it is possible to adopt the most often used temperature coefficient  = 2 % / oC or in case of higher accuracy requirements determine the kind of measured solution and choose for it the correct value of this coefficient.

To introduce the  coefficient one should:

-	in the conductivity measuring mode press the  button till displaying the screen with value of the  coefficient (pic. 20). A  (temp. coefficient  ) symbol will be displayed;

-	with  ,  buttons enter the value of the coefficient .

-	return to the measuring mode by pressing the  button.

Pic. 20.

The measurement result will be counted with use of the introduced temperature coefficient .

- 42 -

16.2.	Introducing the reference temperature value

The range of introducing the reference temperature values is 10.0 ÷ 40.0 oC with possibility of setting with accuracy of 0.1 oC. The most often used value is 25 oC.

To introduce the reference temperature one should:

-	in the conductivity measuring mode press the  button till displaying a screen with value of the reference temperature (pic. 21). A   sign will be displayed (reference temperature);

-	with the  ,  buttons introduce the temperature value;

-	return to the measuring mode by pressing the  button.

Pic. 21.

The measurement result will be calculated to the introduced reference temperature.

Caution:	simultaneous pressing of  and  buttons sets the value to 25 0C.

- 43 -

17. THE SIMPLIFIED WAY OF DETERMINING THE  COEFFICIENT

The knowledge of  coefficient has a crucial significance during measurements in temperatures different than the reference temperature.

This coefficient is changing together with the temperature and concentration. Below we are giving values of  coefficient in 25 oC for few compounds with determined weight concentration.

# Table 4.

| substance | Weight Concentr. | coefficient |
| --- | --- | --- |

| HCl | 10 % | 1.56 |
| --- | --- | --- |

| KCl | 10 % | 1.88 |
| --- | --- | --- |

| H2SO4 | 50 % | 1.93 |
| --- | --- | --- |

| NaCl | 10% | 2.14 |
| --- | --- | --- |

| HF | 1.5 % | 7.20 |
| --- | --- | --- |

| HNO3 | 31 % | 1.39 |
| --- | --- | --- |

In the table 5 there are rough values of  coefficient for KCl and NaCl depending on the temperature and concentration of the measured liquid.

# Table 5.

| temp. `0C | coefficient |
| --- | --- |
|  | KCl solution | Saturated NaCl |
|  | 0,01M | 0,1M | 1,0M |  |
| 5 | 2,68 | 2,68 | 2,39 | 2,77 |
| 10 | 2,45 | 2,36 | 2,20 | 2,53 |
| 15 | 2,27 | 2,19 | 2,04 | 2,38 |
| 20 | 2,11 | 2,06 | 1,89 | 2,21 |
| 25 | 1,91 | 1,86 | 1,75 | 2,03 |
| 30 | 1,80 | 1,77 | - | 1,91 |

# In practice the  coefficient may be determined with simplification that on the value of this coefficient greater influence has the temperature change than changes in the composition of the salts and their concentration  (what shows the data in table 5).

- 48 -

19. SALINITY AND TDS MEASUREMENT

Salts and minerals dissolved in natural water influence the conductivity, which in principle is proportional to the quantity of dissolved substances. This dependence enables, after some calculations, to determine the salinity of the measured solution in concentration units (g/l or %), or determine the TDS (Total Dissolved Solids). The received values are always approximated and the final accuracy depends on the way of making the calculations, concentration of the measured solution and it’s temperature. In most salinity meters a simplification is used, that dependence between the conductivity and salinity in the solution is linear in the whole measuring range. Usually a 0.5 coefficient is used, the conductivity result in mS/cm is multiplied by this coefficient and the result of salinity in g/l is received, ex. if the conductivity value is 2 mS/cm the salinity is 1g/l. In reality the dependence between the conductivity and salinity isn’t linear the conversion coefficient is changing together with the concentration and temperature. The table 6 shows the dependence between conductivity and real salinity of NaCl solution in temperature 25 0C and values of salinity counted for constant coefficient 0.5. This comparison shows that for greater concentrations using the constant coefficient introduces large error.

Table 6.

| Conductivity (mS/cm) | Real salinity (g/l) | Salinity (g/l) Counted for coefficient = 0.5 | Error (%) by using the coefficient = 0.5 |
| --- | --- | --- | --- |
| 1.00 | 0.495 | 0.500 | 0.01 |
| 2.00 | 1.006 | 1.000 | 0.60 |
| 4.00 | 1.976 | 2.000 | 1.21 |
| 10.00 | 5.400 | 5.000 | 7.40 |
| 30.00 | 18.174 | 15.000 | 17.46 |

In CPC-501 and CPC-502 microcontroller take into consideration the real dependence between the conductivity and salinity what greatly lowers the error. There is possibility of counting the salinity in NaCl or KCl, because the dependence for this two salts is a bit different.

The results will be more accurate for homogeneous solutions (NaCl, KCl). Determining the concentration of salts mixture with unknown composition in most cases is counted to NaCl . To check the usefulness of water for home or industrial use usually determining of TDS is used.

- 45 -

18. THE CONDUCTIVITY MEASUREMENT

# 18.1. The conductivity measurement without temperature compensation

An accurate conductivity measurement should be done without the temperature compensation. The measured solution should be brought to the value of earlier introduced reference temperature. During the control it is possible to use the temperature probe. In case of work without temperature probe it is necessary to introduce the temperature value with  ,  buttons.

One should:

-	connect the conductivity and temperature probes to the right connectors cond and temp  (pic. 2);

-	turn the meter on with  button.

-	choose the conductivity measurement and the unit (point 13.1);

-	if the conductivity cell wasn’t calibrated earlier make the calibration according to chapter 15;

-	Place both probes in the measured solution, the conductivity cell can’t touch the walls and bottom. The measuring cell should be completely filled with the sample solution and shouldn’t include any air bubbles, the electrode’s surface should be evenly moistened*;

-	bring the temperature of the measured solution to 25 oC.

-	read the result after it’s stabilisation (pic. 22).

# Pic. 22.

* -	air bubbles may be removed by moving the immersed cell. To make the moistening of the electrodes easier it is recommended to immerse the cell in distilled water with washing-up liquid, and than wash it in distilled water.

- 46 -

18.2.	Conductivity measurement with automatic temperature compensation

# In case of measurement with automatic temperature compensation one should:

-	connect the conductivity and temperature probes to the right connectors cond and temp (pic. 2);

-	turn the meter on with  button.

-	choose the conductivity measurement with    button and the unit (pt. 13.1);

-	if the conductivity cell wasn’t calibrated earlier make the calibration according to chapter 15;

-	control or change the value of the temperature coefficient  and reference temperature;

-	Place both probes in the measured solution, the conductivity cell can’t touch the walls and bottom. The measuring cell should be completely filled with the sample solution and shouldn’t include any air bubbles, the electrode’s surface should be evenly moistened*;

-	wait till the stabilisation and read the result (pic. 23).

# Pic. 23.

Notice: in case of exceeding the value of temperature taken during compensation into consideration the result and  symbol start blinking although the conductivity measuring range wasn’t exceeded.

If near the value a  symbol is displayed instead of  it informs that the temperature probe is broken or inaccurately connected.

* -	air bubbles may be removed by moving the immersed cell. To make the moistening of the electrodes easier it is recommended to immerse the cell in distilled water with washing-up liquid, and than wash it in distilled water.

- 47 -

18.3.	Conductivity measurement with manual temperature compensation

The measurement with manual temperature compensation may be done during stable work conditions, ex. during measurements in laboratory, especially with use of thermostat, or in case of the temperature probe damage. Disconnecting the temperature probe switches the meter to manual temperature compensation.

In case of measurement with manual temperature compensation one should:

-	connect the conductivity probe to the right connector  cond (pic. 2);

-	disconnect the temperature probe from the meter;

-	turn the meter on with  button.

-	choose the conductivity measurement and the unit (point 13.1);

-	if the conductivity cell wasn’t calibrated earlier make the calibration according to chapter 15;

-	control or change the value of the temperature coefficient   and the reference temperature;

-	Place the probe in the measured solution, the conductivity cell can’t touch the walls and bottom. The measuring cell should be completely filled with the sample solution and shouldn’t include any air bubbles, the electrode’s surface should be evenly moistened*;

-	measure the temperature of the solution and introduce it’s value with the  ,  buttons.

-	after the stabilisation read the result (pic. 24).

Notice:	simultaneous pressing of  and  buttons sets the temperature to set reference temperature.

Pic. 24.

* -	air bubbles may be removed by moving the immersed cell. To make the moistening of the electrodes easier it is recommended to immerse the cell in distilled water with washing-up liquid, and than wash it in distilled water

- 52 -

20. ORP (VOLTAGE) MEASUREMENT

Multifunction meters CPC-501, CPC-502 are accurate mV meters (Oxidation Reduction Potential). The measurement may be done with special redox electrode or during titration. The result may be checked after choosing the mV mode with  button (pic. 25).

Pic. 25.

21. TEMPERATURE MEASUREMENT

The temperature measurement is done as follows:

-	to the Chinch connector connect the temperature probe;

-	by pressing the  button switch the meter on;

-	put the temperature probe to the measured solution;

-	wait till the stabilisation of the value and read  the lower row of numbers.

The meter co-operates with the PT-1000 probe. Depending on it’s class the accuracy of the measurement changes.

NOTICE: break in the circuit of the temperature probe switches the meter to the manual temperature compensation mode. It is signalised by changing the  symbol to  symbol. On the display the value of the temperature entered by the user is displayed.

Blinking -50°C value during measurement in positive temperatures informs about short circuit in the temperature probe.

- 49 -

To use the conductivity measurement for determining the TDS it is necessary to specify the WTDS coefficient, which introduced to the meter’s memory enables automatic conversion. To determine the WTDS coefficient it is necessary to specify the weight of dissolved substances. The laboratory method of determining the dissolved solids content consists of taking a given volume of the water, evaporating the filtered sample, drying it to constant weight in temperatures 103 ( 1050C, weighting and counting in volume ratio (mg/dm3). The received weight is lower than total dry mass in water because besides the dissolved solids there can be also substances which were not dissolved which are removed by filtering before evaporating. Using the conductivity measurement it is possible to determine the Total Dissolved Solids. It is assumed that the salt’s composition in the taken samples is not changing significantly.

19.1. 	Salinity measurement with conversion to NaCl or KCl content

The measurement of salinity with conversion to NaCl or KCl content should be done as follows:

-	choose the salinity measurement with conversion to NaCl or KCl content according to chapter 13.1;

-	choose the unit (g/l or %);

-	then act as during the conductivity measurement (chapter 18);

-	after stabilisation read the result.

19.2. 	Determining the WTDS coefficient

For salinity measurement with conversion to TDS content it is necessary to determine the WTDS coefficient and enter it to the meter’s memory. To do so it is necessary to make the conductivity measurement of the tested water with exactly given volume or weight. In traditional way determine the Total Dissolved Solids in this water and than count the coefficient according to the below given formulas.

- 50 -

1. If the result is to be displayed in  g/l:

WTDS  =

where:

WTDS  - TDS coefficient

TDS – Total Dissolved Solids in g/l;

( - conductivity of the sample in mS/cm;

Caution: the TDS value should be counted to volume of sample equal to 1l.

2.	If the result is to be displayed in % of weight concentration:

WTDS  =

where:

WTDS  - TDS coefficient

TDS – Total Dissolved Solids in g/kg;

( -conductivity of the sample in  mS/cm;

Caution: the TDS value should be counted to weight of the sample equal  1kg.

19.3. 	The measurement of salinity with conversion to TDS

The measurement of salinity with conversion to TDS should be done as follows:

- according to chapter 13.2 introduce the WTDS coefficient;

- choose the measurement of salinity with conversion to TDS and the displayed unit (g/l or %);

- than act as during conductivity measurement (chapter 18).

- after stabilisation read the result in chosen units (g/l or %)

- 51 -

IV. ORP and temperature measurement

- 56 -

22.4.	Setting the time and date

The setting mode of the actually displayed parameter (time or date) one enters by pressing and holding the  button. The position which we are going to make the change starts blinking, the value is changed with  or  buttons. The pulsating position is chosen by short pressing of the  button.

The setting mode is left by pressing the  button.

The clock is powered by lithium battery which will last for about 10 years. When afterturning the meter on the clock digits will blink it informs that the time settings were lost and this requires setting the date and time. If the situation is repeated after each turning the meter on, it is necessary to send it to our service for replacing the battery.

- 53 -

V. Other

- 54 -

22. CLOCK WITH DATE

After choosing the time mode with  button the meter will display the actual time. By pressing the  button one may display the date, auto off time and battery condition.

# 22.1. Time display

The hour is displayed in two rows. In the upper one the hour and minutes are displayed and in the lower one seconds (Pic. 26). The way of changing the hours is described in the point 22.4.

Pic.26.

# 22.2. Date display

The date is displayed as follows: Month – Day – Year. (pic. 27) In the upper row the month and day are displayed and in the lower the year

Pic. 27.

- 55 -

22.3. Readout of the meter’s software version

In the  time function press the  button till displaying a screen, as on the picture 28 for the CPC-501 or on the picture 29 for the CPC-502. In the upper row the software version is displayed.:

Pic. 28.

In case of CPC-502 additionally in the lower row a P.int sign is displayed what informs thatthe internal printer is installed.

Pic. 29.

Return to the time mode is done by pressing the  button.

- 60 -

23.5. Reviewing of the results

Reviewing of the stored results is started from the measuring mode, by pressing and holding the  button until the number of last stored result on turns with it’s value will be showed.

Every pressing of the  or  button shows the next or previous number and result with time and date if  function was chosen (p. 29.2.e.).

In this mode the buttons ,  work with repetition and after longer holding the numbers change very quickly till stopping on the highest or lowest number.

The reviewing mode may be left by pressing any of the function buttons.

23.6. Deleting the stored results

To delete the stored results one has to:

- press and hold the  button;

- with buttons  or  set the number of measurement from which we want do delete the memory;

-	press the  button what will delete the stored results from the chosen one till the end of the memory. On the display in the place of results there will be --- sign displayed what confirms the deleting.

-	The reviewing mode may be left by pressing any of the function buttons.

In case of using all numbers of memory more results won’t be stored. To store new results it is necessary to delete the old ones acting as it was described above.

If one wants to clear all the memory the deleting should be started from the first number.

- 57 -

23. STORAGE AND READOUT OF THE RESULTS

23.1. Storing or printing

The meter enables storing, readout or printing of actually measured function. The results are stored in EEPROM memory, which is non-volatile, so the data isn’t lost even after complete lack of power. There is also possibility to printout on request the result of measurement which is being done at this moment. For printing an Interface EI-401 (Elmetron) is required. Before starting the work it is necessary to choose the storing or printing function and determine the way of displaying the result.

23.2.	Parameters of storing and reading from the memory

The parameters are changed in the readout mode. It may be entered from every measuring function by pressing and holding the  button, until on the display the number of last stored result will be displayed. This number is displayed on turns with the stored result.

Before starting the storing one has to choose the kind of results collecting: on request or as series, and also the way of displaying the result.

Next pressing of the  button shows screens with following functions which may be changed:

- - taking or printing series or single results (pic. 30)

Pic. 30.

- 58 -

With  or  button choose in lower row   or  symbols (pic. 30).

Choosing  activates automatic storing of the results and  single, after every pressing of . button

b.	 - time interval during series taking (pic.31).

Pic. 31.

Value of the time interval is displayed in the lower row of digits, and the informative symbol   in the upper row.

Buttons  and  are used to set the chosen time in minutes and seconds. The shortest time is 1 second and the longest 60 minutes. Holding the buttons make changing quicker. (repetition).

In case of choosing the  parameter to  setting the interval isn’t necessary.

c.  - printing - yes or no.

With   buttons one has to choose  when the results should be printed, or , when they should be stored in the memory.

- - Way of displaying the stored results.

- successively number of sample, result, time and date of storing the result.

- successively number of sample and result

Changing with    buttons.

Return to the results readout display after pressing  button.

Exit from the readout mode after pressing any of the function buttons.

- 59 -

23.3. Storing the single measurements in the memory

If according to previous section storing of single results was chosen and that the results will be collected in the memory and not printed or sent to a PC every pressing of  button stores the measured result. The results are stored as the next ones after the last stored. If someone was checking the earlier stored results and didn’t return to the last one the results won’t be deleted and the value will be stored after the last measurement. If the user wants to store the result from the chosen number first he has to delete the results (as described in the point 23.6) and next start storing the results by pressing the  button.

During storing the results the number of result will be displayed for a moment.

If after pressing the  button instead of number an  sign will be displayed it informs that the maximal number of results was stored.

23.4. Storing the measuring series

There is possibility to store series of measurements. To do so one has to :

-	choose the way of collecting the measurements (p. 23.2a);

- enter the time interval (p. 23.2b)

- delete the stored results starting from the chosen one (p. 23.6);

- with the proper button choose the function which results one wants to store

- with  button start taking the series. The measurements will be stored starting from the first free number.

Taking the series is signalised by blinking LED diode in the  button. Before each  storing of the result the number of measurement will be displayed for a moment. Taking of the series may be stopped by pressing  button, any of the function buttons or by filling the memory. If taking the series was stopped earlier than the memory was filled one may start to take the next series till the end of memory.

- 64 -

26. TECHNICAL DATA

# pH MEASUREMENT:

| range | resolution | Accuracy (±1digit) |
| --- | --- | --- |

| -2.000 16.000 pH | 0.001 / 0.01 pH | ±0.002 pH |
| --- | --- | --- |

INPUT IMPEDANCE:				1012

TEMPERATURE COMPENSATION:		manual/automatic

RANGE OF COMPENSATION:			-5.0 ÷ 110.0 oC

CALIBRATION OF THE pH ELECTRODE:		automatic, in 1 ( 5 points

RANGE OF RECOGNITION AND ENTERING OF THE pH BUFFER SOLUTIONS

| Calibration points | Range |
| --- | --- |
| 1 | 0,800 ( 2,100 |
| 2 | 3,900 ( 4,100 |
| 3 | 6,800 ( 7,100 |
| 4 | 8,900 ( 9,400 |
| 5 | 11,500( 14,000 |

AUTOMATIC CHANGE OF THE pH BUFFER VALUE TOGETHER WITH THE TEMPERATURE CHANGE, FOR SAMPLES CONSISTENT WITH NIST, TABLE (page 19) IN RANGE:			 0 ÷ 60 oC

THERMAL STABILITY OF ZERO:				0.0005 pH/ oC

# mV (ORP) MEASUREMENT:

| range | resolution | Accuracy (±1 digit) |
| --- | --- | --- |

| -1000 1000 mV | 0.1 mV | ±0.1 mV |
| --- | --- | --- |

INPUT IMPEDANCE:				1012

- 61 -

24. PRINTOUTS ON THE PRINTERS

In the right wall of the meter there is a Centronics output which enables joining the meter with printer using standard cable. In the CPC-502 the results may be also printed on the internal printer.

There is possibility to print the measured value, or the results stored in the memory.

24.1. Choosing the direction of sending the data

Before sending the current results or those stored to the printer or PC it is necessary, in the readout mode (point 23.2) press the  button till appearing in the upper row of LCD an  (output) symbol, and next with  ,  buttons choose in the lower row one of the symbols:

-	output for PC by RS-232,

-	output for external printer by Centronics;

-	output for internal printer (only CPC-502).

24.2. Printout of the result – single or serial

To print the result of the current measurement one should:

-	connect the printer with the meter;

-	switch the meter and the printer on;

- choose the direction of sending the data (chapter 24.1);

- set the printout parameter  to  (chapter 23.2.c);

- choose the printout form (chapter 23.2.d) and way of printing – single or serial printouts (chapter 23.2.a);

-	if the single kind of taking measurements was chosen, every pressing of the  button will cause printing the actually measured value temperature and number;3

- if taking the series of measurements was chosen, after pressing the  button printing of actually measured value, temperature and number of measurement with chosen time interval will be started.

If the parameter  was set to  together with the value of actually measured result and  temperature the date and hour will be printed.

Printing of the series may be stopped by pressing the  button or any of the function buttons.

There are no limits in number of printed results.

- 62 -

24.3. Printout of the results stored in the memory

To print the results stored in the memory:

-	join the printer with the meter using the cable (optionally);

- switch the meter and printer on;

- choose the direction of sending the data (chapter 24.1);

- choose the printout format (point 23.2.d) and way of printing – serial or single (point 23.2.a)

-	enter the results viewing mode (point 23.5), with buttons ,  set the number of measurement from which the printing should be started. Pressing the .starts printing.

The results of measurements with numbers will be printed.

In case of setting the  parameter to  all results from the set one to the end of memory will be printed, in case of setting the  parameter to  only the result stored under the  chosen number will be printed.

In case of setting the  parameter to , together with the result of measured function and temperature the hour and date will be printed.

24.4. Internal printer maintenance in CPC-502

The printer maintenance basically is based on installing the paper. To do this one should:

- change the direction of sending the data to  (point 24.1);

- move up the lever on the right side of the meter;

- into the slot on the back of the meter insert an evenly cut paper simultaneously pressing the  button, till the end of paper will appear in the upper slot of the meter;

- move down the lever, the LED diode on the  button will light what confirms the readiness of printer to work;

- Through the roll of paper put the handle and the bent part place in the slot on the back of the meter and press it down.

The paper will move out, by one line, after each pressing of the  button. Holding the  button causes continuous moving out of the paper. The excess of the paper may by tear off on the notched edge of the slot.

In case of not using the printer for a longer time or using standard PC printer it is necessary to move up the lever and leave it in this position. The LED on the  button will go out, what confirms switching the printer to stand by mode. End of the paper and improper lever position is signalised by blinking LED on the  button.

- 63 -

25. CO-OPERATION WITH THE PC

Connecting the meter with a PC enables storing data directly on the computer, what eliminates the memory limitation. The PC should be equipped with serial RS-232 connector (typically COM2) configured for 9600 b/s, 8 bit, 1 even bit, 1 stop bit, lack of steering the transmission.

On the picture 32 there is the way of joining the meter with the computer showed. On the right side of the picture there is the way of joining the cables with the 9 or 25 pin plug with relevant pin numbers.

For the transmission a special software prepared by our company may be ordered or residential Windows program - HyperTerminal (in menu Accessories / Communication). If it isn’t there it should be installed from the Windows installation CD. The program should be configured with option „direct connection with port....COMx”.

In the right wall of the meter there is a RS connector which enables joining the meter with a PC by a cable.

Plug                                                                     Plug

9 pin                                                                     9 pin    25 pin

Pic. 32.

After connecting one should switch the meter and PC on and enter the results readout mode by longer pressing of the  button, next press the  button till displaying in the upper row of LCD a  (output) sign and with   ,  buttons chose in the lower row of LCD a  symbol output for PC by  RS-232C,

After pressing the  button transmission of the current result and temperature will start. The mode and format of transmitted data depends on the set ,  and  parameters.

The picture 26 shows way of connecting the meter with a PC. On the right side of picture the connections from the PC side are showed for 9 or 25 pin connector.

Caution:	the meter and PC should be switched on after connecting the cable to both of them.

- 65 -

# CONDUCTIVITY MEASUREMENT:

| ranges | resolution | Accuracy* (±1 digit) | Frequency |
| --- | --- | --- | --- |

| 0.000 ÷ 19.999 S/cm | 0.001 / 0.01 S/cm | ±0.1 % | 100 Hz |
| --- | --- | --- | --- |

| 20.00 199.99 S/cm | 0.01 / 0.1 S/cm | ±0.1 % | 1 kHz |
| --- | --- | --- | --- |

| 200.0 1999.9 S/cm | 0.1 / 1 S/cm | ±0.1 % | 2 kHz |
| --- | --- | --- | --- |

| 2.000 19.999 mS/cm | 0.001 / 0.01 mS/cm | ±0.1 % | 5 kHz |
| --- | --- | --- | --- |

| 20.00 199.99 mS/cm | 0.01 / 0.1 mS/cm | ±0.25 % | 10 kHz |
| --- | --- | --- | --- |
| 200.0 1999.9 mS/cm | 0.1 / 1 mS/cm | ±0.25 % | 10 kHz |

*	Accuracy given for the end value of the range.

Ranges of frequency changes were given for constant  K = 1. For other values of the constant K the values will change proportionally to changes of this constant.

TEMPERATURE COMPENSATION:	manual/automatic

COMPENSATION RANGE:	-5.0 ÷ 70.0 oC

CONSTANT K RANGE:	0.010  19.999 cm-1

α COEFFICIENT RANGE:	0.00  10.00 %/ oC

TDS COEFFICIENT RANGE:	0.20  1.00

MEASURING RANGE KCl:	0  200 g/l

MEASURING RANGE NaCl:	0  250 g/l

PROBE CALIBRATION:	one point

1. by entering the constant K of the probe

2. using the calibration solution

# TEMPERATURE MEASUREMENT:

| Range | Resolution | Accuracy* (±1 digit) |
| --- | --- | --- |

| - 50.0 ÷ 199.9 oC | 0.1 oC | 0.1 oC |
| --- | --- | --- |

*	accuracy of the meter. Final accuracy of the measurement depends on the accuracy of the used PT-1000 probe

TEMPERATURE PROBE:	platinum resistor Pt-1000

ACCURACY OF THE PROBE IN RANGE:	0 ÷ 100 0C:

FOR PT1000B RESISTOR:		±0.8 0C

FOR PT10001/3B RESISTOR:		±0.27 0C

- 66 -

MEMORY OF THE RESULTS

# Standard:	200 results

Optionally:	450  or  950 results

(when ordering the meter)

# INTERNAL PRINTER  (only CPC-502):

# MECHANISM:	LTP1245M

PAPER WIDTH :	60 mm

RESOLUTION:	384 pts/line

PRINTING DENSITY:	23 symbols/line

# OTHER:

# WORKING TEMPERATURE:	0 - 40 oC

POWER:	stabilised power adapter

CPC-501 - 12V/100mA, CPC-502 - 6V/2A.

POWER CONSUMPTION:	60 mW

SCREEN:	LCD 55 x 45 mm

DIMENSIONS:	200 x 180 x 50 mm

WEIGHT:	CPC-501 – 650g, CPC-502 - 720 g

- 67 -

27. EQUIPMENT

The standard equipment of the meter is:

- Standard combination glass pH electrode

- Conductivity cell (K=0,5cm-1)

3.	Temperature probe  Pt-1000B (standard);

4.	Power adapter: CPC-501  12V/100mA, CPC-502 6V/2A;

5.	User’s manual with warranty.

The additional equipment available for this meter is:

1.	Other conductivity cells, different kinds for different conductivity ranges;

3.	5XX-PC cable;

4.	Software for collecting large number of data on the PC;

5.	Temperature probe Pt-1000 1/3Bwith higher accuracy;

6.	Ion selective electrodes;

- ORP measuring electrodes;