##### Model SX716 DO Meter

Operation Manual

Shanghai San-Xin Instrumentation Inc.

SX700 Series of Portable Meters

1. Model SX711 pH / mV Meter

2. Model SX712 ORP Meter

3. Model SX713 Conductivity / TDS / Salinity / Resistivity Meter

4. Model SX716 DO Meter

5. Model SX721 pH / ORP Meter

6. Model SX723 pH /mV / Conductivity Meter

7. Model SX725 pH / mV / DO Meter

8. Model SX726 Conductivity / DO Meter

9. Model SX731 pH / ORP / Conductivity Meter

10. Model SX736 pH / mV / Conductivity / DO Meter

11. Model SX751 pH / ORP / Conductivity / DO Meter

Table of Contents

1. Brief Introduction -----------------------------------------------------------------------------------

2. Technical Parameters -------------------------------------------------------------------------------

3. Instructions to the Meter ---------------------------------------------------------------------------

3.1. LCD Display -------------------------------------------------------------------------------------

3.2. Operation Keys ----------------------------------------------------------------------------------

3.3. The Storage, Recall and Elimination of the Measuring Information ---------------------

4. DO Measurement -----------------------------------------------------------------------------------

4.1. Preparation Work --------------------------------------------------------------------------------

4.2. Meter Calibration -------------------------------------------------------------------------------

4.3. Water Sample Test -------------------------------------------------------------------------------

4.4. Considerations -----------------------------------------------------------------------------------

4.5. Parameter Setting --------------------------------------------------------------------------------

4.6. Change the Membrane Cap --------------------------------------------------------------------

4.7. Zero Oxygen Calibration -----------------------------------------------------------------------

4.8. Salinity Calibration -----------------------------------------------------------------------------

4.9 Barometric pressure Calibration ---------------------------------------------------------------

5. Meter’s Complete Kit ------------------------------------------------------------------------------

6. Warranty ----------------------------------------------------------------------------------------------

ChartⅠ Meter Parameter Setting Schedule

ChartⅡ Content of Saturated Oxygen in Water at Different Temperature

ChartⅢ Content of Saturated Oxygen at Different Barometric pressure

Chart ⅣContent of Saturated Oxygen at Different Altitude

1. Brief Instruction:

Thanks for buying and using the model SX716 DO Meter (the following called “meter” in short).

Before using this meter, please read the operation manual carefully in order to help use and maintain it correctly. On the basis of improving instrument of performance constantly, we reserve the right of changing the content of this manual and accessories in case of not notifying in advance.

This meter is a perfect combination with the most advanced electronic technology, sensor technology and software design. The meter can measure the parameters of DO and temperature for high accuracy water solution. It is the best portable DO meter with the highest performance and the lowest cost. It is suitable for the trade such as the mining industry, power plant and environmental protection, etc., especially has more extensive application in the field and spot test.

Built-in microprocessor chip, beautiful appearance, multi-functional and easy to use, this meter has the following prominent features:

1.1.	Built-in microprocessor chip, with the intelligent functions of automatic calibration, automatic temperature compensation, data storage, function setting and auto power off and low voltage display etc. It equips with calibration solution and special carrying case, easy to use.

1.2.	Adopts digital filter and step slipping technology to intelligently improve meter’s response speed and result accuracy. “    ” will appear when reading to be stable.

1.3. Equips new type of DO electrode with temperature and salinity sensor which makes DO measuring mode has auto temperature salinity compensation and manual Barometric pressure compensation. It’s easy to use and accurate to measure.

1.4. Polarogram type DO electrode has special electrode calibration cap. Only needs 3 to 5 minutes for electrode polarization. Electrode adopts combination diaphragm cap which very convenient to use. Each DO electrode equips with three spare diaphragms caps.

1.5. Meter’s circuit board adopts SMT film-covering technology to improve meter’s production reliability.

1.6. Meter has the back light LCD display monitor.

1.7. Dustproof and waterproof meter meets IP57 rating.

2. Technical Parameters:

2.1. Dissolved Oxygen:

| Measuring range | (0 to 19.99) mg/L(ppm) (0 to 200.0) % |
| --- | --- |
| Resolution | 0.1/0.01 mg/L(ppm) 1/0.1% |
| Accuracy | Meter:±0.10 mg/L, connect meter with electrode:±0.40 mg/L |
| Response time | ≤30 s (25℃, 90% response) |
| Residual current | ≤0.1 mg/L |
| Temp. compensation range | (0 to 45) ℃ (automatic) |
| Salinity compensation range | (0 to 45) ppt (automatic) |
| Barometric pressure compensation range | (66 to 200) kPa (manual) |
| Automatic calibration | Air saturated by water; water saturated by air |
| Electrode type | Polarogram type |

2.2. Other Technical Parameters:

| Data storage | 100 groups |
| --- | --- |
| Storage content | Series number, measuring value measuring unit and temperature, |
| Power | Two AA batteries (1.5V x2) |
| Size and weight | Meter: (65×120×31)mm/180g Carrying case:(255 x 210x 50)mm/790g |
| Quality and safety certification | ISO9001:2000, CE and CMC |

2.3. Working Condition:

| Environment temperature | 5 to 35 ℃ (0.01 grade) |
| --- | --- |
| Environmental humidity | ≤85% |
| IP rating | IP57 Dustproof and waterproof |

3. Instructions to the Meter:

3.1. LCD Display:

① —— Parameters mode icon

② —— Measuring value

③ —— Serial number and icon as measuring value to be stored and recalled and indication icon for special states.

M+ — measuring value to be stored icon;

RM — reading to be recalled icon;

④ —— Measuring unit

⑤ —— Temperature measuring value and unit

⑥ —— Measuring value to be stable icon

⑦ —— Electrode calibration indicate icon

⑧ —— Indication icon of low voltage, appears when the voltage less than 2.6V, call attention to change the batteries.

3.2. Operation Keys:

The meter has 5 operation keys in all.

3.2.1.      — Switch key

3.2.2.      — Calibration key,.

(a) When in the measurement state, press this key to enter into calibration mode.

(b) When in the parameter setting state, press this key to change the number or the ON/OFF state.

3.2.3.      — Function key

Press the key to enter into the parameter setting modes P1, P2, P3…

3.2.4.      — Back light and entrance key

(a) When in the measuring state, short press (less than 1.5s) to open or close the back light display;

(b) When in the calibration state or the parameter setting state, press this key to confirm, and then the meter enters into measuring state;

(c) When in      mode, long press the key to in turn change unit symbol: mg/L→ppm→%, and then release the key.

3.2.5.      — The combination key of memory and recall

(a) Short time press (press time less than 1.5s) to save the measuring data, long time press (press time more than 2s) to recall the saved measuring data when in the measurement state.

(b) When in the parameter setting state, press this key to alter the number or ON/OFF state.

3.3. The Storage, Recall and Elimination of the Measuring Information:

3.3.1. Store the measuring information:

In the measuring mode, when the measuring data is stable and appear the “    ”, short-time press      key(＜1.5s), LCD will display “ M+ ” icon and storage serial number, and meanwhile memory all the measuring information. Meter can totally store 100 groups of measuring information.

3.3.2. Recall measuring information:

(a) Under the measuring mode, depress the      key, meter will recall the last stored information, and the storage number and “RM” icon will appear in the upper right corner of the LCD. Again press      or       key, meter will in turn recall all the measuring information, depress      or      key can quickly query;

(b) In the recalling mode (there are “ RM ” and storage serial number in the lower right corner of the LCD), press      key to return to the measuring mode.

3.3.3. Eliminate the stored measuring information:

In the recalling mode, depress the      key for 5 seconds, LCD will appear “      ” for 2 seconds. It means the storage has been eliminated and returns to measuring mode.

4. DO Measurement:

4.1. Preparation Work:

4.1.1. Press      key to turn on.

4.1.2. Choose units: Depress      to choose unit: mg/L, ppm or %; then release.

4.1.3. Check the DO500 DO electrode: see picture (4-1), unplug the bottom lid of electrode calibration cover, inspect whether the water stored sponge inside is soaked. If it is not soaked, please drop some appropriate purified water, then upside down the calibration cover to remove of rest purified water (Attention: no stored water in calibration cover); repeatedly inspect the electrolyte within electrode membrane cap can’t have air bubbles (except the smaller air bubbles). If there are lager air bubbles, please unplug the membrane cap and add electrolyte and plug on, then connect the electrode into meter to polarize for 5 minutes.

4.2. Meter Calibration:

Press      key, the meter enters into calibration mode, upper-right corner of LCD will display a twinkling “ CAL” . Access the DO electrode to calibration cover, and plug the cap of calibration cover tightly, place vertically to polarize 3 to 5 minutes, after the display values is stable and when appear the “     ” icon, again press       key to calibrate. LCD will display flashing 100%, the calibration is finished after several seconds and return to the measuring mode. If the display value is unstable, please wait for several minutes and again press       key until the display value is completely stable.

4.3. Water Sample Test:

4.3.1. To measure in the flowing water (water sample velocity of flow >5cm/s): Insert the DO electrode into water, water surface should excess location of the thermistor which is on the electrode, the direction for electrode and water is 45°to 75°angle, and swing slightly the electrode, then hold 3 to 5 minutes to take the reading after stable displayed.

4.3.2. To measure in static water: Insert DO electrode into water, water surface should excess location of the thermistor, the direction for electrode and water is 45°to 75°angle, quickly remove the electrode in water, The remove speed is >5cm/s, hold 3 to 5 minutes to take the reading after stable displayed;

4.3.3. To test in the slower flowing water: according to the method of item 4.3.1., but the remove speed of electrode should be faster.

4.4. Considerations:

4.4.1. The air temperature and water temperature should closer when test (≤10℃), if it is biggish different, please immerse the electrode into the sample water for about 10 minutes, then insert the electrode into the calibrate cover for 5 to 6 minutes according to item 4.2.1 and press the      key to do calibration.

4.4.2. After each start-up meter must be carried out as required electrode polarization and calibration, so do not turn off the meter while using (when in the DO mode, the time for producer setting for auto power off is 0).

4.4.3. During DO test, temperature effects a lot on measuring value, because the thermistor of DO electrode is installed on the electrode housing, direct touches with water, not equips in the electrolyte of internal electrode, both of them have difference for sensibility on the water temperature, generally will need 3 to 5 minutes to make the thermistor temperature and the electrolyte within electrode of the actual temperature reached unanimity, so the reading time must be > 3 minutes, or else they would have a greater error, especially when the electrode temperature and water temperature has larger difference, will need to extend the reading time;

4.4.4. DO measurement also can be affected by Barometric pressure, and has biggish inaccuracy (please see chart Ⅲ and Ⅳ). This meter has automatic Barometric pressure compensation function, and the measuring value has been amended by the Barometric pressure compensation.

4.4.5. DO electrode can not be static test in the static water, otherwise it will lead the results to be lower.

4.4.6. When measuring, there can not have any bubbles for DO electrode and the sensitive membrane which touches with water, otherwise will affect measuring accuracy.

4.4.7. Electrode within the electrolyte can not have air bubbles (except for the smaller air bubbles), otherwise it will affect the response speed and accuracy. If a large air bubble, should screw out membrane cap, then add the electrolyte and membrane cap is screwed up;

4.4.8. Should keep the surface of DO electrode sensitive membrance to be wet to prevent the cathode electrolyte withered, there is a piece of water stored sponge in calibration cover bottom, see picture (4-1), user must always maintain a moist sponge when the sponge is dried should be dropping some water on it (Let sponge saturated absorption, but you can not have a water flow). And tightly twist electrode cap, dissolved oxygen electrode should be kept in humid conditions.

4.4.9. The salinity electrode be installed in DO electrode (see picture (4-1)), the surface is plated with a layer of metal platinum black to reduce electrode polarization. The surface of platinum can not be rubbed. It only can be washed by shaking in water in case of damaging the platinum black coating. Wash the organic component pollution with tepid water containing detergent or with alcohol.

4.4.10. When it is abnormal for calibrating or displaying, please set the mode P7 to be “On” to restore meter to producer setting state, and then do calibrate and measure.

4.5. Parameter Setting:

4.5.1. DO parameter setting schedule (Chart (4-1))

Chart (4-1)

| Prompt Mark | Parameter Setting Items | Code | Parameter |
| --- | --- | --- | --- |
| P1 | Resolution selection |  | 0.01/0.1(mg/L and ppm) 0.1/1(%) |
| P2 | Salinity calibration |  |  |
| P3 | Barometric pressure setting |  |  |
| P4 | Temperature unit setting |  | ℃ ℉ |
| P5 | Back light display time setting |  | 0-1-3-6 min |
| P6 | Auto power off time setting |  | 0-10-20 min |
| P7 | Restore to producer setting |  | OFF-On (shut-set) |

4.5.2. Resolution selection (P1)

(a) Depress      key and enter into P1 mode, see picture (4-2);

(b) Press      key to choose resolution: 0.01→0.1;

(c) Press      key to enter into the next parameter setting or    press      return to the measuring interface.

4.5.3. Salinity calibration setting (P2)

(a) Under P1 mode, press      key to enter into P2 mode, show as picture (4-3).

(b) Immerse the DO electrode into 12.88mS/cm calibration solution, solution level should excess salinity electrode, still it after stirring, then press the      key when the value to be stable and appears an icon “    ”, the LCD displays a twinkling “12.9”. The calibration is finished after several minutes and displays a stable salinity value.

(c) Press      key to enter into the next parameter setting or press      key to confirm and return to the measuring mode.

4.5.4. Barometric pressure setting (P3)

(a) Under mode P2, press      key to enter into P3 mode, the LCD displays a barometric pressure measuring value set earlier (unit is kPa), such as 101.3 kPa, show as picture (4-4).

(b) Press      or       to alter according to standard barometric pressure value, and depress the      or      to quickly alter the value.

(c) Press      key to enter into next parameter setting or press      key to confirm and return to measuring mode.

4.5.5. Temperature unit ℃/℉ setting (P4)

(a) Short press      key in mode P3 to enter into mode P4, see picture (4-5).

(b) Press      or      key to choose temperature unit: ℃or℉.

(c) Press      key to enter into next parameter setting or press      key to conform and return to measuring mode.

4.5.6. Back light display time setting (P5)

(a) Short press      key in mode P4 to enter into mode P5, see picture (4-6).

(b) Press      or      key to choose the time of back light auto power off: 0 min, 1 min, 3 min or 6 min. The back light display function will be closed if choosing a 0min.

(c) Press      key to enter into next parameter setting or press      key to conform and return to measuring mode.

(d) The producer setting for P5 is 1min.

4.5.7.Auto power off time setting (P6)

(a) Short press      key in mode P5 to enter into mode P6, see picture (4-7).

(b) Press      or      key to choose the time: 0min, 10min or 20min. The auto power off function will be closed if choosing 0min.

(c) Press      key to enter into next parameter setting or press      key to conform and return to measuring mode.

(d) The producer setting for P6 is 0 min (see item4.4.2).

4.5.8. Restore to producer setting (P7)

(a) Short press      key in mode P6 to enter into mode P7, see picture (4-8).

(b) Press      or      key to choose “    ”, means that parameters have been restored to the producer setting mode, and restore to measuring mode after 2 seconds.

4.6. Change the Membrane Cap:

When electrode response time to be longer and displaying value appear obviously error, or when sensitive membrane of DO electrode has a wrinkle, crack or damaged, should replace a membrane cap according to the following steps.

(a) Loose the membrane cap and remove;

(b) Wash the electrode without membrane cap with purified water and throw off the water on it.

(c) Polish lightly on cathode surface (gold slice) with a piece of clear velvet or tissue;

(d) Inject electrolyte into a new membrane cap slowly and do not has any air bubbles appeared. After injected, should check the electrolyte to see whether there are air bubbles in it. If found there are air bubbles, please carefully knock on the membrane cap to dispel them.

(e) Put membrane cap on the table and set in the electrode vertically, slowly clockwise twist, and then twist tightly. Temporality, redundant electrolyte will squeeze out, please wipe it out with tissue and clear it in the purified water.

(f) Check there is no air bubbles in electrolyte (beside the smaller air bubbles), otherwise, re-assembly is needed.

(g) No touching the sensitive membrane when use the electrode and replace the membrane cap for the sweat and grease will affect the quality of membrane and decrease the rate of oxygen permeability.

4.7. Zero Oxygen Calibration:

Generally need to do zero oxygen calibration in the situation of replacing electrode, replacing membrane cap, and long time without using. Usually no need to do zero oxygen calibration. To do zero oxygen calibration by the following below steps:

(a) Preparing 100mL anaerobic water: In 100 mL beaker add 5g anhydrous sodium sulfite (Na2SO3), plus 100 mL purified water stir to dissolve, anaerobic water is effective within 24 hours.

(b) To polarize electrode by connecting it into meter for 15min, and to do calibration according to the item 4.2.1;

(c) Put the electrode into anaerobic water, press      key to make meter enter into calibration mode, when displaying value is ≤ 0.15mg/L (≤5min), again press

key to calibrate, then the LCD displays a twinkling 0.0%, the calibration finishes after several seconds and the meter displays 0.00 mg/L, then wash the electrode with purified water;

(d) If the meter displayed value ≤0.02mg/L within 5 minutes, means the response speed of meter and residual current are good, need not to do zero oxygen calibration, press      key to return to measuring mode.

(e) If the value of meter is displayed as >0.15 mg/L after 5 minutes, means the response speed of meter is slow and large residual current, at this time can change the membrane cap or remove the membrane cap and take out the polishing paper in the accessories to lightly wipe the catholic of gold plane (to the surface along the arc of gold-polished), and clean surface of cathode with soft cloth to tissue), and wash electrode with purified water and throw off the water on it. Add some electrolyte into the membrane cap, re-assembly and screw tightly, then to do full scale and zero oxygen calibration according to item 4.2. and 4.7.

4.8. Salinity Calibration:

The meter has salinity automatic compensation, but this salinity calibration is just need to apply in replacing a new electrode and long-time without using. In general do not need to do salinity calibration. The meter has already been made salinity calibration before leaving factory, so it is no need to do salinity calibration for fist time using. For the procedure of salinity calibration, please refer to P13 item 4.5.3.

4.9. Barometric pressure Setting:

The meter has the function of manual Barometric pressure compensation. It is suggested to reset the value according to the standard Barometric pressure table when there is a biggish change of Barometric pressure while using to ensure the accuracy of Barometric pressure compensation（Please see the Chart Ⅲ and Ⅳ. For the procedure of Barometric pressure calibration, please see P13 item 4.5.4.

5. Meter’s Complete Kit:

5.1. Model SX716 DO meter

5.2. DO500 DO electrode

5.3. DO503 membrane cap (DO electrode use)

5.4. DO502 DO electrode inner solution (30mL)

5.5. Cathode polishing paper

5.6. Screw driver

5.7. Spare AA batteries

5.8. Operation manual

5.9. Carrying case

6. Warranty:

6.1. We warrant this meter to be free of charge maintain, replace the parts or products under normal using circumstances, from purchased time within one year caused by manufacturing bad and unable to work.

6.2. Attached electrodes do not belong to this warrant range. But, if the newly purchased electrode went wrong without using, it’s free of charge to maintain and replace.

6.3. The above warranty is not apply to defects resulting from action of user such as misuse, improper wiring, operation outside of specification, improper maintenance or repair, or unauthorized modification.

Chart І Meter Parameter Setting Schedule

| Mode | Prompt Mark | Parameter Setting Items | Code | Parameter |
| --- | --- | --- | --- | --- |
| Dissolved Oxygen | P1 | Resolution selection |  | 0.01/0.1(mg/L and ppm) 0.1, 1(%) |
|  | P2 | Salinity calibration |  |  |
|  | P3 | Barometric pressure setting |  |  |
|  | P4 | Temperature unit setting |  | ℃ ℉ |
|  | P5 | Back light display time setting |  | 0-1-3-6min |
|  | P6 | Auto power off time setting |  | 0-10-20min |
|  | P7 | Restore to producer setting |  | OFF-On (shut-set) |

Chart Ⅱ Content of Saturated Oxygen in Water at Different Temperature

| Temperature ℃ | DO mg/L | Temperature ℃ | DO mg/L | Temperature ℃ | DO mg/L |
| --- | --- | --- | --- | --- | --- |
| 0 | 14.64 | 16 | 9.86 | 32 | 7.30 |
| 1 | 14.22 | 17 | 9.66 | 33 | 7.18 |
| 2 | 13.82 | 18 | 9.46 | 34 | 7.07 |
| 3 | 13.44 | 19 | 9.27 | 35 | 6.95 |
| 4 | 13.09 | 20 | 9.08 | 36 | 6.84 |
| 5 | 12.74 | 21 | 8.90 | 37 | 6.73 |
| 6 | 12.42 | 22 | 8.73 | 38 | 6.63 |
| 7 | 12.11 | 23 | 8.57 | 39 | 6.53 |
| 8 | 11.81 | 24 | 8.41 | 40 | 6.43 |
| 9 | 11.53 | 25 | 8.25 | 41 | 6.34 |
| 10 | 11.26 | 26 | 8.11 | 42 | 6.25 |
| 11 | 11.01 | 27 | 7.96 | 43 | 6.17 |
| 12 | 10.77 | 28 | 7.82 | 44 | 6.09 |
| 13 | 10.53 | 29 | 7.69 | 45 | 6.01 |
| 14 | 10.30 | 30 | 7.56 |  |  |
| 15 | 10.08 | 31 | 7.43 |  |  |

ChartⅢ Content of Saturated Oxygen at Different Barometric pressure

| Barometric pressure | Dissovled Oxygen Concentration (mg/L) |
| --- | --- |
| mmHg | kPa | 15℃ | 25℃ | 35℃ |
| 750 | 100.00 | 9.94 | 8.14 | 6.85 |
| 751 | 100.13 | 9.96 | 8.15 | 6.86 |
| 752 | 100.26 | 9.97 | 8.16 | 6.87 |
| 753 | 100.40 | 9.98 | 8.17 | 6.88 |
| 754 | 100.53 | 9.99 | 8.18 | 6.89 |
| 755 | 100.66 | 10.00 | 8.20 | 6.90 |
| 756 | 100.80 | 10.01 | 8.21 | 6.91 |
| 757 | 100.93 | 10.03 | 8.22 | 6.92 |
| 758 | 101.06 | 10.04 | 8.23 | 6.93 |
| 759 | 101.20 | 10.07 | 8.24 | 6.94 |
| 760 | 101.33 | 10.08 | 8.25 | 6.95 |
| 761 | 101.46 | 10.09 | 8.26 | 6.96 |
| 762 | 101.60 | 10.11 | 8.27 | 6.97 |
| 763 | 101.73 | 10.12 | 8.28 | 6.98 |
| 764 | 101.86 | 10.14 | 8.30 | 6.99 |
| 765 | 102.00 | 10.15 | 8.31 | 7.00 |
| 766 | 102.13 | 10.16 | 8.32 | 7.01 |
| 767 | 102.26 | 10.18 | 8.33 | 7.02 |
| 768 | 102.40 | 10.19 | 8.34 | 7.02 |
| 769 | 102.53 | 10.21 | 8.35 | 7.03 |
| 770 | 102.66 | 10.22 | 8.36 | 7.04 |
| 771 | 102.80 | 10.23 | 8.37 | 7.05 |
| 772 | 102.93 | 10.25 | 8.39 | 7.06 |
| 773 | 103.06 | 10.26 | 8.40 | 7.07 |
| 774 | 103.19 | 10.28 | 8.41 | 7.08 |
| 775 | 103.33 | 10.29 | 8.42 | 7.09 |

Conversion between mmHg and kPa: mmHg×0.13333=kPa

DOpt = P×DOt ÷ 760

Note:  DOpt — DO concentration under temperature t, Barometric pressure P, mg/L;

P — Barometric pressure, mmHg;

DOt — DO concentration under temperature t, Barometric pressure 760mmHg, mg/L;760 — Barometric pressure, mmHg.

ChartⅣ Content of Oxygen at Different Altitude

| Altitude | Barometric pressure | DO | Altitude | Barometric pressure | DO |
| --- | --- | --- | --- | --- | --- |
| Foot | meter | kPa | mmHg | mg/l | Foot | meter | kPa | mmHg | mg/l |
| 0 | 0 | 101.3 | 760 | 8.25 | 7500 | 2287 | 77.1 | 579 | 6.28 |
| 500 | 152 | 99.34 | 746 | 8.09 | 8000 | 2439 | 75.63 | 568 | 6.16 |
| 1000 | 305 | 97.6 | 733 | 7.95 | 8500 | 2591 | 74.44 | 559 | 6.06 |
| 1500 | 457 | 95.87 | 720 | 7.81 | 9000 | 2744 | 72.97 | 548 | 5.94 |
| 2000 | 610 | 94.28 | 708 | 7.68 | 9500 | 2896 | 71.64 | 538 | 5.83 |
| 2500 | 762 | 92.54 | 695 | 7.54 | 10000 | 3049 | 70.17 | 527 | 5.71 |
| 3000 | 915 | 90.95 | 683 | 7.41 | 10500 | 3201 | 68.84 | 517 | 5.61 |
| 3500 | 1067 | 89.35 | 671 | 7.28 | 11000 | 3354 | 67.38 | 506 | 5.49 |
| 4000 | 1220 | 87.75 | 659 | 7.15 | 12000 | 3659 | 66.58 | 500 | 5.42 |
| 4500 | 1372 | 86.15 | 647 | 7.02 | 13000 | 3963 | 65.78 | 494 | 5.36 |
| 5000 | 1524 | 84.56 | 635 | 6.89 | 14000 | 4268 | 64.98 | 488 | 5.29 |
| 5500 | 1677 | 83.09 | 624 | 6.77 | 15000 | 4573 | 64.18 | 482 | 5.23 |
| 6000 | 1829 | 81.63 | 613 | 6.65 | 16000 | 4878 | 63.38 | 476 | 5.16 |
| 6500 | 1982 | 80.03 | 601 | 6.52 | 17000 | 5183 | 62.58 | 470 | 5.10 |
| 7000 | 2134 | 78.56 | 590 | 6.40 | 18000 | 5488 | 61.79 | 464 | 5.03 |

Shanghai San-Xin Instrumentation, Inc.

Add: 3/F. Building No.4, No. 471 Guiping Road, Shanghai, China. 200233

Tel: +86-21-63362480                   Fax: +86-21-64956880

Website: www.shsan-xin.com            E-mail: wxmab@shsan-xin.com

# Text boxes

3

4

5

5

6

7

8

8

9

10

10

12

15

16

17

17

17

18

1unit

1pc

1pc

1bottle

6 — Temperature electrode

7 — Membrane cap

8 — Cathode (gold slice)

9 — Anode

10 — Salinity electrode

1 — DO500 DO electrode

2 — Calibration cover cap

3 — Calibration cover

4 — Water stored sponge

5 — Calibration cover bottom lip

1pc

1pc

2pcs

1pc

1unit