Model SX721 pH/ORP	Meter

#### Operation Manual

Shanghai San-Xin Instrumentation Inc.

SX700 Series of Portable Meters

1. Model SX711 pH / mV Meter

2. Model SX712 ORP Meter

3. Model SX713 Conductivity / TDS / Salinity / Resistivity Meter

4. Model SX716 DO Meter

5. Model SX721 pH / ORP Meter

6. Model SX723 pH /mV / Conductivity Meter

7. Model SX725 pH / mV / DO Meter

8. Model SX726 Conductivity / DO Meter

9. Model SX731 pH / ORP / Conductivity Meter

10. Model SX736 pH / mV / Conductivity / DO Meter

11. Model SX751 pH / ORP / Conductivity / DO Meter

Content

1. Summary ------------------------------------------------------------------------------------------------

2. Technical Parameters ---------------------------------------------------------------------------------

3. Meter Overview-------------------------------------------------------------------------------------------

3.1. LCD Display ---------------------------------------------------------------------------------------

3.2. Operation Keys ------------------------------------------------------------------------------------

3.3. The Storing, displaying and clearing of the Measuring Information -----------------------

4. pH Measurement --------------------------------------------------------------------------------------

4.1. Preparation Work ----------------------------------------------------------------------------------

4.2. Meter Calibration ----------------------------------------------------------------------------------

4.3. Sample Test ----------------------------------------------------------------------------------------

4.4. Parameter Setting ----------------------------------------------------------------------------------

4.5. Considerations -------------------------------------------------------------------------------------

4.6. The Self-diagnose Information ------------------------------------------------------------------

5. ORP Measurement ------------------------------------------------------------------------------------

5.1. Sample Test ----------------------------------------------------------------------------------------

5.2. Considerations -------------------------------------------------------------------------------------

5.3. Parameter Setting ----------------------------------------------------------------------------------

6. Meter’s Complete Kit --------------------------------------------------------------------------------

7. Warranty ------------------------------------------------------------------------------------------------

TableⅠ Meter Parameter Setting Table

TableⅡ Meter Restore to Producer Setting Table

Table Ⅲ Code Icon and Abbreviation Table

1. Summary:

Thanks for buying and using the model SX721 pH/mV Meter (the following called “meter” in short).

Before using this meter, please read the operation manual carefully in order to help use and maintain it correctly. On the basis of improving instrument of performance constantly, we reserve the right of changing the content of this manual and accessories in case of not notifying in advance.

This meter is a perfect combination with the most advanced electronic technology, sensor technology and software design The meter can measure the parameters of pH, mV and temperature for high accuracy solution. It is the best portable pH meter with the highest performance and the lowest cost. It is suitable for the trade such as the mining industry, power plant, water treatment projects and environmental protection, etc., especially has more extensive application in the field and spot test.

Built-in microprocessor chip, beautiful appearance, multi-functional and easy to use, this meter has the following prominent features:

1.1. Double parameters meter, equips with pH and ORP electrode,calibration solution and special carrying case, easy to use.

1.2.	Built-in microprocessor chip, with the intelligent functions of automatic calibration, automatic temperature compensation, data storage, function setting and automatic self-diagnose, auto power off and low voltage display etc.

1.3.	Adopts digital filter and step slipping technology to intelligently improve meter’s response speed and result accuracy, “    ” will appear when reading to be stable.

1.4.	Automatically recognize 13 kinds of pH standard buffer solution. User can choose anyone from three series of buffer solutions: Europe & U.S.A. series, NIST series, or China series.

1.5.	Purified water and ammonia added purified water pH measuring mode can be set up, for this two special pH measuring modes, besides the general slope temperature compensation, also added function of nonlinearity solution temperature compensation to make measurement more accurate, especially suitable for the fields of electric power and petrochemical etc.

1.6. Meter’s circuit board adopts SMT film-covering technology to improve meter’s production reliability.

1.7. Meter has the back light LCD display monitor.

1.8. Dustproof and waterproof meter meets IP57 rating.

2. Technical Parameters:

2.1. pH:

| Measuring range | (-2.00 to 19.99) pH |
| --- | --- |
| Resolution | 0.1/0.01 pH |
| Accuracy | Meter:±0.01pH; Overall: ±0.02pH |
| Input current | ≤2×10-12 A |
| Input impendance | ≥1×1012 Ω |
| Stability | ±0.01 pH/3h |
| Temp. compensation range | (0 to 100) ℃ (automatic) |

2.2. ORP:

| Measuring range (mV/ORP/EH) | -1999 mV to 0 to 1999mV |
| --- | --- |
| Resolution | 1mV |
| Accuracy | Meter:±0.1% FS; Overall: ±15mV |

2.3. Other Technical Parameters:

| Data storage | 200 groups |
| --- | --- |
| Storage content | Series number, measuring value, measuring unit and temperature |
| Power | Two AA batteries (1.5V x2) |
| Size and weight | Meter: (65×120×31)mm/180g Carrying case:(360 x 270x 76)mm/1.5Kg |
| Quality and safety certification | ISO9001:2000, CE and CMC |

2.4. Working Condition:

| Environment temperature | 5 to 35 ℃ (0.01 grade) |
| --- | --- |
| Environmental humidity | ≤85% |
| IP rating | IP57 Dustproof and waterproof |

3. Meter Overview:

3.1. LCD Display:

① —— Parameters mode icon

② —— Measuring value

③ —— Serial number and icon as measuring value to be stored and displayed.

M+ — measuring value to be stored icon;

RM — measuring value to be displayed icon;

④ —— Measuring unit

⑤ —— Temperature measuring value and unit

⑥ —— Measuring value to be stable icon

⑦ —— Electrode calibration indicating icon

⑧ —— Indicating icon of low voltage, appearing when the voltage is less than 2.6V, calling attention to change the batteries.

3.2. Operation Keys:

The meter has 5 operation keys in all.

3.2.1.      — Switch on/off key

3.2.2.      — Calibration key

(a) When in pH measurement state, press this key to enter into calibration mode.

(b) When in the parameter setting state, press this key to change the number or the ON/OFF state.

3.2.3.      — Function key

(a) In pH measuring mode, short-time press (<1.5s) to change the measuring parameters, the meter will display cyclically      →       ,press (>2s) to enter into the parameter setting mode, cyclically displaying P1, P2, P3…

(b) In ORP measuring mode,press this key to enter into parameters setting mode, cyclically displaying P1, P2, P3…

3.2.4.      — Back light and entrance key

(a) When in the measuring state, short-time press (<1.5s) to turn on or off the back light display;

(b) When in the calibration state or the parameter setting state, press this key to confirm, and then the meter enters into measuring state;

(c) When in      mode, press the key without releasing to change pH resolution: 0.01→0.1 pH cyclically; Then release after choosing made.

3.2.5.      — The combination key of storing and displaying

(a) When it is in the measurement state, short-time press (<1.5s) to store the measuring data; long-time press (>2s) to display the stored measuring data.

(b) When in the parameter setting state, press this key to change the number or ON/OFF state.

3.3. The Storing, Displaying and Clearing of the Measuring Information:

3.3.1. Store the measuring information:

In the measuring mode, when the measuring data is stable and appears the “    ”, short-time press      key(<1.5s), LCD will display “ M+ ” icon and stored serial number, and meanwhile will store all the measuring information. Meter can separately store 100 groups of measuring information in the mode of pH and ORP, it can totally store 200 groups.

3.3.2. Display measuring information:

(a) Under the measuring mode, long-time press the      key, meter will display the last stored information, and the stored number and “RM” icon and the complete measuring information will appear in the lower right corner of the LCD. Again press       or

key, meter will in turn display all the measuring information, press       or

key without releasing can make a quick query of measuring information under other serial numbers;

(b) In the displaying mode (there are “ RM ” and stored serial number in the upper right corner of the LCD), press      key to return to the measuring mode.

3.3.3. Clear the stored measuring information:

In the displaying mode, press the      key for 5 seconds, LCD will appear “      ” for 2 seconds. It means the memory has been cleared, and meter returns to measuring mode.

4. pH Measurement:

4.1. Preparation Work:

Press      key to turn on the meter. Plug the model 201T-S pH/ATC three in one combination pH electrode into meter’s socket, then the meter enters into pH measuring mode automatically.

4.2. Meter Calibration:

4.2.1. Press       to enter into calibration mode, LCD displays the flickering “    ” indicating to enter into the first point calibration.

4.2.2. Wash the pH electrode in purified water and shake off water, then insert probe into the pH7.00 buffer solution, stir the probe and then let it still, waiting for the data stable and appearing “      ”, then again press key      , the LCD will appear a flickering 7.00 pH, calibration finishes after several seconds and meter appears a stable pH value and a flickering “     ”, indicating the first point calibration has been finished and meter enters into the second point calibration.

4.2.3. Wash the pH electrode in purified water and shake off water, then insert it into the pH4.00 buffer solution, stir the probe and then let it still, waiting for the data stable and appearing “      ”, then again press key      , the LCD will appear a flickering 4.00 pH, calibration finishes after several seconds, LCD will appear a stable pH value and a flickering “    ”, indicating the second point calibration has been finished and meter enters into the third point calibration.

4.2.4. Wash the pH electrode in purified water and shake off water, then insert probe into the pH10.01 buffer solution, stir the probe and then let it still, waiting for the data stable and appearing “     ”, then again press key      , the LCD will appear a flickering 10.01 pH, calibration finishes after several seconds, appearing a stable pH value and “            ” three calibration indication icon, see picture (4-1), indicating the three-point calibration has been finished and meter enters into measuring mode.

4.2.5. Note:

(a) This meter can adopt random one-point, two-point or three-point automatic calibration, after the first point calibration (see item 4.2.2.), press      to confirm one-point calibration and enters into measuring mode. The indication icon “     ” for one-point calibration will appear on the lower left corner of LCD. When the measuring accuracy is ≥±0.1pH, user just need to choose one kind buffer solution to take one-point calibration according to the measuring range.

(b) After the second point calibration, (see item 4.2.3.), press      to confirm two-point calibration and enter into measuring mode. The indication icon “        ” for tow-point calibration will appear on the lower left corner of LCD. User can choose pH4.00 and pH7.00 to calibrate if the measurement is just within the acidity range and choose pH7.00 and pH10.01 to calibrate if just within the alkalinity range.

(c) User should choose three-point calibration so as to reach a more accurate measurement if the measuring range is wide or if the electrode has been used for long or has ageing phenomenon. As to the new electrode which be used for the first time, it must be calibrated by three-point calibration to keep the unanimity of the meter slope adjustment with the pH electrode.

4.3. Sample Test:

Insert pH probe into the sample solution after washing and shaking off water, stir the probe and then let it still, when the LCD appears the icon “     ” to take the reading ,which is the pH value of the sample.

Note: According to the pH equal temperature measuring theory: the closer the temperature of the sample solution with the calibration solution, the more accurate the measuring value will be acquired. So please comply with this theory.

4.4. Parameter Setting:

4.4.1. pH measuring parameter setting Table (Table (4-1))

Table (4-1)

| Prompt Mark | Parameter Setting Items | Code | Parameters |
| --- | --- | --- | --- |
| P1 | pH buffer solution series selection |  | USA(Europe & U.S.A series) NIS (NIST series) CH (China series) |
| P2 | Purified water pH temperature compensation setting |  | OFF-On (shut-set) |
| P3 | Ammonia added purified water pH temperature compensation setting |  | OFF-On (shut-set) |
| P4 | Temperature unit setting |  | ℃ ℉ |
| P5 | Back light display time setting |  | 0-1-3-6min |
| P6 | Auto power off setting |  | 0-10-20min |
| P7 | Restore to producer setting |  | OFF-On ( shut-set) |

4.4.2. pH buffer solution series selection (P1)

(a) Long-time press      key(>2s) to enter into P1 mode: see picture (4-2).

(b) Press      or      key to choose buffer solution series:

( Europe & U.S.A series) —1.68, 4.00, 7.00,10.01 and 12.45 pH

(NIST series) — 1.68, 4.01, 6.86, 9.18 and 12.45 pH

(China series) — 1.68, 4.00, 6.86, 9.18 and 12.46 pH

(c) Press      key to enter into next parameter setting or press      key to conform and return to measuring mode.

4.4.3. Purified water pH temperature compensation setting (P2)

(a) Press      key in the mode P1 to enter into mode P2, see picture (4-3).

(b) Press      or      to choose “     ” (purified water pH temperature compensation setting) or “      ” (shut).

(c) Press      key to enter into next parameter setting or press      key to conform and return to measuring mode.

(d) The producer setting is “ OFF ”.

Note: There will appear“       ” icon in the upper right corner of the LCD if purified water temperature compensation function was set.

4.4.4. Ammonia added purified water pH temperature compensation setting (P3)

(a) Press      key in mode P2 to enter into mode P3, see picture (4-4).

(b) Press      or      key to choose “    ”（ammonia added purified water pH temperature compensation setting）or “      ” (OFF).

(c) Press      key to enter into next parameter setting or press       key to conform and return to measuring mode.

(d) The producer setting is “ OFF ”.

Note: If set the ammonia purified water pH temperature compensation function, there will appear“       ”icon in the right upper corner of the LCD when in the measuring mode.

4.4.5. Temperature unit ℃/℉ setting (P4)

(a) Press       key in mode P3 to enter into mode P4, see picture (4-5).

(b) Press      or      key to choose temperature unit: ℃or℉.

(c) Press      key to enter into next parameter setting or press      key to conform and return to measuring mode.

4.4.6. Back light display time setting (P5)

(a) Press      key in mode P4 to enter into mode P5, see picture (4-6).

(b) Press      or      key to choose the time of back light auto power off: 0 min, 1 min, 3 min or 6 min. The back light display function will be closed if choosing a 0min.

(c) Press      key to enter into next parameter setting or press      key to conform and return to measuring mode.

(d) The producer setting for P5 is 1min.

4.4.7.Auto power off time setting (P6)

(a) Press      key in mode P5 to enter into mode P6, see picture (4-7).

(b) Press      or      key to choose the time: 0min, 10min or 20min. The auto power off function will be closed if choosing 0min.

(c) Press      key to enter into next parameter setting or press      key to conform and to return to measuring mode.

(d) The producer setting for P6 is 10min.

4.4.8. Restore to producer setting (P7)

(a) Press      key in mode P6 to enter into mode P7, see picture (4-8).

(b) Press      or      key to choose “     ”, means that all the pH measuring parameters have been restored to the producer setting mode, and restore to measuring mode after 2 seconds.

4.5. Considerations:

4.5.1.Calibration time of meter depends on the sample, electrode performance and required accuracy. For high accurate measurement (≤±0.02pH), which should be calibrated immediately with high accurate standard buffer solution, for general accuracy measuring (≥±0.1pH), which can be used almost one week or long time once be calibrated one time.

4.5.2. The meter must be recalibrated in the following situations:

(a) New changed or unused electrode for a long time;

(b) After measuring acid (pH<2) or alkaline (pH>12) solution;

(c) After measuring solution which contains fluoride and concentrated organic solution;

(d) The solution’s temperature is much different with calibration temperature.

4.5.3. The soaking solution in the protecting bottle of front pH electrode is to keep the glass bulb and junction activating. Loose the capsule, pull out the electrode and wash it in purified water before measuring. Insert the electrode and screw tight the capsule after measuring to prevent the solution leaking. If the soaking solution is turbid or moldy, please wash and change a new one at once.

4.5.4. The preparation of the soaking solution: take 25g analytical pure KCL, dissolved with purified water and dilute to 100mL. Electrode should avoid soaking in purified water protein solution and acid fluoride solution for a long time as well avoid getting touching with organic silicon lipidic matters.

4.5.5. To calibrate the meter with the given value pH buffer solution, the pH value of the standard buffer solution must be reliable so as to improve the accuracy. Buffer solution should be changed in time after many times using.

4.5.6. Always keep the meter clean and dry, especially for the socket of meter and electrode, otherwise it may lead to an inaccurate measurement or invalidity. To clean and dry them with medical cotton with dehydrated alcohol if there are any dirty.

4.5.7. The sensitive glass bulb in the front of combination electrode should not touch with hard things, any broken and rough will make the electrode invalidity. Before and after measuring, the electrode should be washed with purified water, and dry electrode after washing, don’t clean glass bulb with tissue for it will effect stability of electrode potential and enlarge response time. The electrode should be washed many times for removing the sample stuck on the electrode, or wash with suitable solvent then clean the solvent with purified water after measuring sticky sample.

4.5.8. An electrode be used for a long time, or measured solution which contains a polluting solute easily for the sensitive bulb, or a substance resulting in jam at the junction, the electrode will be getting passivated, its sensitivity will decrease and its response is getting slow, the reading are not correct. It could adopt the following method for various cases:

(a) The glass bulb is contaminated and aging: Put the electrode into 0.1mol/L dilute hydrochloric acid (Preparation: diluted 9mL hydrochloric acid to 1000mL with purified water) for 24h. Rinse it with purified water, then dipped it into the electrode dipping solution for 24h. If the passivation is serious, then user can also put the bulb of electrode into 4% HF (hydrofluoric acid) or the electrode activation solution for 3 to 5 seconds, rinsing it with purified water, and dipped it in the electrode soaking solution for 24h to renew it.

(b) Wash for contaminated glass bulb and junction: (For reference)

Contamination                         Abluent

Inorganic metal oxide               diluted acid less than 1mol/L

Organic lipidic matter              dilute washing (weak alkaline)

Resin macromolecule matter           dilute alcohol, acetone, ether

Proteinic haematocyte sediment    Acidic enzymatic solution (such as dried yeast)

Kinds of paint                   dilute bleacher, peroxide

4.5.9. pH electrode using period is about 1 year, but its life will be shortened if using condition is poor or incorrect maintenance. So it should be replaced immediately after electrode become aging or invalid.

4.5.10. When it appears an abnormal reading when calibration or displaying, please set P7 as “ON” to restore the meter to producer setting mode, and then to calibrate and measure again.

4.6. The Self-diagnose Information:

When using, there might appear the following icons. This is the meter’s self-diagnose information, which can help to know some information about the meter or the electrode when using:

4.6.1. The stable icon -2.00 pH or the 19.99 pH — this icon appeared when the value has surpassed the measuring range. There will also appear such signs when the electrode is not well connected with the meter or when the electrode is not insertd into the solution. This is a normal phenomenon.

4.6.2. “       ” — Electrode zero potential to be exceeded (<-60mV or >60mV)

4.6.3. “       ” — Electrode slope to be exceeded (< 85% or >105%)

When appears “       ” or “       ”, the meter can not work, please take the following check:

(a) Check if the electrode bulb has air bubble, if has, please shake it hardly.

(b) Check the quality of buffer solution, if it goes bad or the value has biggish error.

(c) Set the meter to producer setting mode (for details see P12 item 4.4.8.), then recalibrate it.

If still can not recover the normal state after doing the above checking, please replace a new pH electrode.

5. mV Measurement:

5.1. Sample Test:

5.1.1. Press      key to turn on the meter, plug the 301Pt-s ORP electrode into the meter’s corresponding socket, and then the meter enters into ORP measuring mode automatically,

5.1.2 Washing the ORP electrode in purified water and shake off the water, insert the probe into the sample solution, slowly stir it and then let it still. When appears a “    ” ,the  stable reading, is the ORP measuring value.

5.2. Considerations:

5.2.1. When you use ORP electrode, the meter needs no calibration, but if you are uncertain about the quality of ORP electrode or the testing result, you can use the ORP standard solution to test the mV value of this meter to check ORP electrode or meter’s accuracy.

The ORP standard solution of meter equipment,222mV(25℃) ±15mV can be kept for one year.

Table (5-1)

| ℃ mV | ℃ mV | ℃ mV | ℃ mV |
| --- | --- | --- | --- |
| 10 242 | 20 227 | 30 215 | 38 205 |
| 15 235 | 25 222 | 35 209 | 40 201 |

5.2.2. The ORP electrode clearing and actavition：ORP electrode after long time used, the polluting on the surface of platinum can lead the measure result is inaccurate and response slowly, so we following these steps to wash:

（a）To inorganic pollution, we can put electrode into 0.1mol/L dilute hydrochloric acid 30 minutes after washing by the purified water, then put it into the electrode solution for 6h, then use it

（b）To organic oil polluting and the oil film polluting, we first use abluent to wash the surface of the platinum, second use the purified water, then put it into electrode solution for 6h, then use it

（c） If the surface of the platinum was terribly polluted and became the oxidative film, we first use toothpaste to polish the surface of the platinum, second use the purified water to wash, then put it into electrode solution for 6h, then use it.

5.3. Parameter Setting:

5.3.1. ORP measurement parameter setting Table (Table (5-2))

Table (5-2)

| Prompt Mark | Parameter Setting Items | Code | Parameters |
| --- | --- | --- | --- |
| P1 | Back light display time setting |  | 0 -1-3-6 min |
| P2 | Auto power off time setting |  | 0 -10-20 min |

5.3.2. Back light display time setting (P1)

Press       key until meter displaying P1, see detailed operations in P12 item4.4.6.

5.3.3. Auto power off time setting (P2):

Press        key in mode P1, the meter enters into mode P2, see detailed operations in P12 item 4.4.7.

6. Meter’s Complete Kit:

6.1. Model SX721 pH/ORP meter

6.2. 201T-S plastic pH/ATC three-in-one combined electrode

6.3. 301Pt-s plastic ORP combined electrode

6.4. pH standard buffer solution

(pH4.00, pH7.00 and pH10.01/50mL)

6.5. 220mV ORP standard solution(50mL)

6.6. Screw driver

6.7. Spare AA batteries

6.8. Operation manual

6.9. Carrying case

7. Warranty:

7.1. We warrant this meter to be free of charge maintain, replace the parts or products under normal using circumstances, from purchased time within one year caused by manufacturing bad and unable to work.

7.2. Attached electrodes do not belong to this warrant range. But, if the newly purchased electrode went wrong without using, it’s free of charge to maintain and replace.

7.3. The above warranty is not apply to defects resulting from action of user such as misuse, improper wiring, operation outside of specification, improper maintenance or repair, or unauthorized modification.

Table І  Meter Parameter Setting Table

| Mode | Prompt Mark | Parameter Setting Items | Code | Parameter |
| --- | --- | --- | --- | --- |
| pH | P1 | pH buffer solution series selection |  | USA (Europe & U.S.A. series) NIS (NIST series) CH (China series) |
|  | P2 | Purified water pH temperature compensation setting |  | OFF-On (shut-set) |
|  | P3 | Purified water added with ammonia pH temperature compensation setting |  | OFF-On (shut-set) |
|  | P4 | Temperature unit setting |  | ℃ ℉ |
|  | P5 | Back light display time setting |  | 0-1-3-6min |
|  | P6 | Auto power off time setting |  | 0-10-20min |
|  | P7 | Restore to producer setting |  | OFF-On (shut-set) |
| mV | P1 | Back light display time setting |  | 0-1-3-6min |
|  | P2 | Auto power off time setting |  | 0-10-20min |

Table Ⅱ Meter Restore to Producer Setting Table

| Mode | Prompt Mark | Parameter Setting Items | Producer Setting Content | Non Producer Setting Content Icon |
| --- | --- | --- | --- | --- |
| pH | P2 | Purified water pH temperature compensation setting | OFF |  |
|  | P3 | Purified water added with ammonia pH temperature compensation setting | OFF |  |

Tabel Ⅲ Code Icon and Abbreviation Table

Shanghai San-Xin Instrumentation, Inc.

Add: 3/F. Building No.4, No. 471 Guiping Road, Shanghai, China. 200233

Tel: +86-21-63362480                   Fax: +86-21-64956880

Website: www.shsan-xin.com            E-mail: wxmab@shsan-xin.com

# Text boxes

### End

### Pure-2

Purified water added with ammonia pH temperature compensation setting

### Pure-1

Purified water pH temperature compensation setting

### NIST

NIST series standard

### USA

Europe & U.S.A series standard

### China

China series standard

### Solution

### Standard solution

3

4

5

5

6

7

8

8

8

9

10

13

15

16

16

16

17

18

18

### Calibration

### Code and Abbreviation

### English

Explanation

1unit

1pcs

1pcs

1btl/each

1pcs

1pcs

2pcs

1pcs

1pcs