##### Model SX723

pH/mV/Conductivity Meter

Operation Manual

Shanghai San-Xin Instrumentation Inc.

SX700 Series of Portable Meters

1. Model SX711 pH / mV Meter

2. Model SX712 ORP Meter

3. Model SX713 Conductivity / TDS / Salinity / Resistivity Meter

4. Model SX716 DO Meter

5. Model SX721 pH / ORP Meter

6. Model SX723 pH /mV / Conductivity Meter

7. Model SX725 pH / mV / DO Meter

8. Model SX726 Conductivity / DO Meter

9. Model SX731 pH / ORP / Conductivity Meter

10. Model SX736 pH / mV / Conductivity / DO Meter

11. Model SX751 pH / ORP / Conductivity / DO Meter

Content

1. Summary ----------------------------------------------------------------------------------------------

2. Technical Parameters -------------------------------------------------------------------------------

3. Meter Overview --------------------------------------------------------------------------------------

3.1. LCD Display -------------------------------------------------------------------------------------

3.2. Operation Keys ----------------------------------------------------------------------------------

3.3. The Storing,Displaying and Clearing of the Measuring Information ---------------------

4. pH Measurement ------------------------------------------------------------------------------------

4.1. Preparation Work --------------------------------------------------------------------------------

4.2. Meter Calibration -------------------------------------------------------------------------------

4.3. Sample Test --------------------------------------------------------------------------------------

4.4. Parameter Setting --------------------------------------------------------------------------------

4.5. Considerations -----------------------------------------------------------------------------------

4.6. The Self-diagnose Information ----------------------------------------------------------------

5. mV Measurement -----------------------------------------------------------------------------------

5.1. Sample Test --------------------------------------------------------------------------------------

5.2. Parameter Setting --------------------------------------------------------------------------------

6. Conductivity Measurement ------------------------------------------------------------------------

6.1. Preparation Work --------------------------------------------------------------------------------

6.2. Meter Calibration ---------------------------------------------------------------------------

6.3. Sample Test --------------------------------------------------------------------------------------

6.4. Important Statement ----------------------------------------------------------------------------

6.5. Parameter Setting --------------------------------------------------------------------------------

6.6. Considerations -----------------------------------------------------------------------------------

7. Meter’s Complete Kit ------------------------------------------------------------------------------

8. Warranty ----------------------------------------------------------------------------------------------

TableⅠ Meter Parameter Setting Table

TableⅡ Meter Restore to Default Setting Table

Table Ⅲ Code Icon and Abbreviation Table

1.Summary:

Thanks for buying and using the model SX723 pH/mV/Conductivity Meter (the following called “meter” in short).

Before using this meter, please read the operation manual carefully in order to help use and maintain it correctly. On the basis of improving instrument of performance constantly, we reserve the right of changing the content of this manual and accessories in case of not notifying in advance.

This meter is a perfect combination with the most advanced electronic technology, sensor technology and software design The meter can measure the parameters of pH, mV, conductivity, TDS, salinity, resistivity and temperature for high accuracy water solution. It is the best portable water analysis meter with the highest performance and the lowest cost. It is suitable for the trade such as the mining industry, power plant, water treatment projects and environmental protection, etc., especially has more extensive application in the field and spot test.

Built-in microprocessor chip, beautiful appearance, multi-functional and easy to use, this meter has the following prominent features:

1.1.	Double parameters meter, with pH electrode, conductivity electrode, and with calibration solution and special carrying case, easy to use.

1.2.	Built-in microprocessor chip, with the intelligent functions of automatic calibration, automatic temperature compensation, data storage, function setting and automatic self-diagnose, auto power off and low voltage display etc.

1.3.	Adopts digital filter and step slipping technology to intelligently improve meter’s response speed and result accuracy. “    ” will appear when reading to be stable.

1.4.	Automatically recognize 13 kinds of pH standard buffer solution. User can choose anyone from three series of buffer solutions: Europe & U.S.A. series, NIST series, China series.

1.5.	High purified water and ammonia added purified water pH measuring mode can be set up, for this two special pH measuring modes, besides the general slope temperature compensation, also added function of nonlinearity solution temperature compensation to make measurement more accurate, especially suitable for the fields of electric power and petrochemical etc.

1.6.	Adopts advanced automatic frequency convertible and voltage regulation technology to enlarge the measuring range of conductivity electrode which constant K=1 to 10 times. It can meet measuring accuracy requirement for range within 0 to 200mS/cm by using only one-point calibration. It is the particular one–point calibration function for this meter.

1.7.	Automatically recognize 8 kinds of conductivity standard solution. User can choose anyone from two series of standard solutions: Europe & USA series and CHINA series.

1.8.	Automatically select conductivity measuring range, the meter will do non-linearity temperature compensation test for purified and high purified water less than 10uS/cm, which greatly improve the measuring accuracy. It is especially suitable for use in fields of electric power, microelectronics and medicine.

1.9. Meter’s circuit board adopts SMT film-covering technology to improve meter’s production reliability.

1.10. Meter has the back light LCD display monitor.

1.11. Dustproof and waterproof meter meets IP57 rating.

2. Technical Parameters:

2.1. pH:

| Measuring range | (-2.00 to 19.99) pH |
| --- | --- |
| Resolution | 0.1/0.01 pH |
| Accuracy | Meter:±0.01pH; Overall: ±0.02pH |
| Input current | ≤2×10-12 A |
| Input impendance | ≥1×1012 Ω |
| Stability | ±0.01 pH/3h |
| Temp. compensation range | (0 to 100) ℃ (automatic) |

2.2. mV:

| Measuring range | -1999 mV to 0 to 1999mV |
| --- | --- |
| Resolution | 1mV |
| Accuracy | ±0.1% FS |

2.3. Conductivity:

| Measuring range | Conductivity: (0.00 to 19.99)μS/cm (20.0 to 199.9)μS/cm (200 to 1999)μS/cm (2.00 to 19.99) mS/cm (20.0 to 199.9) mS/cm TDS: (0 to 100) g/L Salinity : (0 to 100) ppt Resistivity:(0 to 100) MΩ·cm |
| --- | --- |
| Resolution | 0.01/0.1/1μS/cm 0.01/0.1 mS/cm |
| Accuracy | Meter:±1.0% FS, Overall:±1.5% FS |
| Temp. compensation range | (0 to 50) ℃ (automatic) |
| Electrode constant | 0.1 / 1 / 10 cm-1 |
| Reference temperature | 25℃, 20℃and 18℃ |

2.4. Other Technical Parameters:

| Data storage | 300 groups |
| --- | --- |
| Storage content | Series number, measuring value measuring unit and temperature, |
| Power | Two AA batteries (1.5V x2) |
| Size and weight | Meter:(65 x 120x 31) mm/180g; Carrying case: (360 x 270x 76) mm/1.7kg |
| Quality and safety certification | ISO9001:2000, CE and CMC |

2.5. Working Condition:

| Environment temperature | 5 to 35 ℃ (0.01 grade) |
| --- | --- |
| Environmental humidity | ≤85% |
| IP rating | IP57 Dustproof and waterproof |

3. Meter Overview:

3.1. LCD Display:

① —— Parameters mode icon

② —— Measuring value

③ —— Serial number and icon as measuring value to be stored and Storinged. and the indication icons for the special measuring state

M+ — measuring value to be stored icon;

RM — measuring value to be displayed icon;

④ —— Measuring unit

⑤ —— Temperature measuring value and unit

⑥ —— Measuring value to be stable icon

⑦ —— Electrode calibration indicating icon

⑧ —— Indication icon of low voltage, appearing when the voltage is less than 2.6V, calling attention to change the batteries.

3.2. Operation Keys:

The meter has 5 operation keys in all.

3.2.1.      — Switch on/off key

3.2.2.      — Calibration key

(a) When in the measurement state, press this key to enter into calibration mode.

(b) When in the parameter setting state, press this key to change the number or the ON/OFF state.

3.2.3.      — Function key

(a) In pH measuring mode,short-time press (<1.5s) to change the measuring parameters, the meter will display      →      .Long-time press (>2s) to enter into the parameter setting mode, and it will display P1,P2, P3…cyclically.

(b) In conductivity measuring mode,press this key to enter into parameters setting mode, and it will display P1,P2, P3…cyclically without releasing this key.

3.2.4.      — Back light and entrance key

(a) When in the measuring state, short press-time (<1.5s) to turn on or off the back light display;

(b) When in the calibration state or the parameter setting state, press this key to confirm, and then the meter enters into measuring state;

(c) When in      mode, press the key without releasing to change pH resolution: 0.01→0.1 pH cyclically, then release after choosing made;

(d) When in      mode, press the key to cyclically change parameter mode:

(total dissolved solid)→      (salinity) →      (resistivity)→     (conductivity), and then release the key when choosing made.

3.2.5.      — The combination key of memory and Storing

(a) When in the measurement state, short-time press (<1.5s) to store the measuring data, long-time press (> 2s) to display the stored measuring data.

(b) When in the parameter setting state, press this key to change the number or ON/OFF state.

3.3. The Storing and Clearing of the Measuring Information:

3.3.1. Store the measuring information:

In the measuring mode, when the measuring data is stable and appears the “    ”, short-time press      key(＜1.5s), LCD will display “ M+ ” icon and stored serial number, and meanwhile will store all the measuring information. Meter can separately store 100 groups of measuring information in the mode of pH, mV, conductivity, and it can totally store 300 groups.

3.3.2. Store measuring information:

(a) Under the measuring mode, long-time press the      key(>2s), meter will display the last stored information, and the stored number and “RM” icon will appear in the upper right corner of the LCD, and the complete measuring information. Again press

or      key, meter will in turn display all the measuring information, press

or      key without releasing can make a quick query;

(b) In the displaying mode (there are “ RM ” and stored serial number in the upper right corner of the LCD), press      key to return to the measuring mode.

3.3.3. Clear the stored measuring information:

In the displaying mode, press the      key for 5 seconds, LCD will appear “      ” for 2 seconds. It means the storage has been cleard, and meter returns to measuring mode.

4. pH Measurement:

4.1. Preparation Work:

Press      key to turn on the meter. Plug the model 201T-S pH/ATC three in one combination pH electrode into meter’s socket, and the meter enters into pH measuring mode automatically.

4.2. Meter Calibration:

4.2.1. Press       to enter into calibration mode, LCD displays the flickering “     ” indicating to enter into the first point calibration.

4.2.2. Wash the pH electrode in purified water and shake off water, then insert it into the pH7.00 buffer solution, stir the electrode and then let it still, waiting for the data stable and appearing “     ”, then again press key      , the LCD will appear a flickering 7.00 pH, calibration finishes after several seconds and appears a stable pH value and a flickering “    ”, indicating the first point calibration has been finished and the meter enters into the second point calibration.

4.2.3. Wash the pH electrode in purified water and shake off water, then insert it into the pH4.00 buffer solution, stir the electrode and then let it still, waiting for a stable data and  “      ” appearing, then again press key       , the LCD will appear a flickering 4.00pH, calibration finishes after several seconds, LCD will appear a stable pH value and a flickering “    ”, indicating the second point calibration has been finished and meter enters into the third point calibration.

4.2.4. Wash the pH electrode in purified water and shake off water, then insert it into the pH10.01 buffer solution, stir the electrode and let it still, waiting for the data stable and

appearing, then again press key      , LCD will appear a

flickering 10.01 pH, calibration finishes after several seconds, appear a stable pH value and “            ” three calibration indication icon, see picture (4-1), indicates the three-point calibration has been finished and meter entered into measuring mode.

4.2.5. Note:

(a) This meter can adopt random one-point, two-point or three-point automatic calibration, after the first point calibration (see item 4.2.2.), press      to confirm one-point calibration and enters into measuring mode. The indication icon “     ” for one-point calibration will appear on the lower left corner of LCD. When the measuring accuracy is≥±0.1pH, user just need to choose one kind buffer solution to take one-point calibration according to the measuring range.

(b) After the second point calibration, (see item 4.2.3.), press      to confirm two-point calibration and enter into measuring mode. The indication icon “        ” for tow-point calibration will appear on the lower left corner of LCD. User can choose pH4.00 and pH7.00 to calibrate if the measurement is just within the acidity range and choose pH7.00 and pH10.01 to calibrate if just within the alkalinity range.

(c) User should choose three-point calibration so as to reach a more accurate measurement if the measuring range is wide, or if the electrode has been used for long or has ageing phenomenon. As to the new electrode which be used for the first time, it must be calibrated by three-point calibration to keep the unanimity of the meter slope adjustment with the pH electrode.

4.3. Sample Test:

Insert pH electrode into the sample solution after washing and shake off water, stir the electrode in the sample solution and still, when the LCD appears the icon “     ”, the displaying stable value is the right one.

Note: According to the pH equal temperature measuring theory: the closer the temperature of the sample solution with the calibration solution, the more accurate the measuring value will be acquired. So please comply with this theory.

4.4. Parameter Setting:

4.4.1. pH measuring parameter setting table (Table (4-1))

Table (4-1)

| Prompt Mark | Parameter Setting Items | Code | Parameters |
| --- | --- | --- | --- |
| P1 | pH buffer solution series selection |  | USA (Europe & U.S.A series) NIS (NIST series) CH (China series) |
| P2 | Purified water pH temperature compensation setting |  | OFF-On (OFF-set) |
| P3 | Ammonia added purified water pH temperature compensation setting |  | OFF-On (OFF-set) |
| P4 | Temperature unit selection |  | ℃ ℉ |
| P5 | Back light display time setting |  | 0-1-3-6min |
| P6 | Auto power off setting |  | 0-10-20min |
| P7 | Restore to default setting |  | OFF-On ( OFF-set) |

4.4.2. pH buffer solution series selection (P1)

(a) Press        key(>2s) until entering into P1 mode: see picture (4-2).

(b) Press      or      key to choose buffer solution series:

(Europe & U.S.A series) — 1.68, 4.00, 7.00, 10.01and 12.45 pH

(NIST series) — 1.68, 4.01, 6.86, 9.18 and 12.45 pH

(China series) — 1.68, 4.00, 6.86, 9.18 and 12.46 pH

(c) Press      key to enter into next parameter setting or press      key to conform and return to measuring mode.

4.4.3. Purified water pH temperature compensation setting (P2)

(a) Press      key in the mode P1 to enter into mode P2, see picture (4-3).

(b) Press      or      to choose “     ” (purified water pH temperature compensation setting) or “      ” (OFF).

(c) Press      key to enter into next parameter setting or press      key to conform and return to measuring mode.

(d) The default setting is “ OFF ”.

Note: There will appear“       ” icon in the upper right corner of the LCD if purified water temperature compensation function was set.

4.4.4. Ammonia added purified water pH temperature compensation setting (P3)

(a) Press      key in mode P2 to enter into mode P3, see picture (4-4).

(b) Press      or      key to choose “    ” (ammonia added purified water pH temperature compensation setting) or “      ” (OFF).

(c) Press      key to enter into next parameter setting or press

key to conform and return to measuring mode.

(d) The default setting is “ OFF ”.

Note: If set the ammonia purified water pH temperature compensation function, there will appear“       ”icon in the right upper corner of the LCD when in the measuring mode.

4.4.5. Temperature unit ℃/℉ setting (P4)

(a) Press      key in mode P3 to enter into mode P4, see picture (4-5).

(b) Press      or      key to choose temperature unit: ℃or℉.

(c) Press      key to enter into next parameter setting or press      key to conform and to return to measuring mode.

4.4.6. Back light display time setting (P5)

(a) Press      key in mode P4 to enter into mode P5, see picture (4-6).

(b) Press      or      key to choose the time of back light auto power off: 0 min, 1 min, 3 min or 6 min. The back light display function will be closed if choosing a 0min.

(c) Press      key to enter into next parameter setting or press      key to conform and return to measuring mode.

(d) The default setting for P5 is 1min.

4.4.7.Auto power off time setting (P6)

(a) Press      key in mode P5 to enter into mode P6, see picture (4-7).

(b) Press      or      key to choose the time: 0min, 10min or 20min. The auto power off function will be closed if choosing 0min.

(c) Press      key to enter into next parameter setting or press      key to conform and return to measuring mode.

4.4.8. Restore to default setting (P7)

(a) Press      key in mode P6 to enter into mode P7, see picture (4-8).

(b) Press      or      key to choose “     ”, means that all the pH measuring parameters have been restored to the default setting mode, and restore to measuring mode after 2 seconds.

4.5. Considerations:

4.5.1.Calibration times of meter rely on the sample, electrode performance and required accuracy. For high accurate measurement (≤±0.02pH), which should be calibrated immediately with high accurate standard buffer solution, for general accuracy measuring (≥±0.1pH), which can be used almost one week or long time once be calibrated one time.

4.5.2. The meter must be recalibrated in the following situations:

(a) New changed or unused electrode for a long time;

(b) After measuring acid (pH<2) or alkaline (pH>12) solution;

(c) After measuring solution which contains fluoride and concentrated organic solution;

(d) The solution’s temperature is much different with calibration temperature.

4.5.3. The soaking solution in the protecting bottle of front pH electrode is to keep the glass bulb and junction activating. Loose the capsule, pull out the electrode and wash it in purified water before measuring. Insert the electrode and screw tight the capsule after measuring to prevent the solution leaking. If the soaking solution is turbid or moldy, please wash and change a new one at once.

4.5.4. The preparation of the soaking solution: take 25g analytical pure KCL, dissolved with purified water and dilute to 100mL. Electrode should avoid soaking in purified water protein solution and acid fluoride solution for a long time as well avoid getting touching with organic silicon lipidic matters.

4.5.5. To calibrate the meter with the given value pH buffer solution, the pH value of the standard buffer solution must be reliable so as to improve the accuracy. Buffer solution should be changed in time after many times using.

4.5.6. Always keep the meter clean and dry, especially for the socket of meter and electrode, otherwise it may lead to an inaccurate measurement or invalidity. To clean and dry them with medical cotton with dehydrated alcohol if there are any dirty.

4.5.7. The sensitive glass bulb in the front of combination electrode should not touch with hard things, any broken and rough will make the electrode invalidity. Before and after measuring, the electrode should be washed with purified water, and dry electrode after washing, don’t clean glass bulb with tissue for it will effect stability of electrode potential and enlarge response time. The electrode should be washed many times for removing the sample stuck on the electrode, or wash with suitable solvent then clean the solvent with purified water after measuring sticky sample.

4.5.8. An electrode be used for a long time, or measured solution which contains a polluting solute easily for the sensitive bulb, or a substance resulting in jam at the junction, the electrode will be getting passivated, its sensitivity will decrease and its response is getting slow, the reading are not correct. It could adopt the following method for various cases:

(a) The glass bulb is contaminated and aging: Put the electrode into 0.1mol/L dilute hydrochloric acid (Preparation: diluted 9mL hydrochloric acid to 1000mL with purified water) for 24h. Rinse it with purified water, then dipped it into the electrode dipping solution for 24h. If the passivation is serious, then user can also put the bulb of electrode into 4% HF (hydrofluoric acid) or the electrode activation solution for 3 to 5 seconds, rinsing it with purified water, and dipped it in the electrode soaking solution for 24h to renew it.

(b) Wash for contaminated glass bulb and junction: (For reference)

Contamination                            Abluent

Inorganic metal oxide                 diluted acid less than 1mol/L

Organic lipidic matter                 dilute washing (weak alkaline)

Resin macromolecule matter              dilute alcohol, acetone, ether

Proteinic haematocyte sediment     Acidic enzymatic solution (such as dried yeast)

Kinds of paint                      dilute bleacher, peroxide

4.5.9. pH electrode using period is about 1 year, but its life will be shortened if using condition is poor or incorrect maintenance. So it should be replaced immediately after electrode become aging or invalid.

4.5.10. When it appears an abnormal reading when calibration or displaying, please set P7 as “ON” to restore the meter to default setting mode, and then to calibrate and measure again.

4.6. The Self-diagnose Information:

When using, there might appear the following icons. This is the meter’s self-diagnose information, which can help to know some information about the meter or the electrode when using:

4.6.1. The stable icon -2.00 pH or the 19.99 pH — this icon appeared when the value has surpassed the measuring range. There will also appear such signs when the electrode is not well connected with the meter or when the electrode is not insertd into the solution. This is a normal phenomenon.

4.6.2. “       ” — Electrode zero potential to be exceeded  (<-60mV or >60mV)

4.6.3. “       ” — Electrode slope to be exceeded (< 85% or >105%)

When appear “       ” or “       ”, the meter can not work, please take the following check:

(a) Check if the electrode bulb has air bubble, if has, please shake it hardly.

(b) Check the quality of buffer solution, if it goes bad or the value has biggish error.

(c) Set the meter to default setting mode (for details see P14 item 4.4.8.), then recalibrate it.

If still can not recover the normal state after doing the above checking, please replace a new pH electrode.

5. mV Measurement:

5.1. Sample Test:

5.1.1. In pH measuring mode, press       to display     ;

5.1.2. In other measuring modes, plug the ORP combination electrode(Sold separately) into meter’s corrsponding socket, then the meter will automatically enter into mV measuring mode, insert the probe into the sample solution, slowly stir and then let it still. When “    ” appears, the stable reading is the ORP value or the potential of the ion electrode.

5.2. Parameter Setting:

5.2.1. mV measurement parameter setting table (Table (5-1))

Table (5-1)

| Prompt Mark | Parameter Setting Items | Code | Parameters |
| --- | --- | --- | --- |
| P1 | Back light display time setting |  | 0 -1-3-6 min |
| P2 | Auto power off time setting |  | 0 -10-20 min |

5.2.2. Back light display time setting (P1)

Press the      key until entering into mode P1, see detailed information in P13 item4.4.6.

5.2.3. Auto power off time setting (P2):

Press the      key in mode P1 to enter into mode P2, see detailed operations in P14 item 4.4.7

6. Conductivity Measurement:

6.1. Preparation Work:

6.1.1. Press      key to turn on the meter; Plug the 2301-S conductivity electrode into the meter’s corresponding socket, then the meter will automatically enter into conductivity measuring mode.

6.1.2. Press      key without releasing it to cyclically display

(Total Dissolved Solid) →      (Salinity) →      (Resistivity) →

(Conductivity), then release the key after choosing made.

6.2. Meter’s Calibration:

Press      key, “      ” is flickering on the LCD, indicating to enter into the calibration mode, wash the conductivity electrode with purified water and shake off water, then insert it into the 1413μS/cm calibration solution, still it after stirring, when “    ” icon appears, meaning the display values is stable and again press      key, LCD will display flickering “1413μS/cm ”, several seconds later the “      ” icon will appear and meter returns to the measuring mode, at this time, LCD will display stable measuring value 1413μS/cm and calibration icon “     ”, indicating the calibration has been finished. If measuring value is unstable, user can repeat calibrating steps until it is stable.

Note: (1) The meter has been calibrated before leaving factory, generally users can use it directly.

(2) Meter only can be calibrated under the mode of      , and can not be calibrated in mode of      ,      and      , must shift to      mode, and again shift to the original mode after finishing calibration.

6.3. Sample Test:

6.3.1. Wash the conductivity electrode and shake off the water on it, then insert it into the solution, let it still after stirring, and take the reading after “    ” appearing, which is the conductivity values of this solution.

6.3.2 Press      key without releasing to display cyclically the value of the TDS, salinity or resistivity value against the conductivity value.

6.4. Important Statement:

6.4.1. This meter build-in below two kinds of calibration solution series, please set up in the parameters P1;

(a)      (Europe & U.S.A. series) — 84μS/cm, 1413μS/cm, 12.88 mS/cm and 111.9 mS/cm

(b)      (China series) — 146.6μS/cm, 1408μS/cm, 12.85mS/cm and 111.3 mS/cm

6.4.2. This meter has an unique one-point calibration function, to choose the calibration solution based on the principle of the water samples and calibration solution conductivity as close as possible, in general the most common calibration solution is 1413 μS/cm. Use the equipped 2301-S conductivity electrode (K = 1 cm-1), and do calibrate with 1413 μS/cm calibration solution, can use within the range of less than 100 mS/cm. Please choose refer to the table (6-1).

Note：When you choose to use DJS-0.1-S conductivity electrode (optional accessories, K=0.1 cm-1 , no temperature element), you can install the meter Manual Temperature Compensation model: long time press       (press time more than 2s), at this time, the value of temperature is twinkling, press       or

to change the value of temperature and then press       is ok.

Table (6-1)

| Measuring range | 0.05 to 20μS/cm | 0.5μS/cm to 200mS/cm |
| --- | --- | --- |
| Electrode constant | K=0.1cm-1(flow test) | K=1.0cm-1 |
| Calibration solution | 84μS/cm | 84μS/cm | 1413μS/cm | 12.88 mS/cm 111.9 mS/cm |
| Calibration indication icon |  |  |  |  |

6.4.3. There are two kinds of method which are standard solution calibration method and constant calibration method for conductivity electrode calibration which have been set in meter, the statement in item 6.2. “electrode calibration” is the standard calibration method, and when the accuracy of standard solution is accurate then suggest the standard solution calibration method is priority selected, it can ensure the best accuracy. If user used to use constant calibration method, please set the electrode constant in the parameter set of P5. (see P24 item6.5.6.). The two methods can choose freely and won’t influence each other. When adopting the constant calibration method, LCD won’t display the calibration indicator symbol.

6.4.4. The temperature compensation coefficient of the meter setting is 2.00%/℃ by default. However, the conductivity temperature coefficient is different for solution of different variety and concentrations, the user can refer chart (6-2), as well as the data which they own get in the experiment, to set in the parameters setting of P4. At the same time, meter will do automatic non-linear temperature compensation in the high purified water which less than 10 μ S / cm.

Note: When the coefficient of temperature compensation is set to be 0.00, that is, there is no temperature compensation when testing, the measuring value based on current temperature.

Table (6-2)

| Solution | Temperature compensation coefficient |
| --- | --- |
| NaCl salt solution | 2.12%/℃ |
| 5%NaOH solution | 1.72%/℃ |
| Dilute ammonia solution | 1.88%/℃ |
| 10% hydrochroric acid solution | 1.32%/℃ |
| 5% sulfuric acid solution | 0.96%/℃ |

6.4.5.The other parameters setting, please see the table (6-3)

6.4.6. Special attention: when the content of parameter setting is consistent with content of default setting, conductivity measuring interface will display as picture (6-1), the upper right corner of LCD will display time, when the content of parameter setting is different with content of default setting, conductivity measuring interface will display as picture (6-2), the code icon will appear on the upper right of LCD, when parameters with code icon to be set more than two, LCD just display one of the code symbols, user should enter into the parameter setting interface and carefully check the contents of setting or restore the meter to default setting and again set needed parameters.

6.5. Parameter Setting:

6.5.1. Conductivity measuring parameter setting table (table (6-3))

Table (6-3)

| Prompt Mark | Parameter Setting Items | Code | Parameter |
| --- | --- | --- | --- |
| P1 | Standard solution series selection |  | USA (84μS/cm, 1413μS/cm, 12.88 mS/cm, 111.9 mS/cm) CH (146.6μS/cm, 1408μS/cm, 12.85mS/cm, 111.3mS/cm) |
| P2 | Electrode constant selection |  | 0.1, 1, 10 |
| P3 | Benchmark temperature selection |  | 25℃ 20℃ 18℃ |
| P4 | Temperature compensation coefficient setting |  | 0.00 to 9.99% |
| P5 | Electrode constant setting |  |  |
| P6 | Temperature unit setting |  | ℃ ℉ |
| P7 | Back light display time setting |  | 0-1-3-6min |
| P8 | Auto power off time setting |  | 0-10-20min |
| P9 | Restore to default setting |  | OFF-On |

6.5.2. Conductivity calibration solution series selection (P1)

(a) Press      key until entering into P1 mode, see picture (6-3);

(b) Press      or      key to choose the standard solution series:

— Europe & U.S.A. series      — China series

(c) Press      key to enter into the next parameter setting or press      key return to measuring mode.

6.5.3. Electrode constant selection (P2)

(a) Under P1 mode,press the      key to enter into P2 mode, see picture (6-4);

(b) Press      or      key to change the constant setting: 0.1→1.0→10.0.

(c) Press      key to enter into the next parameter setting or press      key return to measuring mode.

(d) The default setting of P2 is K=1.0.

6.5.4. Benchmark temperature selection (P3)

(a) Under P2 mode, press      key to enter into the P3 mode, see picture (6-5);

(b) Press     or     key to choose the benchmark temperature: 25℃→20℃→18℃;

(c) Press      key to enter into the next parameter setting or press      key return to measuring mode.

(d) The default setting of P3 is 25℃.

Attention: If change the benchmark temperature to 20℃ or 18℃, the LCD will display“      ”icon in the upper-right corner when in measuring mode.

6.5.5. Temperature compensation coefficient setting (P4)

(a) Under P3 mode, press      key to enter into P4 mode, shown as the picture (6-6);

(b) Press      or      key to change the number, the changing range of data is from 0.00 to 9.99; Press without releasing to quickly change it.

Attention: When the number set as 0.00, means there is no temperature compensation, for details please see P21 item 6.4.4;

(c) Press      key to enter into the next parameter setting or press      key return to measuring mode.

(d) The default setting of P4 is 2.00%.

Attention: If change the temperature coefficient to other numerical value, upper-right corner of LCD will display“      ”icon when in measuring mode.

6.5.6. Electrode constant setting (P5), please refer to P22 item 6.4.3.

(a) Under P4 mode, press      key to enter into P5 mode, LCD displays last calibration constant. For example the 0.96cm-1 shown as picture (6-7);

(b) Press      or      key to change number when the number is flickering, and finish setting the constant which marked on the housing of conductivity electrode.

(c) Press      key to enter into the next parameter setting or press      key return to measuring mode.

(d) If need to do calibration for other conductivity electrode which constant is not 1, such as conductivity electrode with constant 10.3, should enter into constant setting P2 to set the constant to be “10”, then enter into P5 mode and set the constant to be 10.3.

6.5.7. Temperature unit ℃/℉ setting (P6)

Under P5 mode, press      key to enter into P6 mode, other operation please refer to P13 item4.4.5.

6.5.8. Back light display time setting (P7)

Under P6 mode, press      key to enter into P7 mode, other operation please refer to P13 item 4.4.6.

6.5.9. Auto power off time setting (P8)

Under P7 mode, press      key to enter into P8 mode, other operation please refer to P14 item 4.4.7.

6.5.10. Restore to default setting (P9)

Under P8 mode, press      key to enter into P9 mode, other operation please refer to P14 item 4.4.8.

6.6. Considerations:

6.6.1. The conductivity electrode has been calibrated before on the market, user can directly use it under normal circumstance;

6.6.2. Recommend calibrating one time every month under the normal circumstances; It is necessary to calibrate the newly purchased conductivity electrode or used one after a period time;

6.6.3. Keeping the conductivity electrode clean and wash it with purified water, then throw off the water on it before and after testing. It’s better to rinse electrodes with sample solution;

6.6.4. The surface of 2301-S conductivity electrode is plated with a layer of metal platinum black to used as lowering the electrode polarization, enlarge measuring range, so can not polish the surface of platinum black, only can wash it in water by shaking, in case to damage the platinum black coating; for organic stained things can be washed with hot water with detergent or alcohol.

6.6.5. Conductivity electrode before use can be immersed in water, to prevent the platinum black to be passivated, if found the platinum black plated electrode is invalid, can immerse it into 10% nitric acid solution or 10% hydrochloric acid for 2 minutes, then rinse with purified water and again test. If the situation does not improve, then need to replace platinum black, or replacing a new conductivity electrode.

6.6.6. When it appears an abnormal reading when calibrating or displaying, please set P9 as “ON” to make the meter restore to default setting mode, and then do calibrate and measure again.

7. Meter’s Complete Kit:

7.1. Model SX723 pH/mV/Conductivity meter

7.2. 201T-S plastic pH/ATC three-in-one combination electrode

7.3. 2301-S plastic conductivity electrode

7.4. pH standard buffer solution

(pH4.00, pH7.00 and pH10.01/50mL)

7.5. 1413μS/cm conductivity standard solution (50mL)

7.6. Screw driver

7.7. Spare AA batteries

7.8. Operation manual

7.9. Carrying case

8. Warranty:

8.1. We warrant this meter to be free of charge maintain, replace the parts or products under normal using circumstances, from purchased time within one year caused by manufacturing bad and unable to work.

8.2. Attached electrodes do not belong to this warrant range. But, if the newly purchased electrode went wrong without using, it’s free of charge to maintain and replace.

8.3. The above warranty is not apply to defects resulting from action of user such as misuse, improper wiring, operation outside of specification, improper maintenance or repair, or unauthorized modification.

Table І Meter Parameter Setting Table

| Mode | Prompt Mark | Parameter Setting Items | Code | Parameter |
| --- | --- | --- | --- | --- |
| pH | P1 | pH buffer solution series selection |  | USA (Europe & U.S.A. series) NIS (NIST series) CH (China series) |
|  | P2 | Purified water pH temperature compensation setting |  | OFF-On (OFF-set) |
|  | P3 | Purified water added with ammonia pH temperature compensation setting |  | OFF-On (OFF-set) |
|  | P4 | Temperature unit selection |  | ℃ ℉ |
|  | P5 | Back light display time setting |  | 0-1-3-6min |
|  | P6 | Auto power off time setting |  | 0-10-20min |
|  | P7 | Restore to default setting |  | OFF-On (OFF-set) |
| mV | P1 | Back light display time setting |  | 0-1-3-6min |
|  | P2 | Auto power off time setting |  | 0-10-20min |
| Conductivity | P1 | Standard solution series selection |  | USA(84μS/cm, 1413μS/cm, 12.88 mS/cm, 111.9 mS/cm) CH(146.6μS/cm, 1408μS/cm, 12.85mS/cm, 111.3 mS/cm) |
|  | P2 | Electrode constant selection |  | 0.1, 1, 10 |
|  | P3 | Benchmark temperature selection |  | 25℃ 20℃ 18℃ |
|  | P4 | Temperature compensation coefficient setting |  | 0.00 to 9.99% |
|  | P5 | Electrode constant setting |  |  |
|  | P6 | Temperature unit setting |  | ℃ ℉ |
|  | P7 | Back light display time setting |  | 0-1-3-6min |
|  | P8 | Auto power off time setting |  | 0-10-20min |
|  | P9 | Restore to default setting |  | OFF-On (OFF-set) |

Table Ⅱ Meter Restore to Default Setting Table

| Mode | Prompt Mark | Parameter Setting Items | Default Setting Content | Non Default Setting Content Icon |
| --- | --- | --- | --- | --- |
| pH | P2 | Purified water pH temperature compensation setting | OFF |  |
|  | P3 | Purified water added with ammonia pH temperature compensation setting | OFF |  |
| Conductivity | P2 | Electrode constant type selection | K=1 |  |
|  | P3 | Benchmark temperature selection | 25℃ |  |
|  | P4 | Temperature compensation coefficient setting | 2.0% |  |

Table Ⅲ Code Icon and Abbreviation Table

Shanghai San-Xin Instrumentation, Inc.

Add: 3/F. Building No.4, No. 471 Guiping Road, Shanghai, China. 200233

Tel: +86-21-63362480                   Fax: +86-21-64956880

Website: www.shsan-xin.com            E-mail: wxmab@shsan-xin.com

# Text boxes

3

5

6

6

7

8

9

9

9

11

11

14

17

18

18

18

18

18

19

19

20

22

25

27

27

1unit

1pc

1pc

1btl/each

1bottle

1pc

2pcs

1pc

1pc

### Constant Calibration

### Constant setting

### Temperature Reference

Benchmark temperature setting

### Calibration

### End

### Constant

### Electrode constant setting

Temperature compensation coefficient setting

Temperature compensation coefficient setting

### Pure2

Purified water added with ammonia pH temperature compensation setting

### Pure1

Purified water pH temperature compensation setting

### NIST

NIST series standard

### USA

Europe & U.S.A series standard

### China

China series standard

### Solution

### Standard solution

### Salinity

### Total Dissolved Solid

### Resistivity

### Conductivity

### Code and Abbreviation

### English

Explanation