Model SX725

pH/mV/DO Meter

#### Operation Manual

Shanghai San-Xin Instrumentation Inc.

SX700 Series of Portable Meters

1. Model SX711 pH / mV Meter

2. Model SX712 ORP Meter

3. Model SX713 Conductivity / TDS / Salinity / Resistivity Meter

4. Model SX716 DO Meter

5. Model SX721 pH / ORP Meter

6. Model SX723 pH /mV / Conductivity Meter

7. Model SX725 pH / mV / DO Meter

8. Model SX726 Conductivity / DO Meter

9. Model SX731 pH / ORP / Conductivity Meter

10. Model SX736 pH / mV / Conductivity / DO Meter

11. Model SX751 pH / ORP / Conductivity / DO Meter

Content

1. Summary ------------------------------------------------------------------------------------------------

2. Technical Parameters ---------------------------------------------------------------------------------

3. Meter Overview --------------------------------------------------------------------------------------

3.1. LCD Display -------------------------------------------------------------------------------------

3.2. Operation Keys -----------------------------------------------------------------------------------

3.3. The Storage, Display and Clearing of the Measuring Information ---------------------

4. pH Measurement -------------------------------------------------------------------------------------

4.1. Preparation Work --------------------------------------------------------------------------------

4.2. Meter Calibration --------------------------------------------------------------------------------

4.3. Sample Test ---------------------------------------------------------------------------------------

4.4. Parameter Setting --------------------------------------------------------------------------------

4.5. Considerations -----------------------------------------------------------------------------------

4.6. The Self-diagnose Information ----------------------------------------------------------------

5. mV Measurement -------------------------------------------------------------------------------------

5.1. Sample Test ---------------------------------------------------------------------------------------

5.2. Parameter Setting --------------------------------------------------------------------------------

6. DO Measurement -------------------------------------------------------------------------------------

6.1. Preparation Work --------------------------------------------------------------------------------

6.2. Meter Calibration --------------------------------------------------------------------------------

6.3. Water Sample Test -------------------------------------------------------------------------------

6.4. Considerations -----------------------------------------------------------------------------------

6.5. Parameter Setting --------------------------------------------------------------------------------

6.6. Change the Membrane Cap --------------------------------------------------------------------

6.7. Zero Oxygen Calibration -----------------------------------------------------------------------

6.8. Salinity Calibration ------------------------------------------------------------------------------

6.9 Barometric Pressure Setting --------------------------------------------------------------------

7. Meter’s Complete Kit --------------------------------------------------------------------------------

8. Warranty ------------------------------------------------------------------------------------------------

TableⅠ Meter Parameter Setting Table

TableⅡ Meter Restore to Default Setting Table

Table Ⅲ Code Icon and Abbreviation Table

TableⅣ Content of Saturated Oxygen in Water at Different Temperature

TableⅤ Content of Saturated Oxygen at Different Barometric pressure

TableⅥ Content of Saturated Oxygen at Different Altitude

1. Summary:

Thanks for buying and using the model SX725 pH/mV/DO Meter (the following called “meter” in short).

Before using this meter, please read the operation manual carefully in order to help use and maintain it correctly. On the basis of improving instrument of performance constantly, we reserve the right of changing the content of this manual and accessories in case of not notifying in advance.

This meter is a perfect combination with the most advanced electronic technology, sensor technology and software design. The meter can measure the parameters of pH, mV, DO and temperature for high accuracy solution. It is the best portable water analysis meter with the highest performance and the lowest cost. It is suitable for the trade such as the mining industry, power plant, water treatment projects and environmental protection, etc., especially has more extensive application in the field and spot test.

Built-in microprocessor chip, beautiful appearance, multi-functional and easy to use, this meter has the following prominent features:

1.1.	Double parameters meter, with pH and DO electrode, also with calibration solution and special carrying case, easy to use.

1.2. Built-in microprocessor chip, with the intelligent functions of automatic calibration, automatic temperature compensation, data storage, function setting and automatic self-diagnose, auto power off and low voltage display etc.

1.3.	Adopts digital filter and step slipping technology to intelligently improve meter’s response speed and result accuracy. “    ” will appear when reading to be stable.

1.4.	It automatically recognizes 13 kinds of pH standard buffer solution. User can choose anyone from three series of buffer solutions: Europe & U.S.A. series, NIST series and China series.

1.5.	Purified water and ammonia added purified water pH measuring mode can be set up, for this two special pH measuring modes, besides the general slope compensation, also added function of nonlinearity temperature compensation for pH values of solution to make measurement more accurate, especially suitable for use in the fields of electric power and petrochemical etc.

1.6. Equipped with new type of DO electrode with temperature and salinity sensor which makes DO measuring mode has auto temperature compensation, auto salinity compensation and manual barometric pressure compensation. It’s easy to use and accurate to measure.

1.7. Polarogram type DO electrode has special electrode calibration cap. Only needs 3 to 5 minutes for electrode polarization. Electrode adopts combination diaphragm cap which very convenient to use. Each DO electrode equips with three spare diaphragms caps.

1.8. Meter’s circuit board adopts SMT film-covering technology to improve meter’s production reliability.

1.9. Meter has the blue back light LCD display monitor.

1.10. Dustproof and waterproof meter meets IP57 rating.

2. Technical Parameters:

2.1. pH:

| Measuring range | (-2.00 to 19.99) pH |
| --- | --- |
| Resolution | 0.1/0.01 pH |
| Accuracy | Meter: ±0.01pH, Overall: ±0.02pH |
| Input current | ≤2×10-12 A |
| Input impendance | ≥1×1012 Ω |
| Stability | ±0.01 pH/3h |
| Temp. compensation range | (0 to 100) ℃ (automatic) |

2.2. mV:

| Measuring range | -1999 mV to 0 to 1999mV |
| --- | --- |
| Resolution | 1mV |
| Accuracy | ±0.1% FS |

2.3. Dissolved Oxygen:

| Measuring range | (0 to 40.00) mg/L (ppm) (0 to 200.0) % |
| --- | --- |
| Resolution | 0.1/0.01 mg/L (ppm) 1/0.1% |
| Accuracy | Meter: ±0.10 mg/L, Overall: ±0.40 mg/L |
| Response time | ≤30 s (25℃, 90% response) |
| Residual current | ≤0.1 mg/L |
| Temp. compensation range | (0 to 45) ℃ (automatic) |
| Salinity compensation range | (0 to 45) ppt (automatic) |
| Barometric pressure compensation range | (66 to 200) kPa (manual) |
| Automatic calibration | Air saturated by water; water saturated by air |
| Electrode type | Polarogram type |

2.4. Other Technical Parameters:

| Data storage | 300 groups |
| --- | --- |
| Storage content | Series number, measuring value, measuring unit and temperature value |
| Power | Two AA batteries (1.5V x2) |
| Size and weight | Meter: (65×120×31)mm/180g; Carrying case:(360 x 270x 76)mm/1.5Kg |
| Quality and safety certification | ISO9001: 2000, CE and CMC |

2.5. Working Condition:

| Environment temperature | 5 to 35 ℃ (0.01 grade) |
| --- | --- |
| Environmental humidity | ≤85% |
| IP rating | IP57 Dustproof and waterproof |

3. Meter Overview:

3.1. LCD Display:

① —— Parameters mode icon

② —— Measuring value

③ —— Serial number and icon as measuring value to be stored and displayed; and indicate icon for special state.

M+ —measuring value to be stored icon;

RM —measuring value to be displayed icon;

④ —— Measuring unit

⑤ —— Temperature measuring value and unit

⑥ —— Measuring value to be stable icon

⑦ —— Electrode calibration indicate icon

⑧ —— Indication icon of low voltage, appears when the voltage less than 2.6V, call attention to change the batteries.

3.2. Operation Keys:

The meter has 5 operation keys in all.

3.2.1.      — Switch on/off key

3.2.2.      — Calibration key

(a) When in the measurement state, press this key to enter into calibration mode.

(b) When in the parameter setting state, press this key to change the number or the ON/OFF state.

3.2.3.      — Function key

(a) In pH measuring mode, press this key to change measuring parameters, the meter will display      →      ,cyclically. Press this key until displaying P1, indicating enter ing into the parameter setting mode, and it will display cyclically,P1,P2,P3…

(b) In      measuring mode, press this key to enter parameter setting mode, and it will  display cyclically,P1, P2, P3…

3.2.4.      — Back light and entrance key

(a) When in the measuring state, press to turn on or off the backlight display;

(b) When in the calibration state or in the parameter setting state, press this key to confirm, and then the meter will enter into measuring state;

(c) When in      mode, press the key to change pH resolution cyclically: 0.01→0.1 pH; then release after choosing made.

(e) When in      mode, press this key to change cyclically unit symbol: mg/L→ppm→%, and then release the key when choosing made.

3.2.5.      — The combination key of storing and displaying

(a) When in the measurement state, press this key to store the measuring data, long-time press (> 2s) to display stored measuring data.

(b) When in the parameter setting state, press this key to change the number or ON/OFF state.

3.3. The Storing, Displaying and Clearing of the Measuring Information:

3.3.1. Store the measuring information:

In the measuring mode, when the measuring data is stable and appears the “    ”, short-time press      key (<1.5s), LCD will display “ M+ ” icon and stored serial number, and meanwhile store all the measuring information. Meter can separately store 100 groups of measuring information in the mode of pH, mV and DO, it can totally store 300 groups.

3.3.2. Display measuring information:

(a) Under the measuring mode, long-time press the      key(>2s), meter will display the last stored information, and the stored number and “RM” icon and the complete measuring information will appear in the lower right corner of the LCD. Again press

or       key, meter will in turn display all the measuring information, press

or      key without releasing can make a quick query for measuring information under other serial number;

(b) In the displaying mode (there are “ RM ” and stored serial number in the upper right corner of the LCD), press      key to return to the measuring mode.

3.3.3. Clear the stored measuring information:

In the displaying mode, press the      key for 5 seconds, LCD will appear “      ” for 2 seconds. It means the memory has been cleared, and meter will return to measuring mode.

4. pH Measurement:

4.1. Preparation Work:

4.1.1. Press      key to turn on the meter. Plug the model 201T-S pH/ATC three-in-one combination pH electrode into meter’s socket, and the meter enters into pH measuring mode automatically.

4.2. Meter Calibration:

4.2.1. Press       to enter into calibration mode, LCD displays the flickering “    ” indicating to enter into the first point calibration.

4.2.2. Wash the pH electrode in purified water and shake off water, then insert it into the pH7.00 buffer solution, stir the electrode/probe and then let it still, waiting for the data stable and “     ” appearing, then again press key      , the LCD will appear a flickering 7.00 pH, calibration finishes after several seconds and appear a stable pH value and a flickering “    ”, indicating the first point calibration has been finished and the meter enters into the second point calibration.

4.2.3. Wash the pH electrode in purified water and shake off water, then insert it into the pH4.00 buffer solution, stir the electrode or the probe and then let it still, waiting for a stable data and “      ” appearing, then again press key         , the LCD will appear a flickering 4.00pH, calibration finishes after several seconds, LCD will appear a stable pH value and a flickering “    ”, indicating the second point calibration has been finished and the meter enters into the third point calibration.

4.2.4. Wash the pH electrode in purified water and shake off water, then insert it into the pH10.01 buffer solution, stir the electrode/ probe and still, waiting for the data stable and “    ”appearing, then again press key      , LCD will appear a flickering 10.01 pH, calibration finishes after several seconds, appears a stable pH value and “            ” three calibration indication icon, see picture (4-1), indicating the three-point calibration has been finished and entered into measuring mode.

4.2.5. Note:

(a) This meter can adopt random one-point, two-point or three-point automatic calibration, after the first point calibration (see item 4.2.2.), press      to confirm one-point calibration and enters into measuring mode. The indication icon “     ” for one-point calibration will appear on the lower left corner of LCD. When the measuring accuracy is≥±0.1pH, user just need to choose one kind buffer solution to take one-point calibration according to the measuring range.

(b) After the second point calibration, (see item 4.2.3.), press      to confirm two-point calibration and enter into measuring mode. The indication icon “        ” for tow-point calibration will appear on the lower left corner of LCD. User can choose pH4.00 and pH7.00 to calibrate if the measurement is just within the acidity range and choose pH7.00 and pH10.01 to calibrate if just within the alkalinity range.

(c) User should choose three-point calibration so as to reach a more accurate measurement if the measuring range is wide, or if the electrode has been used for long or has ageing phenomenon. As to the new electrode which be used for the first time, it must be calibrated by three-point calibration to keep the unanimity of the meter slope adjustment with the pH electrode.

4.3. Sample Test:

Immerge pH electrode into the sample solution after washing and dry it, rock the electrode holder and still, when the LCD appears the icon “     ” to take the reading after displaying value to be stable.

Note: According to the pH equal temperature measuring theory: the closer the temperature of the sample solution with the calibration solution, the more accurate the measuring value will be acquired. So please comply with this theory.

4.4. Parameter Setting:

4.4.1. pH measuring parameter setting table (Table (4-1))

Table (4-1)

| Prompt Mark | Parameter Setting Items | Code | Parameters |
| --- | --- | --- | --- |
| P1 | pH buffer solution series selection |  | USA(Europe & U.S.A series) NIS (NIST series) CH (China series) |
| P2 | Purified water pH temperature compensation setting |  | OFF-On (shut-set) |
| P3 | Ammonia added purified water pH temperature compensation setting |  | OFF-On (shut-set) |
| P4 | Temperature unit setting |  | ℃ ℉ |
| P5 | Backlight display time setting |  | 0-1-3-6 min |
| P6 | Time setting for auto power off setting |  | 0-10-20 min |
| P7 | Restore to default setting |  | OFF-On (shut-set) |

4.4.2. pH buffer solution series selection (P1)

(a) Press      key until entering into P1 mode: see picture (4-2).

(b) Press      or      key to choose buffer solution series:

( Europe & U.S.A series) —1.68, 4.00, 7.00, 10.01 and 12.45 pH

(NIST series) — 1.68, 4.01, 6.86, 9.18 and 12.45 pH

(China series) — 1.68, 4.00, 6.86, 9.18 and 12.46 pH

(c) Press      key to enter into next parameter setting or press      key to confirm and to return to measuring mode.

4.4.3. Purified water pH temperature compensation setting (P2)

(a) Short-time press      key in the mode P1, the meter enters into mode P2, see picture (4-3).

(b) Press      or      to choose “     ” (purified water pH temperature compensation setting) or “      ” (OFF).

(c) Press      key to enter into next parameter setting or press      key to conform and return to measuring mode.

(d) The default setting is “ OFF ”.

Note: There will appear“       ” icon in the upper right corner of the LCD if purified water temperature compensation function was set.

4.4.4. Ammonia added purified water pH temperature compensation setting (P3)

(a) Short-time press      key in mode P2 to enter into mode P3, see picture (4-4).

(b) Press      or      key to choose “    ”（ammonia added purified water pH temperature compensation setting）or “      ” (OFF).

(c) Press      key to enter into next parameter setting or press       key to conform and return to measuring mode.

(d) The default setting is “ OFF ”.

Note: If set the ammonia purified water pH temperature compensation function, there will appear “       ” icon in the right upper corner of the LCD when in the measuring mode.

4.4.5. Temperature unit ℃/℉ setting (P4)

(a) Short-time press      key in mode P3 to enter into mode P4, see picture (4-5).

(b) Press      or      key to choose temperature unit: ℃ or ℉.

(c) Press      key to enter into next parameter setting or press      key to conform and return to measuring mode.

4.4.6. Backlight display time setting (P5)

(a) Short-time press      key in mode P4 to enter into mode P5, see picture (4-6).

(b) Press      or      key to choose the time of backlight auto power off: 0 min, 1 min, 3 min or 6 min. The backlight display function will be closed if choosing a 0min.

(c) Press      key to enter into next parameter setting or press      key to conform and return to measuring mode.

(d) The default setting for P5 is 1 min.

4.4.7.Auto power off time setting (P6)

(a) Short-time press      key in mode P5 to enter into mode P6, see picture (4-7).

(b) Press      or      key to choose the time: 0 min, 10 min or 20 min. The auto power off function will be closed if choosing 0 min.

(c) Press      key to enter into next parameter setting or press      key to conform and return to measuring mode.

(d) The default setting for P6 is 10 min.

4.4.8. Restore to default setting (P7)

(a) Short-time press      key in mode P6 to enter into mode P7, see picture (4-8).

(b) Press      or      key to choose “     ”, means that all the pH measuring parameters have been restored to the default setting mode, and restore to measuring mode after 2 seconds.

4.5. Considerations:

4.5.1.Calibration times of meter rely on the sample, electrode performance and required accuracy. For high accurate measurement (≤±0.02pH), which should be calibrated immediately with high accurate standard buffer solution, for general accuracy measuring (≥±0.1pH), which can be used almost one week or long time once be calibrated.

4.5.2. The meter must be recalibrated in the following situations:

(a) New changed or unused electrode for a long time;

(b) After measuring acid (pH<2) or alkaline (pH>12) solution;

(c) After measuring solution which contains fluoride and concentrated organic solution;

(d) The solution’s temperature is much different with calibration temperature.

4.5.3. The soaking solution in the protecting bottle of front pH electrode is to keep the glass bulb and junction activating. Loose the capsule, pull out the electrode and wash it in purified water before measuring. Insert the electrode and screw tight the capsule after measuring to prevent the solution leaking. If the soaking solution is turbid or moldy, please wash and change a new one at once.

4.5.4. The preparation of the soaking solution: take 25g analytical pure KCl, dissolved with purified water and dilute to 100mL. Electrode should avoid soaking in purified water protein solution and acid fluoride solution for a long time as well avoid getting touching with organic silicon lipidic matters.

4.5.5. To calibrate the meter with the given value pH buffer solution, the pH value of the standard buffer solution must be reliable so as to improve the accuracy. Buffer solution should be changed in time after many times using.

4.5.6. Always keep the meter clean and dry, especially for the socket of meter and electrode, otherwise it may lead to an inaccurate measurement or invalidity. To clean and dry them with medical cotton with dehydrated alcohol if there are any dirty.

4.5.7. The sensitive glass bulb in the front of combination electrode should not touch with hard things, any broken and rough will make the electrode invalidity. Before and after measuring, the electrode should be washed with purified water, and dry electrode after washing, don’t clean glass bulb with tissue for it will effect stability of electrode potential and enlarge response time. The electrode should be washed many times for removing the sample stuck on the electrode, or wash with suitable solvent then clean the solvent with purified water after measuring sticky sample.

4.5.8. An electrode be used for a long time, or measured solution which contains a polluting solute easily for the sensitive bulb, or a substance resulting in jam at the junction, the electrode will be getting passivated, its sensitivity will decrease and its response is getting slow, the reading are not correct. It could adopt the following method for various cases:

(a) The glass bulb is contaminated and aging: Put the electrode into 0.1mol/L dilute hydrochloric acid (Preparation: diluted 9mL hydrochloric acid to 1000mL with purified water) for 24h. Rinse it with purified water, then dipped it into the electrode dipping solution for 24h. If the passivation is serious, then user can also put the bulb of electrode into 4% HF (hydrofluoric acid) or the electrode activation solution for 3 to 5 seconds, rinsing it with purified water, and dipped it in the electrode soaking solution for 24h to renew it.

(b) Wash for contaminated glass bulb and junction: (For reference)

Contamination                         Abluent

Inorganic metal oxide               diluted acid less than 1mol/L

Organic lipidic matter              dilute washing (weak alkaline)

Resin macromolecule matter           dilute alcohol, acetone, ether

Proteinic haematocyte sediment    Acidic enzymatic solution (such as dried yeast)

Kinds of paint                   dilute bleacher, peroxide

4.5.9. pH electrode using period is about 1 year, but its life will be shortened if using condition is poor or incorrect maintenance. So it should be replaced immediately after electrode become aging or invalid.

4.5.10. When it appears an abnormal reading when calibration or displaying, please set P7 as “ON” to restore the meter to default setting mode, and then to calibrate and measure again.

4.6. The Self-diagnose Information:

When using, there might appear the following icons. This is the meter’s self-diagnose information, which can help to know some information about the meter or the electrode when using:

4.6.1. The stable icon -2.00 pH or the 19.99 pH — this icon appeared when the value has surpassed the measuring range. There will also appear such signs when the electrode is not well connected with the meter or when the electrode is not immerged into the solution. This is a normal phenomenon.

4.6.2. “       ” — Electrode zero potential to be exceeded (<-60mV or >60mV)

4.6.3. “       ” — Electrode slope to be exceeded (< 85% or >105%)

When appear “       ” or “       ”, the meter can not work, please take the following check:

(a) Check if the electrode bulb has air bubble, if has, please shake it hardly.

(b) Check the quality of buffer solution, if it goes bad or the value has biggish error.

(c) Set the meter to default setting mode (for details see P14 item 4.4.8.), then recalibrate it.

If still can not recover the normal state after doing the above checking, please replace a new pH electrode.

5. mV Measurement:

5.1. Sample Test:

5.1.1. In pH measuring mode, short-time press      key to show      mode;

5.1.2. In other measuring mode,plug the ORP combined electrode(Sold separately) into the meter corresponding socket, wash the probe with purified water and shake off water, insert it into the sample solution, slowly stir and then let it still. When appears a “    ” ,the stable reading is the right value.

5.2. Parameter Setting:

5.2.1. mV measurement parameter setting table (Table (5-1))

Table (5-1)

| Prompt Mark | Parameter Setting Items | Code | Parameters |
| --- | --- | --- | --- |
| P1 | Backlight display time setting |  | 0-1-3-6 min |
| P2 | Time setting for auto power off |  | 0-10-20 min |

5.2.2. Backlight display time setting (P1)

Press the      key until entering into mode P1, see detailed operations in P14 item4.4.6.

5.2.3. Time setting for auto power off (P2):

Press the      key in mode P1, the meter enters into mode P2, see detailed operations in P14 item 4.4.7.

6. DO Measurement:

6.1. Preparation Work:

6.1.1. Check the DO500 DO electrode: see picture (6-1), screw off the bottom lid of electrode calibration cover, check whether the water stored sponge inside is soaked or not. If not, please drop some appropriate purified water, then upside down the calibration cover to remove of rest purified water (Attention: no stored water in calibration cover); repeatedly check the electrolyte within electrode membrane cap can’t have air bubbles (except the smaller air bubbles). If there are lager air bubbles, please screw off the membrane cap, add electrolyte and screw on.

6.1.1. Press      key to turn on the meter and to enter into DO mode, plug the electrode

into meter to polarize for 5 minutes.

6.1.2. Units Choosing: press      to choose unit: mg/L, ppm and %; then release after choosing made.

6.2. Meter Calibration:

Press      key, the meter enters into calibration mode, upper-right corner of LCD will display a flickering “ CAL”. Insert the DO electrode into calibration cover, and screw the cap of calibration cover tightly, place vertically to polarize 3 to 5 minutes, after the display value is stable and when the “      ” icon appears, again press

key to calibrate. LCD will display flickering “100%”, the calibration is finished after several seconds and return to the measuring mode. If the display value is unstable, please wait for several minutes and again press      key until the display value is completely stable.

6.3. Water Sample Test:

6.3.1. To measure in the flowing water (water sample velocity of flow >5cm/s): Insert the DO electrode into water, water surface should excess location of the thermistor which is on the electrode, the direction for electrode and water is 45° to 75° angle, and swing slightly the electrode, then hold 3 to 5 minutes to take the reading after stable displayed.

6.3.2. To measure in static water: Insert DO electrode into water, water surface should excess location of the thermistor, the direction for electrode and water is 45° to 75° angle, quickly remove the electrode in water, the remove speed is >5cm/s, hold 3 to 5 minutes to take the reading after stable displayed;

6.3.3. To test in the slower flowing water: according to the method of item 6.3.1, but the remove speed of electrode should be faster.

6.4. Considerations:

6.4.1. The air temperature and water temperature should closer when test (≤10℃), if it is biggish different, please immerse the electrode into the sample water for about 10 minutes, then insert the electrode into the calibrate cover for 5 to 6 minutes according to item 6.2.1 and press the      key to do calibration.

6.4.2. After each start-up meter must be carried out as required electrode polarization and calibration, so do not turn off the meter while using (when in the DO mode, the time for default setting for auto power off is 0).

6.4.3. During DO test, temperature effects a lot on measuring value, because the thermistor of DO electrode is installed on the electrode housing, direct touches with water, not equips in the electrolyte of internal electrode, both of them have difference for sensibility on the water temperature, generally will need 3 to 5 minutes to make the thermistor temperature and the electrolyte within electrode of the actual temperature reached unanimity, so the reading time must be > 3 minutes, or else they would have a greater error, especially when the electrode temperature and water temperature has larger difference, will need to extend the reading time;

6.4.4. DO measurement also can be affected by barometric pressure, and has biggish inaccuracy (please see TableⅤ and Ⅵ). This meter has automatic barometric pressure compensation function, and the measuring value has been amended by the barometric pressure compensation.

6.4.5. DO electrode can not be static test in the static water, otherwise it will lead the results to be lower.

6.4.6. When measuring, there can not have any bubbles for DO electrode and the sensitive membrane which touches with water, otherwise will affect measuring accuracy.

6.4.7. Electrode within the electrolyte can not have air bubbles (except for the smaller air bubbles), otherwise it will affect the response speed and accuracy. If a large air bubble, should screw out membrane cap, then add the electrolyte and membrane cap is screwed up;

6.4.8. Should keep the surface of DO electrode sensitive membrance to be wet to prevent the cathode electrolyte withered, there is a piece of water stored sponge in calibration cover bottom, see picture (6-1), user must always maintain a moist sponge when the sponge is dried should be dropping some water on it (Let sponge saturated absorption, but you can not have a water flow). And tightly twist electrode cap, dissolved oxygen electrode should be kept in humid conditions.

6.4.9. The salinity electrode be installed in DO electrode (see picture (6-1)), the surface is plated with a layer of metal platinum black to reduce electrode polarization. The surface of platinum can not be rubbed. It only can be washed by shaking in water in case of damaging the platinum black coating. Wash the organic component pollution with tepid water containing detergent or with alcohol.

6.4.10. When it is abnormal for calibrating or displaying, please set the mode P7 to be “On” to restore meter to default setting state, and then do calibrate and measure.

6.5. Parameter Setting:

6.5.1. DO parameter setting table (Table (6-1))

Table (6-1)

| Prompt Mark | Parameter Setting Items | Code | Parameter |
| --- | --- | --- | --- |
| P1 | Resolution selection |  | 0.01/0.1(mg/L and ppm) 0.1/1(%) |
| P2 | Salinity calibration |  |  |
| P3 | Barometric pressure setting |  |  |
| P4 | Temperature unit selection |  | ℃ ℉ |
| P5 | Backlight display time setting |  | 0-1-3-6 min |
| P6 | Time setting for auto power off |  | 0-10-20 min |
| P7 | Restore to default setting |  | OFF-On (shut-set) |

6.5.2. Resolution selection (P1)

(a) Press      key until entering into P1 mode, see picture (6-2);

(b) Press      key to choose resolution: 0.01→0.1;

(c) Press      key to enter into the next parameter setting or press      return to the measuring interface.

6.5.3. Salinity calibration (P2)

(a) Under P1 mode, press      key to enter into P2 mode, shown as picture (6-3).

(b) Insert the DO probe into 12.88mS/cm calibration solution, solution level should excess salinity electrode, still it after stirring, then press the      key when the value to be stable and appears an icon “    ”, the LCD displays a flickering “12.9”. The calibration is finished after several minutes and displays a stable salinity value.

(c) Press      key to enter into the next parameter setting or press      key to confirm and return to the measuring interface.

6.5.4. Barometric pressure setting (P3)

(a) Under mode P2, press      key to enter into P3 mode, the LCD displays a barometric pressure measuring value set earlier (unit is kPa), such as 101.3 kPa, show as picture (6-4).

(b) Press      or        to change according to standard barometric pressure value showed on standard barometric pressure table, and press the      or      without releasing to quickly change the value.

(c) Press      key to enter into next parameter setting or press      key to confirm and return to measuring mode.

6.5.5. Temperature unit ℃/℉ selection (P4)

Press      key in mode P3 to enter into mode P4, the rest operations see P13 Item 4.4.5.

6.5.6. Backlight display time setting (P5)

Press      key in mode P4 to enter into mode P5, the rest operations see P14 item 4.4.6.

6.5.7. Time setting for auto power off (P6)

Press      key in mode P5 to enter into mode P6, the rest operations see P14 item 4.4.7.

Note: when in the measuring mode of DO, the default setting time for auto power off is 0, it means the auto power off functiong is closed.

6.5.8. Restore to default setting (P7)

Press      key in mode P6 to enter into mode P7, the rest operation see P14 item 4.4.8.

6.6. Change the Membrane Cap:

When electrode response time to be longer, and displaying value appear obviously error, or when sensitive membrane of DO electrode has a wrinkle, crack or damaged, should replace a membrane cap according to the following steps.

(a) Remove the membrane cap;

(b) Wash the electrode without membrane cap with purified water and throw off the water on it.

(c) Polish lightly on cathode surface (gold slice) with a piece of clear velvet or tissue;

(d) Inject electrolyte into a new membrane cap slowly and do not has any air bubbles appeared. After injected, should check the electrolyte to see whether there are air bubbles in it. If found there are air bubbles, please carefully knock on the membrane cap to dispel them.

(e) Put membrane cap on the table and set in the electrode vertically, slowly clockwise twist, then twist tightly. Temporality, redundant electrolyte will squeeze out, please wipe it out with tissue and clear it in the purified water.

(f) Check there is no air bubbles in electrolyte (beside the smaller air bubbles), otherwise, re-assembly is needed.

(g) No touching the sensitive membrane when use the electrode and replace the membrane cap for the sweat and grease will affect the quality of membrane and decrease the rate of oxygen permeability.

6.7. Zero Oxygen Calibration:

Generally need to do zero oxygen calibration in the situation of replacing new electrode, replacing membrane cap, and long time without using. Usually no need to do zero oxygen calibration. To do zero oxygen calibration by following below steps:

(a) Preparing 100mL anaerobic water: Add 5g anhydrous sodium sulfite (Na2SO3) in 100 mL beaker, plus 100 mL purified water stir to dissolve, anaerobic water is effective within 24 hours.

(b) To polarize electrode by connecting it into meter for 15min, and to do calibration according to the item 7.2.1;

(c) Put the electrode into anaerobic water, press      key to make meter enter into calibration mode, when displaying value is ≤ 0.15mg/L (≤5min), again press

key to calibrate, then the LCD displays a flickering 0.0%, the calibration finishes after several seconds and the meter displays 0.00 mg/L, then wash the electrode with purified water;

(d) If the meter displayed value ≤0.02mg/L within 5 minutes, means the response speed of meter and residual current are good, need not to do zero oxygen calibration, press

key to return to measuring mode.

(e) If the value of meter is displayed as >0.15 mg/L after 5 minutes, means the response speed of meter is slow and large residual current, at this time can change the membrane cap or remove the membrane cap and take out the polishing paper in the accessories to lightly wipe the catholic of gold plane (to the surface along the arc of gold-polished), and clean surface of cathode with soft cloth to tissue, and wash electrode with purified water and shake off the water on it. Add some electrolyte into the membrane cap, re-assembly and screw tightly, then to do full scale and zero oxygen calibration according to item 6.2. and 6.7.

6.8. Salinity Calibration:

The meter has function of salinity automatic compensation, but this salinity calibration is just need to apply in replacing a new electrode and long-time without using. In general do not need to do salinity calibration. The meter has already been made salinity calibration before leaving factory, so it is no need to do salinity calibration for the first time using. For the procedure of salinity calibration, please refer to P23 item 6.5.3.

6.9. Barometric Pressure Setting:

The meter has the function of barometric pressure manual compensation. It is suggested to reset the value according to the standard barometric pressure table when there is a biggish change of barometric pressure while using to ensure the accuracy of barometric pressure compensation（see the Table Ⅴ and Ⅵ）. For the procedure of barometric pressure calibration, please see P24 item 6.5.4

7. Meter’s Complete Kit:

7.1. Model SX725 pH / mV / DO meter

7.2. 201T-S plastic pH/ATC three-in-one combined electrode

7.3.DO500 DO electrode

7.4. pH standard buffer solution

( pH4.00, pH7.00 and pH10.01/50mL)

7.5. DO503 membrane cap (DO electrode use)

7.6. DO502 DO electrode inner solution (30mL)

7.7. Cathode polishing paper

7.8. Screw driver

7.9. Spare AA batteries

7.10. Operation manual

7.11. Carrying case

8. Warranty:

8.1. We warrant this meter to be free of charge maintain, replace the parts or products under normal using circumstances, from purchased time within one year caused by manufacturing bad and unable to work.

8.2. Attached electrodes do not belong to this warrant range, besides temperature electrode. But, if the newly purchased electrode went wrong without using, it’s free of charge to maintain and replace.

8.3. The above warranty is not apply to defects resulting from action of user such as misuse, improper wiring, operation outside of specification, improper maintenance or repair, or unauthorized modification.

Table І Meter Parameter Setting Table

| Mode | Prompt Mark | Parameter Setting Items | Code | Parameter |
| --- | --- | --- | --- | --- |
| pH | P1 | pH buffer solution series selection |  | USA (Europe & U.S.A. series) NIS (NIST series) CH (China series) |
|  | P2 | Purified water pH temperature compensation setting |  | OFF-On (shut-set) |
|  | P3 | Purified water added with ammonia pH temperature compensation setting置 |  | OFF-On (shut-set) |
|  | P4 | Temperature unit setting |  | ℃ ℉ |
|  | P5 | Backlight display time setting |  | 0-1-3-6 min |
|  | P6 | Time setting for auto power off |  | 0-10-20 min |
|  | P7 | Restore to default setting |  | OFF-On (shut-set) |
| mV | P1 | Backlight display time setting |  | 0-1-3-6 min |
|  | P2 | Time setting for auto power off |  | 0-10-20 min |
| Dissolved Oxygen | P1 | Resolution selection |  | 0.01/0.1(mg/L and ppm) 0.1, 1(%) |
|  | P2 | Salinity calibration |  |  |
|  | P3 | Barometric pressure setting |  |  |
|  | P4 | Temperature unit selection |  | ℃ ℉ |
|  | P5 | Backlight display time setting |  | 0-1-3-6 min |
|  | P6 | Time setting for auto power off |  | 0-10-20 min |
|  | P7 | Restore to default setting |  | OFF-On (shut-set) |

Talbe Ⅱ Meter Restore to Default Setting Table

| Mode | Prompt Mark | Parameter Setting Items | Default Setting Content | Non Default Setting Content Icon |
| --- | --- | --- | --- | --- |
| pH | P2 | Purified water pH temperature compensation setting | OFF |  |
|  | P3 | Purified water added with ammonia pH temperature compensation setting | OFF |  |

Table Ⅲ Code Icon and Abbreviation Table

Table Ⅳ Content of Saturated Oxygen in Water at Different Temperature

| Temperature ℃ | DO mg/L | Temperature ℃ | DO mg/L | Temperature ℃ | DO mg/L |
| --- | --- | --- | --- | --- | --- |
| 0 | 14.64 | 16 | 9.86 | 32 | 7.30 |
| 1 | 14.22 | 17 | 9.66 | 33 | 7.18 |
| 2 | 13.82 | 18 | 9.46 | 34 | 7.07 |
| 3 | 13.44 | 19 | 9.27 | 35 | 6.95 |
| 4 | 13.09 | 20 | 9.08 | 36 | 6.84 |
| 5 | 12.74 | 21 | 8.90 | 37 | 6.73 |
| 6 | 12.42 | 22 | 8.73 | 38 | 6.63 |
| 7 | 12.11 | 23 | 8.57 | 39 | 6.53 |
| 8 | 11.81 | 24 | 8.41 | 40 | 6.43 |
| 9 | 11.53 | 25 | 8.25 | 41 | 6.34 |
| 10 | 11.26 | 26 | 8.11 | 42 | 6.25 |
| 11 | 11.01 | 27 | 7.96 | 43 | 6.17 |
| 12 | 10.77 | 28 | 7.82 | 44 | 6.09 |
| 13 | 10.53 | 29 | 7.69 | 45 | 6.01 |
| 14 | 10.30 | 30 | 7.56 |  |  |
| 15 | 10.08 | 31 | 7.43 |  |  |

TableⅤ Content of Saturated Oxygen at Different Barometric pressure

| Barometric Pressure | Dissovled Oxygen (mg/L) |
| --- | --- |
| mmHg | kPa | 15℃ | 25℃ | 35℃ |
| 750 | 100.00 | 9.94 | 8.14 | 6.85 |
| 751 | 100.13 | 9.96 | 8.15 | 6.86 |
| 752 | 100.26 | 9.97 | 8.16 | 6.87 |
| 753 | 100.40 | 9.98 | 8.17 | 6.88 |
| 754 | 100.53 | 9.99 | 8.18 | 6.89 |
| 755 | 100.66 | 10.00 | 8.20 | 6.90 |
| 756 | 100.80 | 10.01 | 8.21 | 6.91 |
| 757 | 100.93 | 10.03 | 8.22 | 6.92 |
| 758 | 101.06 | 10.04 | 8.23 | 6.93 |
| 759 | 101.20 | 10.07 | 8.24 | 6.94 |
| 760 | 101.33 | 10.08 | 8.25 | 6.95 |
| 761 | 101.46 | 10.09 | 8.26 | 6.96 |
| 762 | 101.60 | 10.11 | 8.27 | 6.97 |
| 763 | 101.73 | 10.12 | 8.28 | 6.98 |
| 764 | 101.86 | 10.14 | 8.30 | 6.99 |
| 765 | 102.00 | 10.15 | 8.31 | 7.00 |
| 766 | 102.13 | 10.16 | 8.32 | 7.01 |
| 767 | 102.26 | 10.18 | 8.33 | 7.02 |
| 768 | 102.40 | 10.19 | 8.34 | 7.02 |
| 769 | 102.53 | 10.21 | 8.35 | 7.03 |
| 770 | 102.66 | 10.22 | 8.36 | 7.04 |
| 771 | 102.80 | 10.23 | 8.37 | 7.05 |
| 772 | 102.93 | 10.25 | 8.39 | 7.06 |
| 773 | 103.06 | 10.26 | 8.40 | 7.07 |
| 774 | 103.19 | 10.28 | 8.41 | 7.08 |
| 775 | 103.33 | 10.29 | 8.42 | 7.09 |

Conversion between mmHg and kPa: mmHg×0.13333=kPa

DOpt = P×DOt ÷ 760

Note: DOpt — DO concentration under temperature t, barometric pressure P, mg/L;

P — Barometric pressure, mmHg;

DOt — DO concentration under temperature t, barometric pressure 760mmHg, mg/L;

760 — Barometric pressure, mmHg.

TableⅥ Content of Oxygen at Different Altitude

| Altitude | Barometric Pressure | DO | Altitude | Barometric Pressure | DO |
| --- | --- | --- | --- | --- | --- |
| Foot | meter | kPa | mmHg | mg/l | Foot | meter | kPa | mmHg | mg/l |
| 0 | 0 | 101.3 | 760 | 8.25 | 7500 | 2287 | 77.1 | 579 | 6.28 |
| 500 | 152 | 99.34 | 746 | 8.09 | 8000 | 2439 | 75.63 | 568 | 6.16 |
| 1000 | 305 | 97.6 | 733 | 7.95 | 8500 | 2591 | 74.44 | 559 | 6.06 |
| 1500 | 457 | 95.87 | 720 | 7.81 | 9000 | 2744 | 72.97 | 548 | 5.94 |
| 2000 | 610 | 94.28 | 708 | 7.68 | 9500 | 2896 | 71.64 | 538 | 5.83 |
| 2500 | 762 | 92.54 | 695 | 7.54 | 10000 | 3049 | 70.17 | 527 | 5.71 |
| 3000 | 915 | 90.95 | 683 | 7.41 | 10500 | 3201 | 68.84 | 517 | 5.61 |
| 3500 | 1067 | 89.35 | 671 | 7.28 | 11000 | 3354 | 67.38 | 506 | 5.49 |
| 4000 | 1220 | 87.75 | 659 | 7.15 | 12000 | 3659 | 66.58 | 500 | 5.42 |
| 4500 | 1372 | 86.15 | 647 | 7.02 | 13000 | 3963 | 65.78 | 494 | 5.36 |
| 5000 | 1524 | 84.56 | 635 | 6.89 | 14000 | 4268 | 64.98 | 488 | 5.29 |
| 5500 | 1677 | 83.09 | 624 | 6.77 | 15000 | 4573 | 64.18 | 482 | 5.23 |
| 6000 | 1829 | 81.63 | 613 | 6.65 | 16000 | 4878 | 63.38 | 476 | 5.16 |
| 6500 | 1982 | 80.03 | 601 | 6.52 | 17000 | 5183 | 62.58 | 470 | 5.10 |
| 7000 | 2134 | 78.56 | 590 | 6.40 | 18000 | 5488 | 61.79 | 464 | 5.03 |

Shanghai San-Xin Instrumentation, Inc.

Add: 3/F. Building No.4, No. 471 Guiping Road, Shanghai, China. 200233

Tel: +86-21-63362480                   Fax: +86-21-64956880

Website: www.shsan-xin.com            E-mail: wxmab@shsan-xin.com

# Text boxes

### Dissolved Oxygen

### End

### Pure2

Purified water added with ammonia pH temperature compensation setting

### Pure1

Purified water pH temperature compensation setting

### NIST

NIST series standard

### Calibration

### USA

Europe & U.S.A series standard

### China

China series standard

### Solution

### Standard solution

4

5

7

7

8

9

10

10

10

11

12

15

17

18

18

18

19

19

20

20

21

23

1 — DO500 DO electrode

2 — Calibration cover cap

3 — Calibration cover

4 — Water stored sponge

5 — Calibration cover bottom lip

6 — Temperature electrode

7 — Membrane cap

8 — Cathode (gold slice)

9 — Anode

10 — Salinity electrode

25

25

27

27

28

28

### Code and Abbreviation

### English

Explanation

1set

1pcs

1pcs

1btl/each

3pcs

1btl

2pcs

1pcs

2pcs

1pcs

1unit