5053 Salinity Pocket Tester Operation Manual

Ⅰ. Technical Parameters:

1.  Range：0 ~ 20.0 %

2.  Resolution: 0.1%

3.  Accuracy：±2%FS±one world

4.  Temperature compensation range：0 ~ 60℃（automatic），no temperature display.

5.  Automatic：1 point calibration（50.0 mS/cm）

6.  Power：AAA Batteries x 2，for more than 100 consecutive hours.

7.  Dimension and Weight：39×25×176mm / 87g

8.  Waterproof Grade：IP57

Ⅱ. Operations:

1. Calibration：Wash the electrode with pure water and shake off the water on the electrode,

place it into 50.0 mS calibration solution, still it after stirring until the reading is stable. Repeat

the above operation several times until the reading value is displayed repeatedly. Press the

key for about 2 seconds until the icon “     ” appears on LCD, then “50.0 mS” is twinkling.

The calibration is done after several seconds.

2. Measurement：Wash the electrode with pure water and shake off the water on the electrode,

place it into the sample solution, still it after stirring, and take the reading when the value is

stable.

3. Wash the electrode with pure water and shake off the water on the electrode, then replace

the electrode cap on it.

Notes:

1. The meter has been calibrated before delivery. Generally, user can use it directly. Under

normal situations, calibrate once every 1-2 months.

2. If the electrode is stained, please clean it by warm water with detergent or alcohol. Please do

not wipe it to avoid damaging the platinum black coating on the electrode surface.

3. The meter is automatically power off under no-operation condition next 10 minutes.

4. Please replace the battery when the icon “   ” appears on LCD.

Shanghai San-Xin Instrumentation, Inc.

Address：3/F, Building No.4, No.471 Guiping Road, Shanghai 200233, China.

Tel：86-21-63362480                Fax：86-21-64956880

Website：www.shsan-xin.com        E-mail:wxmab@shsan-xin.com