Model MP521

Lab pH/Conductivity Meter

Operation Manual

Shanghai San-Xin Instrumentation, Inc.

MP500 Series of Electrochemical Meters

I．Multi-parameter Meter

Model MP551pH/mV/ISE/Cond/DO Meter

II．Double Parameter Meters

1. Model MP521 Lab pH/Cond Meter (pH: ±0.01 pH; Conductivity: ±1.0% FS)

2. Model MP522 precision pH/Cond Meter (pH: ±0.002 pH; Conductivity:±0.5% FS)

3. Model MP523 pH/ISE Meter

4. Model MP525 pH/DO Meter

5. Model MP526 Cond/DO Meter

III．Single Parameter Meters

- MP511 Lab pH Meter（±0.01 pH）

- MP512 Precision pH Meter（±0.002 pH）

- MP513 Lab Conductivity Meter（±1.0% FS）

- MP515 Precision Conductivity Meter（±0.5% FS）

- MP516 DO Meter

- MP517 Sodium Ion Concentration Meter

- MP518 Calcium Ion Concentration Meter

- MP519 Fluoride Ion Concentration Meter

Table of Contents

1. Brief Introduction -------------------------------------------------------------------------

2. Technical Parameters ---------------------------------------------------------------------

3. Instructions to the Meter -------------------------------------------------------------------

3.1. LCD Display ---------------------------------------------------------------------------

3.2. Operation Keys ------------------------------------------------------------------------

3.3. The Storage, Recall and Elimination of the Measuring Information -----------

3.4. RS-232 Communication --------------------------------------------------------------

3.5. Sockets ----------------------------------------------------------------------------------

4. pH Measurement --------------------------------------------------------------------------

4.1. Preparation Work ----------------------------------------------------------------------

4.2. Electrode Calibration ------------------------------------------------------------------

4.3. Sample Test -----------------------------------------------------------------------------

4.4. Parameter Setting ----------------------------------------------------------------------

4.5. Considerations -------------------------------------------------------------------------

4.6. The Self-diagnose Information ------------------------------------------------------

5.	mV and ORP Measurement -------------------------------------------------------------

5.1. Sample Test -----------------------------------------------------------------------------

5.2. Considerations -------------------------------------------------------------------------

5.3. Parameter Setting ----------------------------------------------------------------------

6. Conductivity Measurement --------------------------------------------------------------

6.1. Preparation Work ----------------------------------------------------------------------

6.2. Electrode Calibration ------------------------------------------------------------------

6.3. Sample Test -----------------------------------------------------------------------------

6.4. Important Statement -------------------------------------------------------------------

6.5. Parameter Setting ----------------------------------------------------------------------

6.6. Considerations -------------------------------------------------------------------------

7. Meter’s Complete Kit -------------------------------------------------------------------

8. Warranty -----------------------------------------------------------------------------------

ChartⅠ Meter Parameter Setting Schedule

ChartⅡ Meter Restore to Producer Setting Schedule

Chart Ⅲ Code Icon and Abbreviation Schedule

1. Brief Instruction:

Thanks for buying and using the model MP521 Lab pH/Conductivity Meter (the following called “meter” in short).

Before using this meter, please read the operation manual carefully in order to help use and maintain it correctly. On the basis of improving instrument of performance constantly, we reserve the right of changing the content of this manual and accessories in case of not notifying in advance.

This meter is a perfect combination with the most advanced electronic technology, sensor technology and software design, it is the best pH, mV and conductivity meter with the highest performance and the lowest cost. It is suitable for measuring the pH and conductivity value of laboratory common aqueous solution, also can test the electric potential of various ion selective electrodes.

Built-in microprocessor chip, beautiful appearance, multi-functional and easy to use, this meter has the following prominent features:

1.1.	Meter meets with the requirement of international GLP standards. Its features up to automatic calibration, automatic temperature compensation, data storage, timing measurement, RS-232 output, clock display, function setting and automatic self-diagnose etc.

1.2.	Adopts digital filter and step slipping technology to intelligently improve meter’s response speed and result accuracy. “     ” will appear when reading to be stable.

1.3.	Automatically recognize pH standard buffer solution. User can choose anyone from three series of buffer solutions: Europe & U.S.A. series, NIST series and CHINA series.

1.4. Adopts advanced automatic frequency convertible and voltage regulation technology to enlarge the measuring range of conductivity electrode which constant K=1 to 10 times. It can meet measuring accuracy requirement for range within 100mS/cm by using only one-point calibration. It is the particular one–point calibration function for this meter.

1.5. Meter’s circuit board adopts SMT film-covering technology to improve meter’s production reliability.

1.6. Meter has the only product series number.

1.7. Dustproof and waterproof meter meets IP54 rating. All sockets are seal protected by the silica gel caps.

2. Technical Parameters:

2.1. pH:

| Measuring range | (-1.99 to 19.99) pH |
| --- | --- |
| Resolution | 0.1/0.01 pH |
| Accuracy | Meter:±0.01 pH, connect meter with electrode:±0.02pH |
| Input current | ≤2×10-12 A |
| Input impendance | ≥1×1012 Ω |
| Stability | ±0.01 pH/3h |
| Temp. compensation range | (0 to 100) ℃（automatic or manual） |

2.2. mV:

| Measuring range (mV/ORP) | -1999mV to 0 to 1999mV |
| --- | --- |
| Resolution | 1mV |
| Accuracy | ±0.1% FS |

2.3. Conductivity

| Measuring range | Conductivity: (0 to 2 x 105)μS/cm, be divided into 5 ranges: (0.00 to 20.00)μS/cm (20.0 to 200.0)μS/cm (200 to 2000)μS/cm (2.00 to 20.00) mS/cm (20.0 to 200.0) mS/cm Resistivity:(0 to 100) MΩ·cm TDS: (0 to 100) g/L Salinity : (0 to 100) ppt |
| --- | --- |
| Resolution | 0.01/0.1/1μS/cm 0.01/0.1 mS/cm |
| Accuracy | Meter: ±1.0% FS, connect meter with electrode: ±1.5% FS |
| Temp. compensation range | (0 to 50) ℃（automatic） |
| Electrode constant | 0.1 / 1 / 10 cm-1 |
| Reference temperature | 25℃ |

2.4. Temperature:

| Measuring range | -10℃to 110℃ |
| --- | --- |
| Resolution | 0.1℃ |
| Accuracy | 5 to 60℃ range: ±0.5℃ other range: ±1℃ |

2.5. Other Technical Parameters:

| Data storage | 900 groups |
| --- | --- |
| Storage content | Measuring value series number, measuring value, temperature, ATC or MTC state, measuring date and time |
| Power | DC9V/300mA |
| Communication connector | RS-232 |
| Size and weight | 160 x 190 x 70mm/880g |
| Quality and safety certification | ISO9001:2000, CE and CMC |

2.6. Working Condition:

| Environment temperature | 5 to 35 ℃ |
| --- | --- |
| Environmental humidity | ≤85% |
| IP rating | IP54 Dustproof and waterproof |

3. Instructions to the Meter:

3.1. LCD Display:

① —— Parameters mode icon

② —— Measuring value

③ —— Time and date

④ —— Measuring unit

⑤ —— Temperature compensation state icon:

ATC — automatic temperature compensation

MTC — manual temperature compensation

⑥ —— Serial number and icon as measuring value to be stored and recalled. M+ — measuring value to be stored icon; RM — reading to be recalled icon; The left digit means serial number.

⑦ —— Temperature measuring value and unit

⑧ —— RS-232 communication icon. When this icon appears, means meter has been connected with computer.

⑨ —— Timing measuring icon

⑩ —— Electrode calibration indicate icon

—— Measuring value to be stable icon

3.2. Operation Keys:

The meter has 8 operation keys in all.

3.2.1. <     > — Switch key

3.2.2. < CAL > — Calibration key, press the key make meter enters into calibration mode.

3.2.3. < MODE > — Function key

(a) Short-time press（time<1.5s）switch the measuring parameters, the meter will display      →     →     (conductivity).

(b) Depress (time >2s) to enter into the parameter setting mode P1, and again short-time press, will in turn display P2, P3…

3.2.4. < UNIT > —Unit and parameter mode key

(a) When in     mode, press the key to change pH resolution: 0.01→0.1 pH;

(b) When in     mode, press the key to in turn change the parameter mode     and     ;

(c) When in     mode, press the key to in turn change parameter mode:

（conductivity）→    （resistivity）→    （total dissolved solid）→     (salinity);

(d) Press the key to choose modified position of parameter when in the parameter setting state.

3.2.5. < ENTER > — Entrance key, when in the state of calibration or parameter setting, press the key to confirm. Then the meter enters into measuring state.

3.2.6. < ▲ >and < ▼ > — Increase and decrease key.

When in MTC state, press the key to increase or decrease the temperature value. Short-time press to alter 0.1℃ each time, while long-time press, the number will alter rapidly. When in the state of parameter setting, press the key to alter the number or ON/OFF state.

3.2.7. < M+/RM > — The compound key of memory and recall, short press (press time less than 1.5s) to save the measuring data, depress (press time more than 2s) to recall the saved measuring data.

3.3. The Storage, Recall and Elimination of the Measuring Information:

3.3.1. Store the measuring information:

(a) In the measuring mode, when the measuring data is stable and appear the “     ”, short-time press < M+/RM > key, LCD will display “ M+ ” icon and storage serial number, and meanwhile memory all the measuring information including time, date, serial number, measuring value, measuring unit, temperature, temperature compensation state and so on. Meter can separately store 300 groups of measuring information in the mode of pH, mV and conductivity, can total store 900 groups.

(b) When setting the timing measuring function, LCD will display the “     ” icon, and meter will to do timing measurement according to the time intervals set, and stores the measuring information;

3.3.2. Recall measuring information:

(a) Under the measuring mode, depress the < M+/RM > key, meter will recall the last stored information, and the storage number and “ RM ” icon will appear in the lower right corner of the LCD, and the complete measuring information, measuring time and date will be displayed alternatively on the upper right corner of LCD. Again press < ▼ > or  < ▲ > key, meter will in turn recall all the measuring information, depress < ▼ > or < ▲ > key can quickly query measuring information under other serial number;

(b) In the recalling mode (there are “ RM ” and storage serial number in the lower right corner of the LCD), press < ENTER > key to return to the measuring mode.

3.3.3. Eliminate the stored measuring information:

In the recalling mode, depress the < ENTER > key for 5 seconds, LCD will appear “     ” for 2 seconds. It means the storage has been eliminated, and returns to measuring mode.

3.4. RS-232 Communication:

3.4.1. This meter adopts MP521 communication software which can achieve RS-232 communication function. The requirements for computer of this application software are: PC can stable operate the Windows XP operating system (Microsoft Excel 2000 or much higher version has been installed in PC). According to different selection of resolution setting (1280 x 1024 or 1280 x 800) to install MP521 communication software by following the indications. Open the MP521 communication software and connect meter, the LCD will display “    ” RS-232 communication icon. Meanwhile all the stored information including measuring value, measuring unit, temperature, temperature compensation state (ATC or MTC) and the last calibration information (calibration time, date and solution) will be transmitted to the computer.

3.4.2. Again press < M+/RM > key after meter being connected with computer, or  timing measuring function has been set, all the measuring information will be transmitted to computer through RS-232 and won’t be stored in meter. And there will also display a curve graph of measurement value changing with time on the interface of computer.

3.4.3. User can do operation such as analyzing, counting or printing etc. for all measuring information in “Microsoft Excel” file through the key “Export” in the interface of computer.

3.5. Sockets:

3.5.1. REF — Reference electrode socket

3.5.2. pH/mV — pH electrode or ORP electrode socket（BNC socket）

3.5.3. TEMP — Temperature electrode socket（RCA microphone socket）

3.5.4. COND — Conductivity electrode socket（four pins socket）

3.5.5. RS232 — RS-232 communication connector socket

3.5.6. DC9V — DC9V power socket, Ф2.5, inner “+ ”outside “ - ”.

4. pH Measurement:

4.1. Preparation Work:

4.1.1. Switch in power, press <     > key to turn on.

4.1.2. Short-press < MODE > key to switch to     mode.

4.1.3. Install electrode holder and debug.

4.1.4. Install the 201T-M plastic pH/ATC three-in-one combination electrode into meter’s socket.

4.2. Electrode Calibration:

4.2.1. Press < CAL > to enter into calibration mode, LCD display the twinkling “        ” indicate to enter into the first point calibration.

4.2.2. Wash the pH electrode in purified water and dry it, then immerge it into the pH7.00 buffer solution, rock the electrode holder and then still, waiting for the data stable and appear “     ”, then again press key < CAL >, the LCD will appear a twinkling 7.00 pH, calibration finishes after several seconds and appear a stable pH value and a twinkling “       ”, indicates the first point calibration has been finished and enters into the second point calibration.

4.2.3. Wash the pH electrode in purified water and dry it, then immerge it into the pH4.00 buffer solution, rock the electrode holder and then still, waiting for a stable data and appear “    ”, then again press key < CAL >, the LCD will appear a twinkling 4.00pH, calibration finishes after several seconds, LCD will appear a stable pH value and a twinkling “      ”, indicates the second point calibration has been finished and enters into the third point calibration.

4.2.4. Wash the pH electrode in purified water and dry it, then immerge it into the pH10.01 buffer solution, rock the electrode holder and still, waiting “     ” for the data stable and appear, then again press key <CAL>, LCD will appear a twinkling “10.01” pH, calibration finishes after several seconds, appear a stable pH value and temperature value (the pH is relative with temperature at this time, such as 25◦C display 10.01pH, 20◦C display 10.06pH), at the same time will display “            ” three calibration indication icon, indicates the three-point calibration has been finished and entered into measuring mode. See picture (4-1).

4.2.5. Note:

(a) This meter can adopt random one-point, two-point or three-point automatic calibration, after the first point calibration (see item 4.2.2.), press < ENTER >  to confirm one-point calibration and enters into measuring mode. The indication icon “    ” for one-point calibration will appear on the lower left corner of LCD. When the measuring accuracy is ≤±0.1pH, user just need to choose one kind buffer solution to take one-point calibration according to the measuring range.

(b) After the second point calibration, (see item 4.2.3.), press < ENTER > to confirm two-point calibration and enter into measuring mode. The indication icon “         ” for tow-point calibration will appear on the lower left corner of LCD. User can choose pH4.00 and pH7.00 to calibrate if the measurement is just within the acidity range and choose pH7.00 and pH10.01 to calibrate if just within the alkalinity range.

(c) User should choose three-point calibration so as to reach a more accurate measurement if the measuring range is wide, or if the electrode has been used for long or has ageing phenomenon. As to the new electrode which be used for the first time, it must be calibrated by three-point calibration to keep the unanimity of the meter slope adjustment with the pH electrode.

4.3. Sample Test:

Immerge pH electrode into the sample solution after washing and dry it, rock the electrode holder and still, when the LCD appears the icon “     ” to take the reading after displaying value to be stable. Note, the closer the temperature of the sample solution and the calibration solution, the more accurate measuring value will be acquired. This is called the pH equal temperature measuring theory.

4.4. Parameter Setting:

4.4.1. pH measuring parameter setting schedule（chart（4-1））

Chart（4-1）

| Prompt Mark | Parameter Setting Items | Code | Parameters |
| --- | --- | --- | --- |
| P1 | pH buffer solution series selection |  | USA(Europe & U.S.A series) NIS (NIST series) CH (China series) |
| P2 | Time for timing measuring setting |  | 0 to 99 minutes |
| P3 | Temperature unit setting |  | ℃ ℉ |
| P4 | Date setting | Date | Month / Day / Year |
| P5 | Time setting | Time | Hour / Minute |
| P6 | Restore to producer setting |  | OFF-On (shut-set) |

4.4.2. pH buffer solution series selection（P1）

(a) Depress < MODE > key, meter enters into P1 mode: see picture（4-2）.

(b) Press < ▲ > or < ▼ > key to choose buffer solution series:

(Europe & U.S.A series) — 1.68, 4.00, 7.00, 10.01 and 12.45 pH

(NIST series) — 1.68, 4.01, 6.86, 9.18 and 12.45 pH

(China series) — 1.68, 4.00, 6.86, 9.18 and 12.46 pH

(c) Press < MODE > key to enter into next parameter setting or press < ENTER > key to conform and return to measuring mode.

(d) The producer setting is U.S.A (Europe & U.S.A series).

4.4.3. Time for timing measuring setting（P2）

(a) Short press < MODE > key in mode P2 to enter into mode P3. See picture（4-3）.

(b) Press < UNIT > key, the “   ” will move rightward in turn and flash. Press < ▲ > or < ▼ > to alter the number.

(c) Press < MODE > key to enter into next parameter setting or press < ENTER > key to conform and return to measuring mode.

(d) The producer setting is “ 0 ” second.

Note: in the form “       ”, the number in the front of    “ :” is minute, and the maximum setting is 99, in the return of “ :” is second, the maximum setting is 59. After setting the timing measurement mode, LCD will display the “     ” icon.

4.4.4. Temperature unit ℃/℉ setting（P3）

(a) Short press < MODE > key in mode P2 to enter into mode P3, see picture（4-4）.

(b) Press < ▲ > or < ▼ > key to choose temperature unit: ℃or℉.

(c) Press < MODE > key to enter into next parameter setting or press < ENTER > key to conform and return to measuring mode.

4.4.5. Date setting（P4）

(a) Short press < MODE > key in mode P3 to enter into mode P4, see picture（4-5）.

(b) Press < UNIT > key, the number will move rightward and flash, press < ▲ > or < ▼ > key to alter the number. The upper right is month-day and lower right is year.

(c) Press < MODE > key to enter into next parameter setting or press < ENTER > key to conform and return to measuring mode.

4.4.6.Time setting（P5）

(a) Short press < MODE > key in mode P4 to enter into mode P5, see picture (4-6).

(b) Press < UNIT > key, the number will move rightward and flash, press < ▲ > or < ▼ > key to alter the number.

(c) Press < MODE > key to enter into next parameter setting or press < ENTER > key to conform and return to measuring mode.

4.4.7. Restore to producer setting (P6)

(a) Short press < MODE > key in mode P5 to enter into mode P6, see picture (4-7).

(b) Press < ▲ > key to choose “     ”, means that all the pH measuring parameters have been restored to the producer setting mode, and restore to measuring mode after 2 seconds.

4.5. Considerations:

4.5.1. In order to make electrode’s response to be faster, should shake it several times and still then after immersing pH combination electrode into solution. Shaking should be more hard especially when using plastic pH combination electrode, because there is a small empty compartment between the electrode protecting cap and pH bulb, it is easy to produce air bulb and make bulb can not touch well with solution, so it can eliminate the air bulb by hardly shaking. Also can take measurement by twisting the protecting cap off, but be attention to the damage of bulb.

4.5.2. Calibration times of meter rely on the sample, electrode performance and required accuracy. For high accurate measurement (≤±0.02pH), which should be calibrated immediately with high accurate standard buffer solution, for general accuracy measuring (≥±0.1pH), which can be used almost one week or long time once be calibrated.

4.5.3. The meter must be renew calibrated in the following situations:

(a) New changed or unused electrode for a long time;

(b) After measuring acid (pH<2) or alkaline (pH>12) solution;

(c) After measuring solution which contains fluoride and concentrated organic solution;

(d) The solution’s temperature is much different with calibration temperature.

4.5.4. If user does not want to use the temperature electrode, can press < ▲ > or < ▼ > key to adjust temperature to do manual temperature compensation.

4.5.5. The soaking solution in the protecting bottle of front pH electrode is to keep the glass bulb and junction activating. Loose the capsule, pull out the electrode and wash it in purified water before measuring. Insert the electrode and screw tight the capsule after measuring to prevent the solution leaking. If the soaking solution is turbid or moldy, please wash and change a new one at once.

4.5.6. The preparation of the soaking solution: take 25g analytical pure KCL, dissolved with purified water and dilute to 100mL. Electrode should avoid soaking in purified water protein solution and acid fluoride solution for a long time as well avoid getting touching with organic silicon lipidic matters.

4.5.7. To calibrate the meter with the given value pH buffer solution, the pH value of the standard buffer solution must be reliable so as to improve the accuracy. Buffer solution should be changed in time after many times using.

4.5.8. Always keep the meter clean and dry, especially for the socket of meter and electrode, otherwise it may lead to an inaccurate measurement or invalidity. To clean and dry them with medical cotton with dehydrated alcohol if there are any dirty.

4.5.9. The sensitive glass bulb in the front of combination electrode should not touch with hard things, any broken and rough will make the electrode invalidity. Before and after measuring, the electrode should be washed with purified water, and dry electrode after washing, don’t clean glass bulb with tissue for it will effect stability of electrode potential and enlarge response time. The electrode should be washed many times for removing the sample stuck on the electrode, or wash with suitable solvent then clean the solvent with purified water after measuring sticky sample.

4.5.10. An electrode be used for a long time, or measured solution which contains a polluting solute easily for the sensitive bulb, or a substance resulting in jam at the junction, the electrode will be getting passivated, its sensitivity will decrease and its response is getting slow, the reading are not correct. It could adopt the following method for various cases:

(a) The glass bulb is contaminated and aging: Put the electrode into 0.1mol/L dilute hydrochloric acid (Preparation: diluted 9mL hydrochloric acid to 1000mL with purified water) for 24h. Rinse it with purified water, then dipped it into the electrode dipping solution for 24h. If the passivation is serious, then user can also put the bulb of electrode into 4% HF (hydrofluoric acid) solution for 3 to 5 seconds, rinsing it with purified water, and dipped it in the electrode soaking solution for 24h to renew it. Note: the electrode activation solution is available in our company.

(b) Wash for contaminated glass bulb and junction:（FYI）

Contamination                            Abluent

Inorganic metal oxide                  diluted acid less than 1mol/L

Organic lipidic matter                  dilute washing（weak alkaline）

Resin macromolecule matter               dilute alcohol、acetone、ether

Proteinic haematocyte sediment       Acidic enzymatic solution（such as dried yeast）

Kinds of paint                     dilute bleacher, peroxide

4.5.11. pH electrode using period is about 1 year, but its life will be shortened if using condition is poor or incorrect maintenance. So it should be replaced immediately after electrode become aging or invalid.

4.5.12. When it appears an abnormal reading when calibration or displaying, please set P6 as “ON” to restore the meter to producer setting mode, and then to calibrate and measure again.

4.5.13. Do not unplug the adapter when the meter is power on, only to unplug when the meter is power off.

4.6. The Self-diagnose Information:

When using, there might appear the following icons. This is the meter’s self-diagnose information, which can help to know some information about the meter or the electrode when using:

4.6.1. The stable icon “         ” pH or the “        ” pH — this is the icon appeared when the value has surpassed the measuring range. There will also appear such signs when the electrode is not well connected with the meter or when the electrode is not immerged into the solution. This is a normal phenomenon.

4.6.2. “       ” — Electrode zero potential to be exceeded  (<-60mV or >60mV)

4.6.3. “       ” — Electrode slope to be exceeded（< 85% or >105%）

When the upper right corner of LCD appears “       ” or “       ”, the meter will work abnormally, please take the following check:

(a) Check if the electrode bulb has air bubble, if has, please shake it hardly.

(b) Check the quality of pH buffer solution, if it goes bad or the value has biggish error.

(c) Set the meter to producer setting mode（for details see P15 item 4.4.7）, then recalibrate it.

If still can not recover the normal state after the above checking, please replace a new pH electrode.

5. mV and ORP Measurement:

5.1. Sample Test:

5.1.1. Press <     > to electrify meter, press < MODE > to switch to      , then press < UNIT > key to choose     or     .

(a)     — This is the measuring mode of electrode potential. mV value  and the temperature will appear at the same time.

(b)     — This is the ORP electrode measuring mode. ORP measurement has no temperature compensation, so there is no temperature appeared in this mode;

Note: — ORP is the abbreviation of “ Oxidation-Reduction Potential ”, signify the Oxidation-Reduction potential of the solution, ORP is the index about the ability of the aqueous solution in the Oxidation-Reduction reaction, its unit is mV.

5.1.2. Connect ORP electrode or ion electrode, immerge into the sample solution, slowly stir and then still. When there appear a “     ” and a stable reading, that is the ORP value or the potential of the ion electrode.

5.1.3. If it is a combination ORP or ion electrode, user just need to insert it into “ pH/mV ” socket, and if it is not combination type, then should also need to choose a proper reference electrode to connect to the “ REF ” socket and measure with both of the two electrodes.

5.2. Considerations:

5.2.1. The meter does not need to be calibrated when do ORP measurement. But if it is interrogative for the ORP electrode quality or the measuring result, it’s better to measure the mV value with the ORP standard solution to identify the accuracy of the ORP electrode or the meter.

Our company can supply the following ORP standard solution: 222mV(25℃)±15mV ORP standard solution (warranty 1 year), the inner solution of reference electrode is 3.5mol/L KCl.

Chart（5-1）

| ℃ mV | ℃ mV | ℃ mV | ℃ mV |
| --- | --- | --- | --- |
| 10 242 | 20 227 | 30 215 | 38 205 |
| 15 235 | 25 222 | 35 209 | 40 201 |

5.2.2. The clean and activation of ORP electrode: the surface of the platinum will be polluted after the ORP electrode has been using for a long time, which will lead to inaccurate measurement and slow response. It could adopt the following method for activation:

(a) For mineral pollution, immerge the electrode into 0.1mol/L dilute hydrochloric acid for 30 minutes, then wash with purified water, and then immerge into the electrode soaking solution for 6 hours before using.

(b) For the organic oil stain pollution, wash the platinum surface with detergent, and then immerge it into the electrode soaking solution for 6 hours before using.

(c) For the seriously polluted platinum with oxidized membrane, polish the surface with toothpaste, then wash it in purified water, and then immerge it into the electrode soaking solution for 6 hours before using.

5.3. Parameter Setting:

5.3.1. mV and ORP measurement parameter setting schedule (chart (5-2))

Chart（5-2）

| Prompt Mark | Parameter Setting Items | Code | Parameters |
| --- | --- | --- | --- |
| P1 | Time for timing measuring setting |  | 0 to 99 minutes |
| P2 | Restore to producer setting |  | OFF-On (shut-set) |

5.3.2. Time for timing measuring setting (P1):

Please see P10 item 4.4.3.

5.3.3. Restore to producer setting（P2）:

Please see P11 item 4.4.7.

6. Conductivity Measurement:

6.1. Preparation Work:

6.1.1. Switch the power and press <     > key to turn on the meter, the meter display   mode;

6.1.2. Press < UNIT > key to choose     （Resistivity）→     （Dissolved Total Solid ）→     （Salinity）→    （Conductivity）;

6.1.3. Install the electrode holder and adjust it properly;

6.1.4. Connect the 2301-M plastic conductivity electrode into the meter.

6.2. Electrode Calibration:

Press < CAL > key, “     ” is flashing on the LCD, indicate enters into the calibration mode, and wash the conductivity electrode with purified water and dry it, then insert it into the 1413μS/cm calibration solution, still it after stirring, when “    ” icon appears, means the display value is stable and again press < CAL > key, LCD will display flashing “ 1.413mS/cm ”, several seconds later the “     ” icon will appear and return to the measuring mode, at this time, LCD will display stable measuring value 1.413mS/cm and calibration icon “     ”, indicates the calibration is finished. If measuring value is unstable, can repeat calibrating until it is stable.

Note: Meter only can be calibrated under the mode of     , and can not be calibrated in mode of     ,     and     , must shift to     mode, and again shift to the original mode after finishing calibration.

6.3. Sample Test:

6.3.1. Wash the conductivity electrode and throw off the water on it, then insert it into the solution, still it after stirring, (If the electrode is installed on the electrode rack to test, then need to short-shaking electrode rack, and take the reading after “    ” appears, which is the conductivity values of this solution.

6.3.2 Press < UNIT > key can display the value of the resistivity, TDS or salinity value which against the conductivity value.

6.4. Important Statement:

6.4.1. This meter build-in below four kinds of calibration solution series: 84μS/cm, 1413μS/cm, 12.88 mS/cm and 111.9 mS/cm. This meter has an unique one-point calibration function, to choose the calibration solution based on the principle of the water samples and calibration solution conductivity as close as possible, in general the most common calibration solution is 1413 μS/cm. Use the equipped 2301 - M conductivity electrode (K = 1 cm-1), and do calibrate with 1413 μS/cm calibration solution, can use within the range of less than 100 mS/cm. Please choose refer to the chart（6-1）.

Chart（6-1）

| Measuring range | 0.05 to 20μS/cm | 0.5μS/cm to 200mS/cm |
| --- | --- | --- |
| Electrode constant | K=0.1cm-1（flow test） | K=1.0cm-1 |
| Calibration solution | 84μS/cm | 84μS/cm | 1413μS/cm | 12.88 mS/cm 111.9 mS/cm |
| Calibration indication icon |  |  |  |  |

6.4.2. There are two kinds of method which are standard solution calibration method and constant calibration method for conductivity electrode calibration which have been set in meter, the statement in item 6.2. is the standard calibration method, and when the accuracy of standard solution is accurate then suggest the standard solution calibration method is priority selected, it can ensure better accuracy. If user used to use constant calibration method, please set the electrode constant in the parameter set of P3. (See P26 item6.5.4.). The two methods can choose freely and won’t influence each other. When adopting the constant calibration method, LCD won’t display the calibration indicator symbol.

6.4.3. The temperature compensation coefficient of the meter setting is 2.0%/℃ by producer. However, the conductivity temperature coefficient is different for solution of different variety and concentrations, the user can refer chart (6-2), as well as the data which they own get in the experiment, to set in the parameters setting of P2.

Note: When the coefficient of temperature compensation is set to be 0, that is, there is no temperature compensation when testing, the measuring value based on current temperature.

Chart（6-2）

| Solution | Temperature compensation coefficient |
| --- | --- |
| NaCl salt solution | 2.12%/℃ |
| 5%NaOH solution | 1.72%/℃ |
| Dilute ammonia solution | 1.88%/℃ |
| 10% hydrochroric acid solution | 1.32%/℃ |
| 5% sulfuric acid solution | 0.96%/℃ |

6.4.4. When the content of parameter setting is consistent with content of producer setting, conductivity measuring interface will display as picture（6-1）, the upper right corner of LCD will display time, when the content of parameter setting is different with content of producer setting, conductivity measuring interface will display as picture（6-2）, the code icon will appear on the upper right of LCD, please be attention when using.

6.5. Parameter Setting:

6.5.1. Conductivity measuring parameter setting schedule（chart（6-3））

Chart（6-3）

| Prompt Mark | Parameter Setting Items | Code | Parameter |
| --- | --- | --- | --- |
| P1 | Electrode constant type selection |  | 0.1, 1, 10 |
| P2 | Temperature compensation coefficient setting |  | 0.00 to 9.99% |
| P3 | Constant calibration setting |  |  |
| P4 | Time for timing measuring setting |  | 0 to 99 minutes |
| P5 | Restore to producer setting |  | OFF-On (shut-set) |

Hh6.5.2. Electrode constant selection（P1）

(a) Press the < MODE > key and meter enters into P1 mode, show as picture (6-3);

(b) Press < ▲ > or < ▼ > key to change the constant setting: 0.1→1→10

(c) Press < MODE > key to enter into the next parameter setting or press < ENTER > key return to measuring mode.

(d) The producer setting of P1 is K=1.

6.5.3. Temperature compensation coefficient setting（P2）

(a) Under P1 mode to short press < MODE > key and enter into P2 mode, show as the picture（6-4）;

(b) Press < UNIT > key to shift, press < ▲ > or < ▼ > key to change the number, the changing range of data is from 0.00 to 9.99;

Attention: When the figure set as 0.00, means there is no temperature compensation, for details please see item 6.4.3;

(c) Press < MODE > key to enter into the next parameter setting or press < ENTER > key return to measuring mode.

(d) The producer setting of P2 is 2.00%.

Attention: If change the temperature coefficient to other numerical value, upper right corner of LCD will display“      ”icon when in measuring mode.

6.5.4. Constant calibration setting（P3）, please refer to item 6.4.2.

(a) Under P2 mode to short press < MODE > key and enter into P3 mode, LCD display last calibration constant, show as picture（6-5）;

(b) Press < UNIT > key, data move to right and flashing, press < ▲ > or < ▼ > key to change number when the number is flashing, and finish setting the constant which marked on the housing of conductivity electrode.

(c) Press < MODE > key to enter into the next parameter setting or press < ENTER > key return to measuring mode.

(d) If need to do calibration for other conductivity electrode which constant K=1, such as conductivity electrode with constant K=10.3, should enter into constant setting P1 to set the constant to be “10”, then enter into P3 mode and set the constant to be 10.3.

6.5.5. Time for timing measuring setting（P4）

(a) Under P3 mode to press < MODE > and enter into P4 mode, see picture (6-6).

(b) Press < UNIT > key, the “  ” will move rightward in turn and flash. Press < ▲ > or < ▼ > key to alter the number.

(c) Press < MODE > key to enter into next parameter setting or press < ENTER > key to conform and return to measuring mode.

(d) The producer setting is “0” second.

Note：in the form “        ”, the number in the front of “ ：” is minute, and the maximum setting is 99; in the back of “ ：” is second, the maximum setting is 59. After setting the regular time interval measuring mode, LCD will display the “     ” icon.

6.5.6. Restore to the producer setting（P5）

(a) Press < MODE > key in mode P4 to enter into mode P5, see picture (6-7).

(b) Press < ▲ > key to choose “     ”, means parameter setting has been restored to the producer setting mode, and will return to measuring mode after 2s.

6.6. Considerations:

6.6.1. The conductivity electrode has been calibrated before meter leaving factory, user can directly use it under normal circumstance;

6.6.2. Recommend calibrating one time every month under the normal circumstances; User should do calibration in time for high accuracy measuring. It is necessary to calibrate the newly purchased conductivity electrode or used after a period time;

6.6.3. Keeping the conductivity electrode clean and wash it with purified water, then throw off the water on it before and after testing. It’s better to rinse electrodes with sample solution;

6.6.4. The surface of 2301-M conductivity electrode is plated with a layer of metal platinum black to used as lowering the electrode polarization, enlarge measuring range, so can not polish the surface of platinum black, only can wash it in water by shaking, in case to damage the platinum black coating; for organic stained things can be washed with hot water with detergent or alcohol.

6.6.5. Conductivity electrode before use can be immersed in water, to prevent the platinum black to be passivated.

6.6.6.When it appears an abnormal reading when calibrating or displaying, please set P5 as “ON” to make the meter restore to producer setting mode, and then do calibrate and measure again.

6.6.7. Do not unplug the adapter when the meter is power on, only to unplug when the meter is power off.

7. Meter’s Complete Kit:

7.1. Model MP521 Lab pH/Conductivity meter

7.2. Model 602 flexible electrode holder

7.3. 201T-M plastic pH/ATC three-in-one combination electrode

7.4. 2301-M plastic conductivity electrode

7.5. pH standard buffer solution                                          1btl/each

（pH4.00, pH7.00 and pH10.01/50mL）

7.6. 1413μS/cm conductivity standard solution（50mL）

7.7. 9V power adapter（meter use）

7.8. RS-232 communication cable

7.9. MP521 communication software disk

7.10. Operation manual

7.11. Brief operation instruction

8. Warranty:

8.1. We warrant this meter to be free of charge maintain, replace the parts or products under normal using circumstances, from purchased time within one year caused by manufacturing bad and unable to work.

8.2. Attached pH electrode does not belong to this warrant range. But, if the newly purchased electrode went wrong without using, it’s free of charge to maintain and replace.

8.3. The above warranty does not apply to defects resulting from action of user such as misuse, improper wiring, operation outside of specification, improper maintenance or repair, or unauthorized modification.

Chart І Meter Parameter Setting Schedule

| Mode | Prompt Mark | Parameter Setting Items | Code | Parameter |
| --- | --- | --- | --- | --- |
| pH | P1 | pH buffer solution series selection |  | USA (Europe & U.S.A. series) NIS (NIST series) CH (China series) |
|  | P2 | Timing measuring setting |  | 0 to 99 minutes |
|  | P3 | Temperature unit setting |  | ℃ ℉ |
|  | P4 | Date setting | Date | Month / Day / Year |
|  | P5 | Time setting | Time | Hour / Minute |
|  | P6 | Restore to producer setting |  | OFF-On (shut-set) |
| mV | P1 | Time for timing measuring setting |  | 0 to 99 minutes |
|  | P2 | Restore to producer setting |  | OFF-On (shut-set) |
| Conductivity | P1 | Electrode constant selection |  | 0.1, 1, 10 |
|  | P2 | Temperature compensation coefficient setting |  | 0.00 to 9.99% |
|  | P3 | Constant calibration setting |  |  |
|  | P4 | Time for timing measuring setting |  | 0 to 99 minutes |
|  | P5 | Restore to producer setting |  | OFF-On |

Chart Ⅱ Meter Restore to Producer Setting Schedule

| Mode | Indication Icon | Parameter Setting Items | Producer Setting Content | Code |
| --- | --- | --- | --- | --- |
| pH | P4 | Time for timing measuring setting | 0 second |  |
| mV | P2 | Time for timing measuring setting | 0 second |  |
| Conductivity | P2 | Electrode constant type selection | K=1 |  |
|  | P4 | Temperature compensation coefficient setting | 2.0% |  |
|  | P6 | Time for timing measuring setting | 0 second |  |

Chart Ⅲ Code Icon and Abbreviation Schedule

Shanghai San-Xin Instrumentation, Inc.

Add: 3/F. Building No.4, No. 471 Guiping Road, Shanghai, China. 200233

Tel: +86-21-63362480               Fax: +86-21-64956880

Website: www.shsan-xin.com        E-mail: wxmab@shsan-xin.com

# Text boxes

3

4

6

6

7

8

9

10

11

11

11

13

13

15

18

19

19

20

21

22

22

22

23

23

25

28

29

29

#### ON

#### OFF

#### ON

#### OFF

1unit

1pc

1pc

1pc

2btls

1pc

1pc

1pc

1pc

1pc

#### ON

#### OFF

#### ON

#### OFF

### Calibration

### End

### Constant Calibration

Constant calibration setting

### Constant

Electrode constant setting

### Temperature Compensation Coefficient

Temperature compensation coefficient setting

### NIST

NIST series standard

### USA

Europe & U.S.A series standard

### China

China series standard

### Solution

Standard solution

### Salinity

### Total Dissolved Solid

### Resistivity

### Oxidation-Reduction Potential

### Conductivity

### Code and Abbreviation

### English

Explanation