Model YD200

Laboratory Water Hardness Meter

Operation Manual

Shanghai San-Xin Instrumentation, Inc.

Table of Content

1. Brief Introduction ------------------------------------------------------------------

2. Technical Parameters ---------------------------------------------------------------

2.1. Definition and Conversion of Water Hardness Unit -----------------------

2.2. Meter’s Technical Parameters -------------------------------------------------

3. Instructions to the Meter ----------------------------------------------------------

3.1. Instruction to the LCD Display -----------------------------------------------

3.2. Instruction to the Operation Key ---------------------------------------------

3.3. Instruction to the Meter’s Socket----------------------------------------------

4. Operation Procedure ---------------------------------------------------------------

4.1. Preparations for Measurement ------------------------------------------------

4.2. Considerations During Measurement Preparation --------------------------

4.3. Meter’s Calibration -------------------------------------------------------------

4.4. Considerations During Meter’s Calibration ---------------------------------

4.5. Water Sample Test --------------------------------------------------------------

4.6. Considerations During Water Sample Test ----------------------------------

4.7. The Preservation of Electrode -------------------------------------------------

5. Meter’s Functional Parameters Setting ------------------------------------------

5.1. Parameter Setting Table ----------------------------------------------------------

5.2. Calibration Way Setting（P1）-----------------------------------------------

5.3. Measurement Resolution Setting（P2）-------------------------------------

5.4. Temperature Unit Setting（P3）----------------------------------------------

5.5. ON/OFF Setting of Activation Compensation （P4）--------------------

5.6. Clock Setting （P5）----------------------------------------------------------

5.7. Date Setting（P6）-------------------------------------------------------------

5.8. ON/OFF Setting of Restoring to Producer Setting（P7）-----------------

6. Storage/Recall/Delete of Measuring Information and RS-232 Communication---

6.1. Store Measuring Information --------------------------------------------------

6.2. Recall Measuring Information ------------------------------------------------

6.3. Delete Stored Measuring Information ----------------------------------------

6.4. RS-232 Communication -------------------------------------------------------

7. Meter’s Complete Kit --------------------------------------------------------------

8. Accessories Options ----------------------------------------------------------------

9. Warranty ----------------------------------------------------------------------------

1. Brief Introduction:

Thanks for your purchasing model YD200 water hardness meter hereinafter called “the meter” from Shanghai San-Xin Instrumentation, Inc.

In order to help you correctly use and maintain this meter, please carefully read this operation manual before using. Based on continually improve the performance of meter, we reserve the right of revise the contents in this operation manual in the case of not notifying in advance.

We call the total concentration of calcium ion (Ca2+) and magnesium ion (Mg2+) in water sample to be the hardness of water. It is an important index for water quality. This meter adopts electrode method to test water hardness, which be proved that it has the same accuracy as the EDTA volumetric method after doing a great deal experiments, meanwhile it can avoid many shortness points of EDTA method, such as complicated operation and judgmental difficulties etc.

There is a microprocessor chip inside the meter conform to international GLP function, the appearance of meter looks beautiful and easily using.

For notable features read as follows:

1.1. Adopts 601-F water hardness combination electrode which consists of measuring electrode, reference electrode and temperature probe. At the same time, its new structure, stable electrode potential, quick response and use easily features are more welcomed by customers.

1.2. The electronic unit adopts intelligent chip design, which has automatic calibration, automatic temperature compensation, data storage; measuring value can be locked automatically and automatic power off and low potential displaying function etc.

1.3. Eight water hardness units can be chosen freely: mmol/L, mg/L (CaCO3), mg/L (CaO), mmol/L (Boiler), mg/L (Ca), ◦fH (France Degree), ◦dH (German Degree) and ◦eH (England Degree), which can choose in the WH-Link communication software.

1.4. It adopts digital filter and step slipping technology to intelligently improve meter’s response speed and result accuracy. “      ” will appear when reading is stable.

1.5. Equips B1, B2 and B3 three calibration solutions.

1.6. LCD display with blue background.

1.7. Equips WH-Link communication software which can transput the data to computer.

1.8. Meter’s circuit board adopts SMT film-covering technology to improve

meter’s production reliability.

2. Technical Parameter:

2.1. Units and Its Conversion of Water Hardness:

mmol/L —basic unit for water hardness

mg/L (CaCO3) — water hardness based on quality concentration of CaCO3, or we call

it ppm and U.S. Degree.

mg/L (CaO) —water hardness based on quality concentration of CaO

mmol/L (Boiler) — special unit for industrial boiler water hardness measuring

mg/L (Ca) —water hardness based on quality concentration of Ca

ºfH—France Degree

ºdH—German Degree

ºeH—England Degree

Units’ Conversion of Water Hardness

|  | mmoL/L | mg/L (CaCO3) | mg/L (CaO) | mmol/L (Boiler) | mg/L (Ca) | ºfH （France Degree） | ºdH （German Degree） | ºeH （England Degree） |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| mmoL/L | 1 | 100 | 56.1 | 2 | 40.1 | 10 | 5.6 | 7.0 |
| mg/L(CaCO3) | 1.00×10-2 | 1 | 0.56 | 2.00×10-2 | 4.01×10-1 | 0.10 | 5.60×10-2 | 7.00×10-2 |
| mg/L (CaO) | 1.78×10-2 | 1.78 | 1 | 3.57×10-2 | 7.15×10-1 | 1.78×10-1 | 9.98×10-2 | 1.25×10-1 |
| mmol/L (Boiler) | 5.00×10-1 | 50 | 28.05 | 1 | 20.05 | 5.00 | 2.8 | 3.50 |
| mg/L (Ca) | 2.49×10-2 | 2.49 | 1.40 | 4.99×10-2 | 1 | 2.49×10-1 | 1.40×10-1 | 1.75×10-1 |
| ºfH (France Degree) | 1.00×10-1 | 10 | 5.61 | 2.00×10-1 | 4.01 | 1 | 5.59×10-1 | 6.99×10-1 |
| ºdH (German Degree) | 1.79×10-1 | 17.9 | 10.02 | 3.57×10-1 | 7.16 | 1.79 | 1 | 1.25 |
| ºeH (England Degree) | 1.43×10-1 | 14.3 | 8.01 | 2.86×10-1 | 5.73 | 1.43 | 8.00×10-1 | 1 |

2.2. Technical Parameter:

2.2.1. Water Hardness:

| Measuring range | （0 ~ 10）mmol/L； （0 ~ 401）mg/L(Ca)； （0 ~ 1000）mg/L(CaCO3)； （0 ~ 100）ºfH(France Degree)； （0 ~ 561）mg/L(CaO)； （0 ~ 56）ºdH(German Degree)； （0 ~ 20）mmol/L(Boiler)； （0 ~ 70）ºeH (England Degree)。 |
| --- | --- |
| Resolution | 0.01 and 0.1 water hardness units |
| Accuracy | ±5% FS |
| Temp. compensation range | （5 ~ 50）℃（automatic） |
| Calibration Solution | B1 Calibration Solution — 2.00×10-2 mmol/L B2 Calibration Solution — 2.00×10-1 mmol/L B3 Calibration Solution — 2.00 mmol/L （B3 Calibration Solution also can be used for the soaking solution of electrode） |
| Calibration Mode | (a) B1/B2 Calibration—using B1 and B2 Calibration Solution，suitable for < 2.00×10-2 mmol/L low concentration water quality，for example, boiler water. (b) B2/B3 Calibration—using B2 and B3 Calibration Solution，suitable for general water quality. |

2.2.2 Temperature：

| Measuring range | 0 ~ 60℃ |
| --- | --- |
| Resolution | 0.1℃ |
| Accuracy | ±0.5℃ |

2.2.3. Other Technical Parameter：

| Data storage | 128 groups |
| --- | --- |
| Content Storage | serial number of measuring value, measuring value, units, temperature, ATC or MTC, date, time and calibration way. calibration way |
| Communication Connector | RS-232 |
| Power | DC9V |
| Dimension and weight | 160×190×70 mm / 750 g |
| Quality and safe certificate | ISO9001:2008 and CE |

2.2.4. working conditions：

| Ambient temperature | （5 ~ 35）℃ |
| --- | --- |
| ambient humidity | ≤85% |

3. Instructions to the Meter:

3.1. Instruction to the LCD Display:

① — Measuring reading display adopts modular type, for example,  means 2.00×10-1.

② — Time and date display

③ — Measuring value lock icon, when this icon appears, it means the measuring value has been locked.

④ — RS-232 communication icon, when this icon appears, it means meter has been connected with PC, RS-232 communication function has been started.

⑤ — Calibration solution indication icon, for example when     flashing, means the electrode should be calibrated with B1 solution.

⑥ — Measuring units, display in turn: mmol/L, mg/L(CaCO3), mg/L(CaO), mmol/L(Boiler) and mg/L(Ca).

⑦ — Temperature compensation icons:

ATC — means automatic temperature measuring and compensation.

MTC —means manual temperature measuring and compensation.

⑧ — Temperature value and temperature unit of the solution.

⑨ — M+: icon for storing measuring value; RM: icon for recalling measuring value; the digit in the left side of icon means storage serial number, meter can store 128 groups of measuring value.

⑩ — Calibration icon, Cal 1 means the first point calibration, Cal 2 means the second point calibration.

3.2．Instruction to the Operation Key:

This meter has eight operation keys in all.

3.2.1.        — Switch button for meter’s ON or OFF.

3.2.2.        — Function button, press this button to enter into parameter setting menu.

3.3.3.        — Calibration button, press this button to enter into calibration state.

3.2.4.        — Compound button for unit and confirmation, when in measuring state, press this button can change the measuring unit. When in in the state of parameter setting. press this button to confirm.

3.2.5.        — Storage button to store the measuring data.

3.2.6.        — Recall button, recall stored measuring data. Press together the buttons      and      to quickly check the stored information.

3.2.7.       — Increase button, in the state of MTC, manually increase temperature, when in the state of parameter setting, this icon can be used as the button for revising function and data.

3.2.8.       — Decrease button, in the state of MTC, manually decrease temperature when in the state of parameter setting, this icon can be used as   functions selecting button.

3.3. Instruction to the Meter’s Back Panel:

3.3.1. “Hardness Electrode” socket: Used for connecting with the model 601-F water hardness electrode (BNC connector), after measuring, should put electrode short circuit protecting cap on to keep socket’s clean.

3.3.2. “Temperature” socket: Used for connecting the model 601-F water hardness electrode, after connecting temperature electrode, meter is in the automatic temperature compensation state, when taking off the temperature sensor, meter is in the manual temperature compensation state, press      or      button to adjust temperature.

3.3.3. “RS232” socket: Connect this meter and PC with RS232 communication cable, measuring information can be sent to PC, and all measuring information can be stored in the internal storage when out of connection.

3.3.4. “DC9V” socket: Used for connecting DC9V adaptor, adaptor equips Φ2.5 inner hole plug, power is inside “ + ” outside “ — ”.

4. Operation Procedure:

4.1. Measurement Preparations:

4.1.1. Electrode’s Activation:

(a) New electrode used for the first time — shake electrode vigorously to make the inner solution of electrode into the compartment of measuring head, there can not be any bubbles in the compartment. Pour a little B3 calibration solution into the “soaking solution” cup, and immerse electrode into it for activating more than 1 hour.

(b) Normal use situation — The activation time is about from several minutes to decade minutes for normally used electrode, as long as the reading is stable then can make calibration and measuring. You can adopt following method to observe: place electrode into B3 solution, when it appears stable icon “     ”, it can be used normally. (Attention the display value is faculative).

4.1.2. Pour a little B2 and B3 calibration solution into the test cup of “B2” and “B3” for using (take B2/B3 calibrating way as an example).

4.1.3. Choose meter’s calibrating way, resolution, water hardness units and other measuring parameters.

4.2. Considerations During Measurement Preparations:

4.2.1. In order to extend use life of electrode sensitive membrane, there do not equip measuring head for new electrode when promote it into market, instead of it is a sensitive screw plug, and replace the measuring head on when use. Do not screw off the measuring head when use it normally, and change the second standby one when electrode invalid.

4.2.2. After calibrating, user can also use B3 as soaking solution.

forcefully, to make the reference solution into the measuring head.

4.3. Meter’s Calibration: (take B2/B3 calibration method and mmol/L unit as an example )

4.3.1. Press        ,      and        icons appear and flashing at the same

time, indicates entering into first point calibration, calibration solution is B2,

(see picture (4-2).

4.3.2. Wash and shake electrode forcefully, put it

into B2 Calibration solution, still it after

shaking, when    appear, press

button, the 2.00×10-1  and  mmol/L icon

flashing on the LCD (see picture (4-3),

calibration is finished after 20 seconds,

and  icons appear and flashing

simultaneity, indicates enter into second point

calibration, calibration solution is B3.

4.3.3. Wash and shake the electrode forcefully, and

place it into B3 calibration solution, stir and

still, when     appear, press       button,

the 2.00 and mmol/L icons flashing on the LCD (see picture (4-4)), calibration is finished after 20 seconds, meter enters into measuring mode.

4.3.4. Wash electrode with purified water and shake it forcefully, then put it into B2 solution, stir and still, and take reading after displaying     icon, normal displaying value is equal or less than 2.00 ±0.25×10-1 mmol/L, if not, repeat calibration in solution of B2 and B3.

4.4. Considerations During Meter’s Calibration:

4.4.1. The new electrode for the first time to use need to be calibrated for two times. Sometimes there will be deviation for the displayed value, even just the unchanged value “1.00×103mmol/L” (Said overflow display), this is normal and can be corrected in the second calibration.

4.4.2. The concentration of B1, B2 and B3 calibration solution is from thin to thick, so the electrode does not need to be washed when using, as long as calibration accordingly and throwing solution on electrode off, if the concentration of solution from thick to thin, electrode need to be washed with high purified water very clear, otherwise it will affect measuring accuracy.

4.4.3. Calibration using two-point calibration method, calibration curve will be established after two-point calibration. Therefore, the correct values will be displayed only after calibration.

4.4.4. Because meter does not have recognition ability for the three calibration

solutions, so user must place electrode into right calibration solution according

to the ,   and    icons flashing on the LCD, otherwise, it will occur

biggish calibration error, if the above situation happens, then user need to

repeat calibration.

4.4.5. Because B1, B2 and B3 calibration solution have no buffer ability, and are easy to be polluted, especially for the low concentration B1 solution. So solution put into test cup generally can be used 4 to 5 times, then need to be changed, and they need to be paid much attention to the clear and be prevented to be dirty, otherwise it will cause big measuring error.

4.5. Sample Test:

Wash electrode with purified water and shake it forcefully, put it into sample solution, stir and still, when the     icon appear, and read the value. Shown in picture (4-5) means: calibrating method is B2 and B3, temperature of solution is 25℃, measuring value is 2.88×10-1, unit is mmol/L, press       can convert it

to other units.

4.6. Considerations During Sample Test:

4.6.1. Electrode needs to be washed clean before measuring sample solution. The

container for diluted and concentrated solution should be separated, and can

not be mixed used together. When testing, try to test diluted solution first and

then to test concentrated solution.

4.6.2. Generally, calibration is necessary before measuring, but for shorter interval test, it does not need, or put the electrode into the calibration solution close to the measured solution, if reading within error range, then it unnecessary to do calibration.

4.6.3. When testing boiler water, should choose B1 and B2 calibrating method, and add Ionic Strength Buffer (TISAB) into B1, B2 calibration solution and soaking solution, otherwise it will cause big measuring error; the proportion is add 50.4mg analysis pure Sodium Bicarbonate (NaHCO3) per 100mL solution.

4.6.4. According to the ion electrode’s isothermal measurement principle, if the temperature of sample solution is much closer to the calibration solution, then the measuring accuracy will be higher. So for generally testing, the temperature discrepancy between calibration solution and sample solution should ≤± 15°C; for high accuracy testing, should keep the temperature of calibration solution the same as sample solution.

4.6.5. The sensitive membrane of electrode might become turbid and blanched after using, this is a normal phenomenon.

4.6.6. During using, if you find the PVC sensitive membrane in the electrode measuring head is seriously sunken, please first screw it loose, then screw it tight, and then shake it forcefully for several times to eliminate the bubble.

4.6.7. When the concentration of sample solution is bigger than 1.00×10mmol/L,  icon will display in the left-bottom corner, indicating to dilute the sample solution10-fold before test, then multiple the test value by 10 to get the correct one.

4.7. Electrode Preserving:

Electrode head should be kept with protection cap after washing and dry, there is a small piece of water storage sponge in the bottom of electrode cap, please dip a little B3 solution into it (make sponge full of water, but water can not flow out), keep electrode in moist condition.

5. Meter’s Functional Parameter Setting:

5.1. Parameter Setting Table

| Prompt | Parameter setting items | Parameter |
| --- | --- | --- |
| P1 | Calibration Method Setting | -------- |
| P2 | Resolution Setting | 0.01-0.1 |
| P3 | Temperature Unit Setting | ℃-℉ |
| P4 | Activity Compensation Setting | OFF-On |
| P5 | Clock Setting | Hour- Minute |
| P6 | Date Setting | Month –Day-Year |
| P7 | Default Setting | OFF-On |

5.2. Calibration Way Setting (P1):

Short press the     button, meter goes into calibration way setting P1. LCD display as picture (5-1), press        button, meter alternately change

and        two calibration ways. Press

button to confirm and return to measuring mode, the producer setting is        .

5.3. Measurement Resolution Setting (P2):

Short press        button under the P1 state, meter enters into measuring resolution setting P2. LCD display as the picture (5-2), press      button to alternately change resolution 0.01 and 0.1, press       button to confirm and return to measuring mode, the producer setting is 0.01.

5.4. Temperature Unit Setting (P3):

Short press        button under the P2 state, meter enters into temperature unit setting P3. LCD display as picture (5-3), press      button alternately change temperature unit °C and °F, press       button to confirm and return to measuring mode, the producer setting is °C.

5.5. ON/OFF Setting of Activation Compensation (P4):

Short press       button under the P3 state, meter enters into ON/OFF setting of activation compensation P4. Picture (5-4) is producer setting “ON”, press      button alternately change temperature compensation button OFF and ON, press        button to confirm and return to measuring mode. When choosing ON, means within the high concentration measuring range of 4.00~10.00mmol/L, and the displaying value has activation compensation, when choosing OFF, the displaying value does not has activation compensation.

5.6. Clock Setting (P5):

Short press       button under the P4 state, meter enters into clock setting P5. LCD display as picture (5-5), press      button to alternatively choose number of hour and minute, press       button change number to set clock, press        button to confirm and return to measuring mode.

5.7. Date Setting (P6):

Short press        button under the P5 state, meter enters into date setting P6. LCD display as picture (5-6), press     button alternately choose the number of month-date-year, press      button change number to set date. Picture (5-6) shows June 16th 2006, press

button to confirm and return to measuring mode.

5.8. ON/OFF Setting of Restore to Default Setting (P7):

Short press       button under the P6 state, meter enters into ON/OFF setting of restore to default setting state P7. The default setting is OFF, LCD display as picture (5-7), press       button to switch ON and to select default settings, then after 2s, it enter into the measuring mode. At this time, P1 to P4 all restore to default setting state and calibration point restore to the theoretical design state.

6. Store/Recall/Delete of Measuring Information and RS-232 Communication:

6.1. Store Measuring Information:

6.1.1. In the state of measurement, short press        button after reading to be stable, a storage number and M+ icon appear on the LCD, which means the measruing information has been stored under this number, measuring information includes measuring time and date, measuring value number, measuring value, unit, temperature and calibration way, as the picture (6-1) showing that a group of measruing information which under the serial number of 28 after short pressing        button.

6.1.2. LCD will display the last time stored measuring serial number and M+ icon, as the picture (6-2) showing         on the LCD, means the meter has already stored 16 groups of measuring information, the meter can store 128 groups of measuring information in all.

6.1.3. Again press       button when stored reach at 128 groups, the icon      will appear on the LCD, as the picture (6-3) showing, if press the button        at this time, meter will clear original 128 groups information, and new measuring value will be stored, and icon     appears on the LCD. If press the button       , meter will keep original 128 groups of information and restore to measuring mode.

6.2. Recall Measuring Information:

In the state of measruing, press        button, meter recall last stored information, the measuring time and date will display alternatively. Again press

button, meter will display measruing information under serial number 27, or press      or      button, can quickly check the measuring information under other serial numbers.

6.3. Delete Stored Measuring Information:

Under the state of measuring, depress        button, the flashing icon

will appear on the LCD,at this time, press the       button, meter will clear all stored measuring information and return to mearuing mode. If press the

button, meter will leave origianl measuring information and return to measuring mode.

6.4. RS-232 Communication:

This meter adopts WH-Link communication software, can achieve RS-232 communication function. This application software requests PC which can stable run Windows 2000/XP operation systerm. Connect the PC and this meter with RS-232 communication calbe, and all stored informaiton by meter will be sent to PC after pressing        button, Meanwhile, the      icon flashing on the LCD, ask if will clear stored measuring information, if want to save it, please press        button, if want to clear it, please press        button, meter return to measuring mode, the RS232 commnication icon     appears on the LCD. At that time, again press        button, the measuring information will be sent to PC all through RS-232, will not be stored in the meter.

The water hardness unit shows on WH-Link is basic unit mmol/L, user can choose other units and interface will automatically switch and display required unit.

After you pressing the “Transfer” button on the interface of WH-Link, a new window will jump on the soft display interface, shows the path of how to store data and file name, user can revise path and file name, and press “Save” button, then PC will transfer all information on this software to Microsoft Excel file and save it, then you can edit, analysis, print the measuring information freely.

7. Meter’s Complete Kit:

7.1. Model YD200 water hardness meter

7.2. Model 601-F water hardness electrode

(measruing head has been fixed on the electrode)

7.3. Model 602 Flexible electrode holder

7.4. 250mL water hardness calibration solution (B1, B2 and B3)

7.5. 20mL test cup

7.6. Test cup pedestal

7.7. WH-Link communication software

7.8. RS-232 communication cable

7.9. Operation manual

8. Accessories Options:

8.1. Model 601-F water hardness electrode

8.2. Water hardness calibration solution (B1，B2 and B3)

Specs: 500mL/bottle and 250 mL/bottle

8.3. 20 mL test cup

8.4. Test cup pedestal

9. Warranty:

9.1. We warrant this instrument to be free of defects in parts and workmanship for one year from date of shipment.

9.2. Attached water hardness electrode does not belong to this warranty range, besides the new one.

9.3. This warranty does not apply to defects resulting from action of the user such as misuse, improper wiring, operation outside of specification, improper maintenance or repair, or unauthorized modification.

9.4. The meter implements the maintenance during all the using periods. Users can send the meter to our factory to repair. Users pay the freight and insurance.

Shanghai San-Xin Instrumentation, Inc.

# Add: 3/F. Building No.4, No. 471 Guiping Road, Shanghai, China. 200233

Tel: +86-21-63362480               Fax: +86-21-64956880

Website: www.shsan-xin.com        E-mail: sanxin@shsan-xin.com

# Text boxes

2

3

3

4

5

5

7

8

8

8

9

9

10

11

11

12

12

12

13

13

13

13

14

14

14

15

15

16

16

17

181818

UNIT

READ

ON

OFF

MODE

CAL

UNIT

READ

M+

UNIT

READ

UNIT

READ

UNIT

READ

MODE

UNIT

READ

UNIT

READ

MODE

MODE

UNIT

READ

MODE

UNIT

READ

MODE

ON

OFF

M+

M+

RM

M+

M+

RM

RM

UNIT

READ

RM

RM

1set

1pc

1set

1btl/each

4pcs

1pc

1pc

1pc

1 book

RM

RM

MODE