LABORATORY

MULTIFUNCTION

METER

CX-505

USER’S MANUAL

USER’S MANUAL

LABORATORY

MULTIFUNCTION METER

CX-505

Before use please read the instruction carefully!

TABLE OF CONTENT

I.	INTRODUCTION	7

1.	EXPLOITATION NOTICES	3

2.	THE CHARACTERISTICS OF THE METER	4

3.	WHAT IS THE METER DESIGNED FOR	5

4.	OUTSIDE VIEW	6

5.	SWITCHING THE METER ON AND OFF	9

6.	PREPARATION TO WORK	10

6.1.	Choosing the kind of temperature compensation	10

6.2.	Changing the resolution of the measurements	11

6.3.	Changing the electrode number	12

6.4.	Readout of the date of last calibration	13

6.5.	Entering the calibration validity date	14

II.	pH MEASUREMENT	15

7.	PREPARATION OF THE pH ELECTRODE	17

8.	CALIBRATION	18

8.1.	Calibration modes	19

8.2.	Calibration with use of NIST sample solutions	20

8.3.	Calibration with introducing the pH value of the sample solutions	21

8.4.	Entering the calibration parameters	22

8.5.	pH electrode calibration	24

9.	CHECKING THE ELECTRODE CONDITION	27

10.	pH MEASUREMENT	28

10.1.	Measurement with automatic temperature compensation	28

10.2.	Measurement with manual temperature compensation	29

11.	NOTICES ABOUT THE TEMPERATURE COMPENSATION AND

INTERPRETATION OF THE pH MEASUREMENTS RESULTS	30

III.	CONDUCTIVITY AND SALINITY MEASUREMENT	33

12.	BASIC INFORMATION ABOUT THE CONDUSTIVITY MEASUREMENT	35

13.	ENTERING THE CONDUCTIVITY METER PARAMETERS	36

13.1.	Choosing the unit	36

13.2.	Entering the WTDS coefficient	39

13.3.	Entering the  coefficient value	40

13.4.	Entering the reference temperature value	41

14.	CHOICE AND MAINTENANCE OF THE CONDUCTIVITY CELL	42

14.1.	The cell choice	42

14.2.	The conductivity cell maintenance	43

15.	CALIBRATION	44

15.1.	Calibration without the sample solution	44

15.2.	Calibration with use of the sample solution	45

16.	THE CONDUCTIVITY MEASUREMENT	48

16.1.	Measurement without the temperature compensation	48

16.2.	Measurement with automatic temperature compensation	49

16.3.	Measurement with manual temperature compensation	50

17.	SALINITY AND TDS (TOTAL DISSOLVED SOLIDS) MEASUREMENT	51

17.1.	Salinity measurement with conversion to NaCl or KCl content	52

17.2.	Salinity measurement with conversion to TDS content	52

17.3.	Determining the WTDS coefficient	53

18.	SIMPLIFIED WAY OF DETERMINING THE  COEFFICIENT	54

IV.	DISSLVED OXYGEN MEASUREMENT	57

19.	BASIC INFORMATION ABOUT THE DISSOLVED OXYGEN MEASUREMENT	59

20.	DISSOLVED OXYGEN SENSOR	61

21.	ENTERING THE OXYGEN METER PARAMETERS	61

21.1.	Change of the unit	61

21.2.	Compensation of the salinity influence	62

21.3.	Manual entering of the salinity value	62

21.4.	Automatic compensation of the air pressure influence	64

22.	CALIBRATION OF THE DISSOLVED OXYGEN SENSOR	65

23.	OXYGEN CONCENTRATION MEASUREMENT	67

23.1.	Measurement with automatic temperature compensation	67

23.2.	Measurement with manual temperature compensation	68

V.	ATMOSPHERIC PRESSURE MEASUREMENT	69

24.	ATMOSPHERIC PRESSURE MEASUREMENT	71

VI.	ORP (mV) AND TEMPERATURE  MEASUREMENT	73

25.	ORP (mV) MEASUREMENT	75

26.	TEMPERATURE MEASUREMENT	75

VII.	OTHER	77

27.	CLOCK WITH DATE	79

27.1.	Clock	79

27.2.	Date	79

27.3.	Setting the time and date	80

27.4.	The LCD backlight mode	80

27.5.	Brightness control	81

27.6.	Readout of the software version number	82

28.	STORING AND READOUT OF THE RESULTS	83

28.1.	Parameters of storing and readout from the memory	83

28.2.	Storing the single measurements to the memory	85

28.3.	Storing the measuring series	85

28.4.	Reviewing the results	86

28.5.	Deleting the stored results	86

29.	CO-OPERATION WITH A PC	87

30.	TECHNICAL DATA	88

31.	EQUIPMENT	91

##### INTRODUCTION

# EXPLOITATION NOTICES

Dear User!

We present you a device distinguished by accuracy according to the technical data and by a high stability of the displayed results. We believe that the measurements will not cause you any trouble and that the meter will operate without any inconvenience. Wide range of additional functions requires careful reading of the manual in other case some of the features may stay unused or using the meter may cause you a troubles.

The employing of good-quality electrodes cells and probes and their replacing after a suitable time ensures obtaining of high measuring parameters. We want to call your attention to the fact, that this equipment has a much shorter working life than the meter. Typical symptoms of an improper operation of the electrode are deterioration of final result stability, it’s flowing as well as a higher measuring error. Part of the users has problems, arising from employing electrodes not being preconditioned before the measurement or making the measurements without removing the shielding ring from the liquid junction or taking measurements with a plugged junction. To avoid this situation it is necessary to choose the proper kind of electrode for solutions which are going to be measured ex. special electrodes for the sewage, liquids with deposits, meat `cheese etc. Therefore, if you observe improper operation of the device, please take control measurements with another electrode or check the used electrode with another pH-meter. Generally the deterioration of the meter work is caused by the electrode and not by the meter.

In case of conductivity measurements it is important to choose the cell with right constant K value for the measuring range. Wrong selection may cause larger error occur, similar situation occurs during measurements with automatic temperature compensation with incorrectly introduced ( coefficient.

Accuracy of the dissolved oxygen measurements depends on the sensor calibration and regular maintenance which consist in replacing the membranes, electrolyte and cleaning the electrodes. Lack of correct maintaining of the sensor after some time will make measurements impossible. Please turn your attention to the fact that stabile measurement is possible only with simulated or natural measured water flow.

The essential feature of our  products is their low failure frequency. However if your pH-meter will fail, our firm immediately performs its warranty repair.

We wish you a pleasant and trouble-free work with our meter.

# THE CHARACTERISTICS OF THE METER

The multifunction meter CX-505 belongs to the newest generation of measuring devices which offer wide range of additional functions. The meter ensures high accuracy and repeatability of the readings. The newest generation electronic elements used in the meter made it’s memory independent to power supply. The internal clock is  powered from battery which is enough for many years. The meter is equipped with custom, backlight LCD display, which enables simultaneous observing of the measured function, temperature value and additional symbols which make working easier.

Main features of the CX-505 are:

- high accuracy and stability of the reading;

- automatic and manual temperature compensation;

- electrode calibration in 1 to 5 points;

- automatic recognition of pH buffers and standards;

- imposed values of standard solutions with possibility of changing their value;

- option of automatic introduction of temperature correction on the value of pH buffer solutions (NIST norm);

- information about the condition of the electrode;

- remembering the date and parameters of last calibration of three sensors (electrodes) for each measuring function;

- wide range of conductivity measurement with 6 automatically switched subranges (autorange);

- counting the conductivity to salinity in NaCl or KCl according to real dependence to conductivity;

- possibility of introducing the TDS coefficient;

- calibration of the conductivity cell by introducing the constant K or on sample solutions;

- function of determining the constant K of cell;

- automatic compensation of the salinity influence on the oxygen measurement using the conductivity measurement function;

- measurement and automatic compensation of the atmospheric pressure influence on the dissolved oxygen concentration;

- remembering the measurements results with time, date, individually or taking series of measurements with set time interval;

- USB port;

- real clock and date.

# WHAT IS THE METER DESIGNED FOR

Multifunction meter is precise and easy to use meter designed for hydrogen ion concentration measurements in pH units, Oxidation Reduction Potential (mV) conductivity in S/cm or mS/cm, dissolved oxygen in water in % of saturation or mg/l and atmospheric pressure measurement in hPa. The meter may be also used for accurate temperature measurement of solutions and air in °C.

The conductivity measurement result may be also displayed in concentration units (g/l or %) counted to NaCl, KCl or TDS. The meter is being used in food, chemical, pharmaceutical, energetic industries, in water treatment stations, laboratories, agriculture, universities, scientific laboratories etc.

CX-505 is prepared to work with all types of combination pH electrodes and conductivity cells, with wide constant K range, equipped with BNC-50 connector. The meters co-operate with Pt-1000 temperature probe with Chinch connector.

The meter may collect measurements taken as single or series of measurements with set time interval.

The built in USB port enables sending the data to the PC. There is possibility of sending current results of measurements or collected in the meter’s memory together with date and time of their collecting.

In case of necessity to collect series longer than the memory capacity results it is possible to use special PC software offered by our company.

# OUTSIDE VIEW

In the upper part of the meter there is a LCD display (Pic. 1), on which depending on the chosen function following symbols are displayed:

- result of the conductivity or salinity measurement;

- result of the pH measurement in pH units;

- result of the mV measurement in mV;

- result of the oxygen measurement in % or mg/l

- time and date.

Choosing the function is done by specific buttons and is signalised by lighting LED diode placed on the button.

Simultaneously with the result a measured temperature value is displayed in oC. Symbols of the units are displayed by the results. In the oxygen measurement mode there is possibility of reading the atmospheric pressure value.

Pic. 1.

Beside the temperature value a  symbol for automatic temperature compensation is displayed or  for manual. CAL symbol on the left side of the display informs that the meter is in calibration mode. On the left side of the display a number of chosen electrode is showed (E1; E2; E3) which inform about the number of characteristic which will be taken into consideration during all calculations. Deleted characteristic, validity time of electrode calibration exceeded and (only in pH function) lost pH electrode efficiency from last calibration is signalised  by flashing symbol of electrode. During calibration between the upper and lower row of numbers number of the calibration point is displayed ( P1, P2, P3, P4, P5).

In the MODE all parameters introduced by the user are displayed and also the value of atmospheric pressure measurement (in oxygen measurement function).

The keyboard  (Pic. 2) placed under the display is used for switching the meter on and off, choosing the measuring function, calibration, entering the parameters, printing and storing the results in the memory.

-	switches the meter on and off;

-	chooses the pH measuring function;

-	chooses the mV measuring function;

-	chooses the conductivity and salinity measuring function;

-	chooses the dissolved oxygen measuring function;

-	displays the time and date;

-	longer pressing of this button enters the calibration mode (CAL symbol displayed). Short pressing in this mode confirms the calibration result;

-	short pressing causes storing single results or measuring series, longer pressing enters the stored results readout mode;

-	chooses the entered parameter;

, 	-	buttons for entering the parameters.

On the back wall of the meter there are inputs with below given symbols:

pH/mV	-	BNC-50 connector for connecting combined or measuring pH electrode and ORP or ion selective electrodes;

Gnd	-	connector for reference electrode;

cond	-	BNC-50 connector for conductivity cell;

O2	-	connector for dissolved oxygen sensor;

temp	-	Chinch connector for temperature probe;

USB	-	USB port for connecting the meter with the PC;

POWER	-	connector for power adapter.

Pic. 2.

# SWITCHING THE METER ON AND OFF

The meter is switched on by pressing the  button. The meter tests the memory and display on which all symbols are displayed (Pic. 3).

Pic. 3.

If the test was successful, after about 1.5 s the meter switches it self automatically to the measuring mode, in which it was switched off. If an  sign will be displayed it means that the meter has lost the factory settings and requires the service repair. If after the 1,5 s all symbols will be continuously displayed it informs that the calibration parameters of electrodes or cells  were lost.

After pressing the  button the meter will take standard characteristics:

- shift = 0 pH, characteristic slope = 100% for pH electrode;

- constant K = 1.000 cm-1 for conductivity cell;

- shift = 0% O2, characteristic slope = 100% O2 for oxygen sensor.

And will enter the measuring mode. It will be necessary to calibrate the conductivity cell, pH electrodes and oxygen sensor.

The meter is switched off by pressing and longer holding of the  button.

# PREPARATION TO WORK

Before starting the work one should:

- join the power adapter plug to the Power input;

- to pH/mV input (BNC-50) join the ready to work combination pH electrode, ion selective electrode or redox electrode;

- to cond input (BNC-50) join the conductivity cell;

- to O2  connector join the dissolved oxygen probe;

- in case of using the temperature probe it should be connected with the chinch temperature input - temp;

- in case of work with the PC join the cable with the USB port;

- switch the meter on by pressing the  button.

The pH electrode is isolated from the conductivity sensor, therefore during the pH and conductivity measurements the pH electrode and conductivity sensor may be plunged together in the same solution.

## Choosing the kind of temperature compensation

The meter switches it self to the automatic temperature compensation mode automatically after joining the temperature probe, after disconnecting it the meter enters the manual temperature compensation mode. In ATC mode near the displayed temperature a  symbol appears. Manual temperature compensation is indicated by  symbol near the value entered by the user, it’s value may be changed using  or  keys.

## Changing the resolution of the measurements

The measurement results may be displayed with chosen resolution. To change it one should:

-	in the measuring mode press the   button, a  (resolution) sign will be displayed. (Pic );

-	using keys  or  one may choose:

- (low) resolution of the measurement;

- (high) resolution of the measurement.

Pic 4

For the pH measurement:

- resolution of the measurement 0.01 pH;

- resolution of the measurement 0.001 pH.

For the conductivity measurement:

- resolution of the measurement 3½ digits;

-  resolution of the measurement 4½ digits.

For the oxygen measurement:

- resolution of the measurement 1% or 0,1mg/l;

- resolution of the measurement 0.1% or 0.01 mg/l.

There is no possibility of change the resolution in mV measuring mode.

Return to the measuring mode by pressing the chosen function button.

## Changing the electrode number

If in the meter’s memory more than one electrode characteristic was stored, it is possible to replace the electrodes without the need of calibration. This option is quite useful in case of using different kinds of electrodes ex. for sewage, and clear water etc. It is necessary to connect the earlier calibrated electrode, marked with the number which responds to the number which was stored in the memory and choose this number.

To do so in the measuring mode:

-	press the   button till the moment of displaying in the upper row of the display the ,  or  symbol, than using the  ,  buttons choose the electrode number. Under this number the calibration results will be stored (Pic. ). Above the electrode number following signs will be displayed:

-	under this number there is no characteristic stored and producers values are provided. In the measuring mode the electrode number will be flashing.

-	under this number there are values of last calibration stored.

- the last calibration showed that the electrode is loosing it’s efficiency and in short time it’s calibration won’t be possible. (only for electrode calibrated in pH function). In the measuring mode the electrode number will be flashing.

Additionally under the electrode number the points in which the electrode was calibrated will be showed.

Pic. . 5.

-	return to the measuring mode by pressing the chosen function button.

## Readout of the date of last calibration

The meter remembers dates of calibration of all electrodes and sensors in every measuring function. Before starting the work one may check the date of last calibration.

To do so in the measuring mode:

- press the   button till the moment of displaying in the upper row of the display the electrode number (,  or  symbol);

- using the  ,  buttons choose the electrode number which we calibration date we want to check and shortly press the  button. The date of last calibration will be displayed in the form: month – day and below the year (Pic. ). On the left side is the number of the electrode displayed.

Pic. 6.

The meter stores the date in the moment of calibration. If after the calibration the date will be changed the meter will signalise with error the exceeding of the calibration validity, so it is necessary to remember about setting the correct date before the calibration.

Return to the readout of the electrode number after pressing the  button, or to measuring mode by pressing any of the function buttons.

## Entering the calibration validity date

In each measuring function the meter stores the validity date of the calibration for each of the electrodes or sensors. Exceeding the set time the meter will signalise with blinking number of the electrode (symbols ,  or ), informing that the electrode calibration characteristic stored under this number is no longer valid and the electrode should be calibrated.

To set the next calibration validity time one should in the measuring mode of the chosen function:

- Press the  button till the moment of displaying in the upper row of LCD a  (time out) symbol. Below the earlier introduced number of days till the next calibration will be displayed (Pic. 7);

- With  ,  buttons enter the requested number of days till the next calibration. If after choosing 1 day time one will press the  button a  symbol will be displayed instead of the digits – this will block the function of calibration time control.

Pic. 7.

Return to the measuring mode by pressing any of the function buttons.

##### pH MEASUREMENT

# PREPARATION OF THE pH ELECTRODE

The electrode should be prepared to work according to the producer instructions. If the instructions weren’t given please follow the steps:

-	new electrode should be put into distilled water or in saturated KCl solution for about 3 hours:

-	Before starting the measurements protecting rings (if used in this kind of electrode) should be removed. The ring placed on the junction – lower part of the electrode should be moved up along the electrode’s body and the upper, which protects the KCl refilling hole, down along the body.

Removing the lower ring is essential, in other case the electrode won’t measure.

Upper ring should be removed during measurements of high temperature solutions or to protect the junction during measurements in solutions with deposits or oils. Sometimes instead of ring a cork is used;

-	during measurements in laboratory it is advisable to use an electrode holder;

-	after every measurement the electrode should be washed in distilled water;

-	excess of liquid on the electrode should be removed by gentle touching the glass with a tissue paper;

-	after work the electrode should be stored in saturated KCl solution what prolongs it’s life. The protecting rings should be moved on the junction and upper hole;

-	in case of long breaks between the measurements the electrode should be stored, after drying, in the package;

-	after taking the electrode of the package the eventual deposit should be removed using water;

-	before using the electrode it should be placed in distilled water for about 1 hour;

-	if the construction of the electrode enables refilling the electrolyte, it should be controlled and refilled periodically by the upper hole in the electrode’s body (usually as the electrolyte a saturated KCl solution is used).

CAUTION:	storing of the electrode in distilled water shortens it’s life time and may cause shifting of the zero point.

# CALIBRATION

Before starting the measurement with new electrode, after long-lasting using, or before making measurements which require higher accuracy the electrode connected with the meter should be calibrated. The results of measurements done without calibration will have a great error. The calibration is done in the buffer or sample solutions with known value. It is based on comparing pH value of the buffer solutions with the value displayed by the meter and next automatic introduction of correction to the meter’s memory and taking it into consideration during measurements. The calibration should be periodically repeated because during work the parameters of the electrode are changing what influences the accuracy. The frequency of this procedure depends on the demanded accuracy, number of the measurements carried out, conditions in which the electrode was used, temperature and value of the measured solutions. When the highest accuracy is required it is recommended to use certified sample solutions. In technical usage buffer solutions are used with lower accuracy but mostly with total values ex. 2.00 pH; 4.00 pH etc. In this manual to make it easier a buffer solution name was used for both buffer and sample solutions. For accurate measurements it is necessary to use fresh good quality sample solutions.

The first action taken before the calibration is entering the values of used pH buffer solutions to the meters memory. This should be done before the first calibration and repeated in case of changing the used buffers to other with different values.

During the calibration after immersing the pH electrode and temperature probe the meter automatically detects the value of the used buffer, if it was earlier stored in the meter’s memory.

There is possibility of calibration in minimum 1 buffer solution and maximum in 5. The more points of calibration is used the higher is the accuracy in the whole measuring range.

Calibration in one buffer solution can’t ensure high accuracy. If only one buffer solution is used it’s value should be close to the forecasted value of measured solution. If the required accuracy isn’t very high and the measurements will be done in the whole range the one point calibration should be done with buffer solution close to 7.00 pH. Thanks this the so called zero electrode shift will be eliminated. In other points a standard characteristic will be taken from the meter’s memory. This characteristic corresponds with theoretical efficiency of the pH electrode.

If measurements are done both in acids and alkalis and the measurements are not done at the ends of measuring range it is enough to calibrate the electrodes in 3 buffer solutions with values in range given in table 2 – calibration points 2, 3 and 4. In case of accurate measurements in the whole range it is recommended to calibrate the electrode in all 5 points given in the table. In CX-505 the characteristic of the electrodes is approximated linearly between the calibration points.

Starting the calibration under one of the electrode numbers does not remove the set buffer values, but irreparably removes the electrode characteristic stored under this number.

There is no possibility of calibrating only one point and leaving the rest of data from the last calibration.

The memory in CX-505 enables independent storing of 3 characteristics of pH electrodes. This feature is especially useful when fast replacing of the electrodes is necessary or when replacing a broken one. Those electrodes should be earlier calibrated and brought into memory under following symbols (,  , ).

The order of using the buffer solutions is freely chosen.

## Calibration modes

The user may choose two independent ways of action:

- use the values of pH sample solutions entered to the memory by the manufacturer. Those values are in conformity with the NIST norm. This type of calibration switches on the automatic correction introduction connected with the change of sample solution along with temperature changes.(description in chapter 8.2).

- enter the value of pH buffer solutions depending on the actually used buffers, in the range given for each point of calibration; (chapter 8.3).

Choosing the first option makes the calibration much easier. The  user doesn’t need to warm up or cool down the sample solutions to the temperature given by the user. This option includes some simplifications and when very high accuracy is required it shouldn’t be used. The meter stores independently 5 values of the buffers for each of the calibration modes described above.

## Calibration with use of NIST sample solutions

In this mode 5 constant sample solutions values, according to NIST, are used. In the meters memory a table with dependence between the temperature and pH values for this 5 sample solutions is stored. This dependence is showed in the table 1. During the calibration the temperature of sample solution is measured and on it’s basis automatically the right pH value responding to this temperature is taken into consideration. Values between the points given in the table are approximated linearly by the meter.

Some sample solutions made according to the appropriate norms may differ on the third decimal place. The meter enables setting of the appropriate buffer solution values in range  ±0.010 pH from the values given in the table for 20(C. The way of entering the buffer solution value was described in chapter 8.4. In case of differences larger than ±0.010 pH the user may not use the mode with use of NIST sample solutions but the mode which enables setting the buffers value and before the calibration set, in the meter’s memory, the accurate value of the used sample solution according to pt. 8.3.

The range of the temperatures, taken into consideration during introducing the correction, is 0 to 60 oC and can’t be exceeded during calibration.

The way of action during calibration way choice and entering sample solutions values other than those in table 1 was described in chapters 8.4.

Table 1.

| Temp. 0C | Kind of buffer solution |
| --- | --- |
|  | 1 oxalate | 2 phthalate | 3 phosphate | 4 di-sodium tetraborate | 5 calcium hydroxide |
| 0 | 1.666 | 4.003 | 6.984 | 9.464 | 13.423 |
| 5 | 1.668 | 3.999 | 6.951 | 9.395 | 13.207 |
| 10 | 1.670 | 3.998 | 6.923 | 9.332 | 13.003 |
| 15 | 1.672 | 3.999 | 6.900 | 9.276 | 12.810 |
| 20 | 1.675 | 4.002 | 6.881 | 9.225 | 12.627 |
| 25 | 1.679 | 4.008 | 6.865 | 9.180 | 12.454 |
| 30 | 1.683 | 4.015 | 6.853 | 9.139 | 12.289 |
| 35 | 1.688 | 4.024 | 6.844 | 9.102 | 12.133 |
| 40 | 1.694 | 4.030 | 6.838 | 9.063 | 11.984 |
| 45 | 1.700 | 4.047 | 6.834 | 9.038 | 11.841 |
| 50 | 1.707 | 4.060 | 6.833 | 9.011 | 11.705 |
| 55 | 1.715 | 4.075 | 6.834 | 8.985 | 11.574 |
| 60 | 1.723 | 4.091 | 6.836 | 8.962 | 11.449 |

## Calibration with introducing the pH value of the sample solutions

Values of the pH buffer solutions set by the producer in the meter’s memory are given in the Table 2, this values may be changed in the range given in the table.  If the values of used buffer solutions are different than the given above the user should change the set values to the actually used.

The meter does not allow to enter the pH values different than those given in the table 2.

Table 2.

| Calibration point | Factory value | Range of changes |
| --- | --- | --- |
| 1 | 2,000 | 0,800 ( 2,100 |
| 2 | 4,000 | 3,900 ( 4,100 |
| 3 | 7,000 | 6,800 ( 7,100 |
| 4 | 9,000 | 8,900 ( 10,200 |
| 5 | 12,000 | 11,800( 14,000 |

In case of using buffer solutions with values different than the given by factory, it is necessary to remember about introducing this values to the meter’s memory. Way of doing it was described in the point 8.4.

The range of introducing the pH values of the buffer solutions in the individual points of calibration is quite wide what enables to use buffer solutions with values which differ from those set by the manufacturer. In each case the buffer will be automatically detected by the meter.

The manufacturers of sample solutions often give the values of the solutions in few temperatures. This data may be used and the meter may be calibrated in other temperature than 20 oC, by introducing to the meters memory value of the buffer which responds to this temperature, this will enable eliminating the necessity of warming up or cooling the solutions. It is possible to introduce to the meter’s memory values of the pH sample solutions with 2 or 3 decimal places, depending on the chosen resolution.

The value of sample solutions introduced to the meter’s memory by the user will be stored till the moment of changing it to other value.

## Entering the calibration parameters

It is necessary to:

a.	Choose or block the function of automatic correction of the change of pH value of buffer solution with the temperature. (described in chapter 8.1).

To do so one should press the  button than a  symbol (points of calibration) will be displayed, (Pic. ) next using  or  buttons choose in the lower row on LCD:

– automatic change of the stored pH value of the buffer solution by it’s temperature change according to table 1.

– setting the value of the buffer solution in the range given in table 2.

Pic. 8.

b.	After choosing the  or the   symbol one should start checking and eventually changing the stored values of calibration points. To do so one should press the  button. In the lower row a  symbol will be displayed, point one of calibration, and in the upper row value of the buffer. (Pic. ) If one uses buffer solution with value a bit different than this displayed one, using  or  buttons bring the displayed value to the value of used buffer solution.

Range of changes for each of calibration points in the  mode is ±0.010 from the values given in the table 1 for 20(C, and the range of changes in the   mode is given in the table 2.

Pic. 9.

To pass to the second point of calibration press the  button in the lower row a , symbol will be displayed (point 2) and in the upper row the stored in this point value of buffer solution. When checking or changing the value in the next steps please follow the instructions given above.

After the end of introducing the pH values of buffer solutions using the  button return to the choosing the calibration points mode  / , or by pressing the  button enter the pH measuring mode and calibration, chapter 8.5.

During next calibrations if the values of the previously used buffer solutions haven’t changed the actions described in this point may be omitted.

## pH electrode calibration

It is necessary to:

a.	choose the resolution with which the pH sample solution value will be introduced, point Błąd! Nie można odnaleźć źródła odsyłacza..

b.	choose the electrode number under which the results of calibration will be stored, point Błąd! Nie można odnaleźć źródła odsyłacza..

c.	prepare the electrode to work according to the producer’s instructions and mark it with number responding to number chosen in the meter (,  , ).

- connect with the meter the prepared combination pH electrode and temperature probe, use the right connectors pH/mV and temp;

- If the function of automatic change of the pH sample solution value is not used it is necessary to measure the temperature of sample  solutions and bring it to the temperature given by the manufacturer to receive the pH value identical like this entered to the meter’s memory. There is no limitation in order of using the samples.

Under each of the electrode numbers it is possible to calibrate different electrode in sample solutions with different values (after entering those values to the meters memory).

If the pH sample solutions with values according to NIST are used the values should not be changed. The meter takes under consideration only the values of sample solutions detected during calibration.

### Calibration with automatic temperature compensation

After preparing the meter to the calibration it is necessary in pH function to:

a	press and hold the  button till the appearance of the CAL (Pic. ), symbol on the display; the old parameters of calibration are deleted;

b	put the electrode and the temperature probe to the buffer solution, a P symbol with number of calibration point will be displayed, wait till stabilisation of the result. The result may be different than pH value of the buffer solution (Pic. ).

Pic. 10.

After stabilisation of the result press the .button.

In this moment the result will blink what will inform about the storing the calibration result, in this same time in the upper row the corrected value of measurement will be displayed – it will be equal the value of used buffer solution (Pic. 111).

If the function of automatic correction of the temperature influence on pH value of the buffer solution is used, this value will be considered in the displayed value of the buffer solution. If by accident someone will use buffer solution different than those set for the calibration the meter won’t be able to detect this value and  sign will be displayed. In this case it is necessary to check the value of the buffer solution or the electrode which may be broken or clogged.

Pic. 11.

- one may finish the calibration at this moment by pressing the  button or continue the calibration in other buffer solutions, washing the electrode and temperature probe before every immersing in the buffer and acting according to the point b.

After calibrating one electrode it is possible to calibrate two others, choosing the two remaining symbols according to point Błąd! Nie można odnaleźć źródła odsyłacza.

If after choosing the electrode number and entering the calibration mode one will exit it without making the calibration the earlier stored characteristic will be deleted and a standard characteristic will be adapted

### Calibration with manual temperature compensation

To start the calibration with manual temperature compensation one has to disconnect the temperature probe. Disconnecting the temperature probe switches the meter to manual compensation. On the display the entered value of the temperature is displayed, not the measured one. The ,  buttons will be unblocked and using them one may enter the value of the buffer solution temperature. This value is displayed in the lower row of numbers on the display.

Next thing is to connect the pH electrode with the meter and act according to points a ( c. of the previous section. It is important to remember that in case of manual temperature compensation the temperature value entered by the user is taken under consideration.

Caution:	simultaneous pressing of both  and  buttons sets the temperature to 20 0C.

# CHECKING THE ELECTRODE CONDITION

If after the calibration the electrode symbol is blinking (E1, E2, E3) it informs that the electrode has lost its efficiency and in short time its calibration won’t be possible. After entering the mode of electrode number changing under the number of electrode a  symbol will be displayed (

Pic. 12).

Pic. 12

It is necessary to prepare a new electrode.

Pic. 12.

# pH MEASUREMENT

Before starting the measurement one has to prepare the meter for work (chapter Błąd! Nie można odnaleźć źródła odsyłacza.) and pH electrode (chapter 7). Good condition of the electrode is the main condition of correct measurements. If the electrode was calibrated and is connected with the meter it is advisable to check whether the number of the electrode is this same as the number of the characteristic chosen from the memory. If not, the number should be changed according to section Błąd! Nie można odnaleźć źródła odsyłacza. and the resolution of measurement should be chosen according to the chapter Błąd! Nie można odnaleźć źródła odsyłacza..

## Measurement with automatic temperature compensation

During measurements with automatic temperature compensation, the meter co-operates with the temperature probe and measures the temperature of the solution simultaneously with the pH measurement and takes it into consideration during compensation.

In case of measurement with automatic temperature compensation one should:

- turn the meter on by pressing the  button;

- with the  button choose the pH measuring function;

- connect the temperature probe and the combination pH electrode to the right connectors on the meter pH/mV and temp, symbol  will be displayed;

- if the electrode wasn’t calibrated or was already used for some time it is advisable to calibrate it. (chapter 8);

- insert the electrode and the temperature probe to the measured solution. During measurements in vessels one should not touch the bottom and the walls with the electrode. Best way is to use an electrode stand;

- After the stabilisation read the result

Accurate laboratory measurements require using of electromagnetic stirrer.

NOTICE: exceeding of the measuring range is indicated by blinking of the digits on the display. Exceeding the automatic temperature compensation range is signalised by blinking digits and  symbol.

## Measurement with manual temperature compensation

Disconnecting the temperature probe from the meter switches the meter to the manual temperature compensation mode (symbol  is displayed).

Measurement with manual temperature compensation is similar to the measurement with ATC, the difference is that using ,  buttons one enters the temperature measured with other thermometer. This value is displayed under the pH value and is taken into consideration during the compensation.

The manual compensation may be used during stable conditions ex. during pH measurements in the laboratory, especially when a thermostat is used, or when the temperature probe was damaged.

During the measurement with manual temperature compensation one should:

- turn the meter on using the  button;

- using the  button choose the pH function;

- insert the pH electrode to the vessel with the measured solution, if the electrode wasn’t calibrated or already was used for a long time the calibration should be done (chapter 8). During measurements in a vessel one shouldn’t touch the bottom and the walls with the electrode. The best way is to use an electrode stand;

- using a laboratory thermometer measure the temperature;

- with    buttons enter the value of measured temperature;

- wait till the stabilisation of the value and read the result.

Caution: simultaneous pressing of  and   buttons sets the temperature value to 20 0C.

# NOTICES ABOUT THE TEMPERATURE COMPENSATION AND INTERPRETATION OF THE pH MEASUREMENTS RESULTS

The multifunction meter CX-505 has a manual and automatic temperature compensation, what enables eliminating errors resulting from changes of the electrode characteristics under the influence of temperature changes. To explain the role of the temperature compensation it is important to remind that pH meter is an mV meter which displays voltage counted to pH unit. In constant temperature for one pH unit a constant mV value is changed. In temperature of 20 0C it is 58,168 mV. The value of mV on one pH unit changes together with the temperature, what is taken into consideration in the formula for “k coefficient” of the pH electrode.

k=0.198422 T

Taking this change into consideration during measurement is called temperature compensation, which is connected with change of the electrode efficiency and not with the pH value change of the measured solution caused by temperature change.

Changes of the pH values of some solutions with the temperature are very small and others big ex. pure water.

When comparing solutions, which change their values together with the temperature change, it should be done in this same temperature.

Sometimes it happens that results of measurement in this same solution in stable temperature are different. When analysing such a situation it is important to take following factors into consideration:

- Differences may occur because of  not highest quality electrode;

- The result was treated as stabilised to fast (medium class electrode needs about 40 seconds to full stabilisation);

- The measured solution may not be homogeneous and lack of magnetic stirrer doesn’t allow for obtaining similar results;

- During measurements in sewage some chemical reactions, which change the result, may occur;

- The source of additional errors, which occur during the calibration, is the not accurate value of the buffer solutions.

- Very small differences may be caused by the meters accuracy. The CX-505 has accuracy of 0.002 pH, 1 digit, what practically means that in extreme situation results of 2 measurements may differ for 0.005 pH and this will be an acceptable error, because 1 measurement will be made with 0.002 pH error and second with -0.002 pH. 1 digit information explains difference caused by rounding up of the result on the last visible place on LCD (discretisation error).

If after making a two point calibration in pH buffers 7.00 pH and 4.00 pH (acidic conditions), the results are checked in 9.00 pH (alkali conditions),

in some cases the result may be 8.90 pH or 9.10 pH. This may occur when the electrode has unsymmetrical characteristic. Making at least a 3 point calibration with alkali, neutral and acidic buffers may prevent from such errors. Sometimes the measurement results are unstable, the quality of the electrode has the crucial influence. In most cases the reason for slow drifting of the result, it’s unstability or prolonging the time of stabilisation is the clogged junction, broken electrode or contaminated membrane.

Often this situation  happens when wrong kind of electrode was chosen for the kind of measured solution.

Leaving the electrode for few hours in the distilled water or placing it in water with detergent may eliminate this symptoms, especially if the measurements were done in solutions with deposits, fats or oils. The electrode which for a long time wasn’t used may have the junction clogged by KCl crystals, what may be removed by placing the electrode in distilled water and if this doesn’t help one may try to immerse the electrode in the thiourea in hydrochloric acid. Heavily contaminated electrode may be cleaned in chloroform and deposits of iron in 2N HCl. The electrode life may be prolonged by storing it in saturated KCl solution. Depending on the kind of measured solution or substance proper kind of electrode should be chosen. They differ one from another with shape, membranes look, kind of junction and body. Using electrode not proper for the kind of measured solution may cause it’s damaging and make the measurement impossible.

##### CONDUCTIVITY AND SALINITY MEASUREMENT

# BASIC INFORMATION ABOUT THE CONDUSTIVITY MEASUREMENT

The conductivity measurement is based on applying a voltage with proper frequency between two electrodes immersed in solution and measuring it. In the CX-505 the frequency depends on the measuring range and may vary from 100 Hz up to 10 kHz. Depending on the kind of the measured liquid, it’s concentration and the temperature, the voltage flow is easier or harder. The conductivity result indirectly informs about the salts concentration in the measured liquid. The greater is the salt concentration the bigger is the conductivity (KCl, NaCl) This dependence doesn’t concern all of the solutions. In some cases after exceeding some salinity value the conductivity starts decreasing. The temperature has also a great influence on the results of a conductivity measurement which increases together with the temperature growth. The measured conductivity may be displayed as salinity in g/l of NaCl or KCl assuming that the measured liquid includes homogeneous salt.

The electrode’s surface and the distance between them have decisive influence on the so called constant K value of the cell. This value has a great influence on the accuracy of the measurement. Depending on the measured conductivity value cells with constant K = 0,1 cm-.1 up to  10 cm-.1 are used. During the measurement the meter multiplies the measured value by the constant K introduced to the meters memory and displays the result in units of conductivity (μS/cm or mS/cm). By the result a shortened symbol of the unit is displayed (μS or mS). The conductivity is changing together with temperature and salts concentration. To make the comparing of the results possible the measured value is counted by the meter to the value which responds to measurement in reference temperature. Usually it is 25°C. Measurement in reference temperature is most accurate. In other temperatures there is so called temperature compensation used, this means that there are some calculations made by the meter which take the actually measured temperature and ( coefficient under the consideration (it should be introduced before the measurement point 13.3). This coefficient describes how much (in %) will the result change when the temperature will change for 10C. The  coefficient in this meter may be changed in range 0 ÷ 10.00% / oC. The value of this coefficient and reference temperature  is introduced to the memory by the user. For NaCl in temperatures close to 25 0C it is 2%/ 0C. Ex. in case of measurements in 30 0C the result changes 5x2%=10%. The calculation is done automatically and takes the introduced by the user value of the  coefficient into consideration. The value of this coefficient may be found in the literature or may be approximated by the user (point 18).  The conductivity measurements, despite the electronic precision, should always be treated as burdened with some error, which depends on the error of conductivity cell (it’s linearity), temperature and first of all depending on the ( coefficient during measurements in temperatures different than the set reference temperature.

# ENTERING THE CONDUCTIVITY METER PARAMETERS

# Before the calibration and measurements it is necessary to make all activities described in the chapter 6. Additionally according to the chapter below it is necessary to choose the unit in which we will make the calibration and measurement.

## Choosing the unit

The result of the measurement may be displayed in units of conductivity and salinity. The salinity may be counted to NaCl, KCl or TDS content.

The result may be displayed in % of weight concentration or in g/l. To choose the unit one has to:

-	in the conductivity measuring mode press the  button till displaying in the lower row on LCD a   (unit) symbol;

-	with buttons  ,  choose in the upper row on LCD sign:

-	measurement in units of conductivity (Pic. 13);

Pic. 13.

-	measurement result calculated to NaCl in g/l (Pic. 14);

Pic. 14.

-	measurement result calculated to  KCl in g/l (Pic. 15);

Pic. 15.

-	measurement result calculated to  TDS in g/l (Pic. 166).

Pic. 16.

In case of salinity measurement (,  or ), short pressing of the  button chooses displaying the result in % of weight concentration or g/l. By the  symbol  a % or g/l symbol will be displayed.

-	by pressing the . enter the measuring mode.

The result of measurement in % of weight concentration may be counted to value in ppm according to dependence:

1% of weight concentration (C) = 10 000 ppm = 10 ppt

Resolution of measurement in % of weight concentration is 0.001% or 10 ppm.

## Entering the WTDS coefficient

In case of measurement of salinity with conversion to TDS it is necessary to introduce the WTDS coefficient.

One should follow the steps:

-	in the conductivity measuring mode few times press the  button till moment of displaying in the upper row on LCD a   symbol – TDS coefficient (Pic. 17);

-	with  ,  buttons enter the correct value of the TDS coefficient;

-	enter the measuring mode by pressing the  button.

Pic. 17.

The way of determining the WTDS coefficient is described in section 17.3.

## Entering the  coefficient value

The  coefficient entering range in CX-505 is 0.00 ÷ 10.00 % / oC with accuracy of change 0.01 % / oC. For the measurements it is possible to adopt the most often used temperature coefficient  = 2 % / oC or in case of higher accuracy requirements determine the kind of measured solution and choose for it the correct value of this coefficient.

To introduce the  coefficient one should:

-	in the conductivity measuring mode press the  button till displaying the screen with value of the  coefficient (Pic. 188). A  (temp. coefficient  ) symbol will be displayed;

-	with  ,  buttons enter the value of the coefficient .

-	return to the measuring mode by pressing the  button.

Pic. 188.

The measurement result will be counted with use of the introduced temperature coefficient .

# The simplified way of determining the  coefficient is described in the chapter 18.

## Entering the reference temperature value

The range of introducing the reference temperature values is 10.0 ÷ 40.0 oC with possibility of setting with accuracy of 0.1 oC. The most often used value is 25 oC.

To introduce the reference temperature one should:

-	in the conductivity measuring mode press the  button till displaying a screen with value of the reference temperature (Pic. 19). A  sign will be displayed (reference temperature);

-	with the  ,  buttons introduce the temperature value;

-	return to the measuring mode by pressing the  button.

Pic. 19.

The measurement result will be calculated to the introduced reference temperature.

Caution:	simultaneous pressing of  and  buttons sets the value to 25 0C.

# CHOICE AND MAINTENANCE OF THE CONDUCTIVITY CELL

## The cell choice

The conductivity measuring range in CX-505 meter is 0 – 1999 mS/cm. The meter co-operates with conductivity cells with constant K = 0.010 ÷ 19.99 cm-1 and BNC-50 connector. Depending on the required measuring range it is necessary to choose the right cell with constant K which enables receiving correct results. Beyond the range the cell looses it’s linearity and the results have greater error. Accurate measurements in the whole range are possible with use of 3 different conductivity cells.

Depending on the expected measuring range one may choose the correct cell using the chart below.

Pic. 20.	Dependence between the measuring range and constant K of the used conductivity cells.

# The cell with constant K ≈ 0,1 cm –1  should be used for measurements of ultra pure and redistilled water. Those waters after contact with air are very quickly changing their conductivity, so their pouring into vessels for measurement with dip cell (immersing) may cause errors. For accurate measurements it is necessary to use flow through probe with build in temperature probe, what enables measurement of water which flows directly from the container.

## The conductivity cell maintenance

To receive stable results it is recommended to store the cell for few hours before the measurement in water, especially this is required in case of distilled water measurements.

The conductivity cell maintenance is mainly based on accurate washing the inside of the measuring cell with distilled water. It is forbidden to clean the platinum electrodes mechanically, because the platinum layer may be rub of, what will result in decreasing the accuracy, may cause lowering the stability and change the constant K.

Measurements of liquids with oils and heavy sediments content may cause platinum contamination, make the measurement impossible and irreparably damage the electrodes. In case of fat content in the measured liquids it is possible to clean the electrodes by immersing the cell in acetone, chloroform, fourhydrofurane or detergent.

Some norms propose universal liquid for cleaning the cells. This is mixture of equal parts of isopropyl alcohol, ethyl ether and hydrocholic acid, diluted with water in 1:1 ratio.

The conductivity probe with broken measuring cell can’t be used for further measurements because the constant K is greatly changed, the result is unstable and the influence of placing the cell in measuring vessel greatly changes the result.

# CALIBRATION

The calibration is done to prepare the meter for co-operation with the used conductivity probe, this is the essential condition for obtaining accurate results. The calibration may be done without the sample solution (by entering to the meter’s memory known value of the constant K of cell) or with use of standard solution with known conductivity. Constant K of the conductivity cell may change as reaction to contamination of the electrodes, that’s why the cell should be kept clean and from time to time to check the meters indications on the sample solution and eventually repeat the calibration. When using three cells it is possible to store their characteristics under the ,  or  symbols.

## Calibration without the sample solution

The meter has the possibility of calibration without the use of sample solution. To make this calibration it is necessary to know the constant K of the conductivity probe. This value may be given by the cell producer or may be determined using the meter, after making the calibration in the sample solution.

To make the calibration one should:

-	in the measuring mode press the  button till displaying the screen with the value of constant K (Pic. 21).

with  ,  buttons choose the number of cell (symbols ,  or );

-	press the  button on the display a CAL symbol will be displayed;

-	with  ,  buttons enter the value of constant K;

-	press the  button and return to the mode of choosing the number of cell or enter the measuring mode by pressing the  button.

Pic. 21.

## Calibration with use of the sample solution

The meter enables one point calibration in the freely chosen sample solution. To decrease the error it is recommended to use solutions with value close to the forecasted value of measurement. It is required to use high quality sample solutions. The calibration is done in the actually chosen unit (subchapter 13.1).

Obtaining accurate results of the calibration depends on following the below given rules:

- The sample temperature should be equal to the reference temperature (in most cases 25 0C), using thermostat is suggested.

- New not used sample solution should be used .

- The cell and the electrodes of the calibrated probe should be clean and without air bubbles.

- The cell should be placed in electrode holder.

### Entering the sample solution value

To enter the value of sample solution one should:

-	choose the unit according to point 13.1;

-	in the conductivity measuring mode press the  button till appearing a   symbol in the lower row on LCD (Pic. 22);

-	with  ,  buttons enter, in the upper row on LCD, the value of sample solution;

-	enter the measuring mode by pressing the  button.

Pic. 22.

### Calibration with automatic temperature compensation

What to do:

-	enter the value of sample solution  (point 15.2.1);

-	connect the conductivity cell and temperature probe;

-	immerse both probes in the sample solution, hold the conductivity cell at least 1 cm away from the bottom and walls of the vessel. The measuring cell should be completely filled with the measured solution, there shouldn’t be any air bubbles and the electrodes should be evenly moistened*;

-	measure the temperature of solution and bring it to the value reference temperature;

-	press and hold the  button till appearing a CAL symbol on the display (Pic. 23).

-	wait till the stabilisation of the value and press the  button. Pulsating of the results informs about storing in the memory. If  symbol will be displayed it is necessary to check the introduced value of the sample solution.

-	exit the calibration mode by pressing the   button.

Pic. 23.

The meter is calibrated and ready for work.

### Calibration with manual temperature compensation

To calibrate the meter it is necessary to:

-	turn the meter on with the  button;

-	choose the measurement of conductivity with  button;

-	disconnect the temperature probe and simultaneously press the ,  buttons. In the lower row of LCD the introduced reference temperature value will appear;

-	introduce the value of the sample solution (chapter 15.2.1);

-	immerse the conductivity cell in the sample solution and hold it at least 1cm away from the bottom and walls of the vessel. The measuring cell should be completely filled with the sample solution and shouldn’t include any air bubbles, the electrode’s surface should be evenly moistened*;

-	with lab thermometer measure the temperature of the sample solution and bring it to the introduced reference temperature value;

-	press and hold the  button till CAL symbol (Pic. 24) will appear on LCD.

-	wait till stabilisation of the result and press the  button. Pulsating of the result informs about storing it. If  sign will be displayed it is necessary to check the entered value of sample solution.

-	exit the calibration mode by pressing the  button.

The meter is calibrated and ready to work.

Pic. 24.

* -	air bubbles may be removed by moving the immersed cell. To make the moistening of the electrodes easier it is recommended to immerse the cell in distilled water with washing-up liquid, and than wash it in distilled water.

# THE CONDUCTIVITY MEASUREMENT

## Measurement without the temperature compensation

An accurate conductivity measurement should be done without the temperature compensation. The measured solution should be brought to the value of earlier introduced reference temperature to control it, it is possible to use the meter’s temperature probe. In case of work without temperature probe it is necessary to introduce the temperature value with  ,  buttons.

One should:

-	connect the conductivity and temperature probes to the right connectors cond and temp;

-	turn the meter on with  button.

-	choose the conductivity measurement and the unit (point 13.1);

-	if the conductivity cell wasn’t calibrated earlier make the calibration according to chapter 15;

-	Place both probes in the measured solution, the conductivity cell can’t touch the walls and bottom. The measuring cell should be completely filled with the sample solution and shouldn’t include any air bubbles, the electrode’s surface should be evenly moistened*;

- bring the temperature of the measured solution to the set reference temperature value.

- read the result after it’s stabilisation (Pic. 25)

Pic. 25

* -	air bubbles may be removed by moving the immersed cell. To make the moistening of the electrodes easier it is recommended to immerse the cell in distilled water with washing-up liquid, and than wash it in distilled water.

## Measurement with automatic temperature compensation

# In case of measurement with automatic temperature compensation one should:

-	connect the conductivity and temperature probes to the right connectors cond and temp;

-	turn the meter on with  button.

-	choose the conductivity measurement with   button and the unit (pt. 13.1);

-	if the conductivity cell wasn’t calibrated earlier make the calibration according to chapter 15;

-	control or change the value of the temperature coefficient  and reference temperature;

-	Place both probes in the measured solution, the conductivity cell can’t touch the walls and bottom. The measuring cell should be completely filled with the sample solution and shouldn’t include any air bubbles, the electrode’s surface should be evenly moistened*;

-	wait till the stabilisation and read the result (Pic. 26)..

# Pic. 26.

CAUTION: in case of exceeding the value of temperature taken during compensation into consideration the result and  symbol start blinking although the conductivity measuring range wasn’t exceeded.

If near the value a  symbol is displayed instead of  it informs that the temperature probe is broken or inaccurately connected.

* -	air bubbles may be removed by moving the immersed cell. To make the moistening of the electrodes easier it is recommended to immerse the cell in distilled water with washing-up liquid, and than wash it in distilled water.

## Measurement with manual temperature compensation

The measurement with manual temperature compensation may be done during stable work conditions, ex. during measurements in laboratory, especially with use of thermostat, or in case of the temperature probe damage. Disconnecting the temperature probe switches the meter to manual temperature compensation.

In case of measurement with manual temperature compensation one should:

-	connect the conductivity probe to the right connector  cond;

-	disconnect the temperature probe from the meter;

-	turn the meter on with  button.

-	choose the conductivity measurement and the unit (point 13.1);

-	if the conductivity cell wasn’t calibrated earlier make the calibration according to chapter 15;

-	control or change the value of the temperature coefficient   and the reference temperature;

-	Place the probe in the measured solution, the conductivity cell can’t touch the walls and bottom. The measuring cell should be completely filled with the sample solution and shouldn’t include any air bubbles, the electrode’s surface should be evenly moistened*;

-	measure the temperature of the solution and introduce it’s value with the  ,  buttons.

-	after the stabilisation read the result (Pic. 27).

Caution:	simultaneous pressing of  and  buttons sets the compensation temperature to the earlier set reference temperature.

Pic. 27.

* -	air bubbles may be removed by moving the immersed cell. To make the moistening of the electrodes easier it is recommended to immerse the cell in distilled water with washing-up liquid, and than wash it in distilled water.

# SALINITY AND TDS (TOTAL DISSOLVED SOLIDS) MEASUREMENT

Salts and minerals dissolved in natural water influence the conductivity, which in principle is proportional to the quantity of dissolved substances. This dependence enables, after some calculations, to determine the salinity of the measured solution in concentration units (g/l or %), or determine the TDS (Total Dissolved Solids). The received values are always approximated and the final accuracy depends on the way of making the calculations, concentration of the measured solution and it’s temperature. In most salinity meters a simplification is used, that dependence between the conductivity and salinity in the solution is linear in the whole measuring range. Usually a 0.5 coefficient is used, the conductivity result in mS/cm is multiplied by this coefficient and the result of salinity in g/l is received, ex. if the conductivity value is 2 mS/cm the salinity is 1g/l. In reality the dependence between the conductivity and salinity isn’t linear the conversion coefficient is changing together with the concentration and temperature. The table 3 shows the dependence between conductivity and real salinity of NaCl solution in temperature 25 0C and values of salinity counted for constant coefficient 0.5. This comparison shows that for greater concentrations using the constant coefficient introduces large error.

Table 3.

| Conductivity (mS/cm) | Real salinity (g/l) | Salinity (g/l) Counted for coefficient = 0.5 | Error (%) by using the coefficient = 0.5 |
| --- | --- | --- | --- |
| 1.00 | 0.495 | 0.500 | 0.01 |
| 2.00 | 1.006 | 1.000 | 0.60 |
| 4.00 | 1.976 | 2.000 | 1.21 |
| 10.00 | 5.400 | 5.000 | 7.40 |
| 30.00 | 18.174 | 15.000 | 17.46 |

In CX-505 microcontroller takes into consideration the real dependence between the conductivity and salinity what greatly lowers the error. There is possibility of counting the salinity in NaCl or KCl, because the dependence for this two salts is a bit different.

The results will be more accurate for homogeneous solutions (NaCl, KCl). Determining the concentration of salts mixture with unknown composition in most cases is counted to NaCl . To check the usefulness of water for home or industrial use usually determining of TDS is used.

To use the conductivity measurement for determining the TDS it is necessary to specify the WTDS coefficient, which introduced to the meter’s memory enables automatic conversion. To determine the WTDS coefficient it is necessary to specify the weight of dissolved substances. The laboratory method of determining the dissolved solids content consists of taking a given volume of the water, evaporating the filtered sample, drying it to constant weight in temperatures 103 ( 1050C, weighting and counting in volume ratio (mg/dm3). The received weight is lower than total dry mass in water because besides the dissolved solids there can be also substances which were not dissolved which are removed by filtering before evaporating. Using the conductivity measurement it is possible to determine the Total Dissolved Solids. It is assumed that the salt’s composition in the taken samples is not changing significantly.

## Salinity measurement with conversion to NaCl or KCl content

The measurement of salinity with conversion to NaCl or KCl content should be done as follows:

-	choose the salinity measurement with conversion to NaCl or KCl content according to chapter 13.1;

-	choose the unit (g/l or %);

-	then act as during the conductivity measurement (chapter 16);

-	after stabilisation read the result.

## Salinity measurement with conversion to TDS content

The measurement of salinity with conversion to TDS should be done as follows:

- according to the chapter 17.3 determine the WTDS coefficient;

- according to chapter 13.2 introduce the WTDS coefficient;

- choose the measurement of salinity with conversion to TDS and the displayed unit (g/l or %);

- than act as during conductivity measurement;

- after stabilisation read the result in chosen units (g/l or %).

## Determining the WTDS coefficient

For salinity measurement with conversion to TDS content it is necessary to determine the WTDS coefficient and enter it to the meter’s memory. To do so it is necessary to make the conductivity measurement of the tested water with exactly given volume or weight. In traditional way determine the Total Dissolved Solids in this water and than calculate the coefficient according to the below given formulas.

1. If the result is to be displayed in  g/l:

WTDS  =

where:

WTDS  - TDS coefficient

TDS – Total Dissolved Solids in g/l;

( - conductivity of the sample in mS/cm;

Caution: the TDS value should be counted to volume of sample equal to 1l.

2.	If the result is to be displayed in % of weight concentration:

WTDS  =

where:

WTDS  - TDS coefficient

TDS – Total Dissolved Solids in g/kg;

( -conductivity of the sample in  mS/cm;

Caution: the TDS value should be counted to weight of the sample equal to 1kg.

# SIMPLIFIED WAY OF DETERMINING THE  COEFFICIENT

The knowledge of  coefficient has a crucial significance during measurements in temperatures different than the reference temperature.

This coefficient is changing together with the temperature and concentration. Below we are giving values of  coefficient in 25 oC for few compounds with determined weight concentration.

Table 4.

| substance | Weight Concentr. | coefficient |
| --- | --- | --- |

| HCl | 10 % | 1.56 |
| --- | --- | --- |

| KCl | 10 % | 1.88 |
| --- | --- | --- |

| H2SO4 | 50 % | 1.93 |
| --- | --- | --- |

| NaCl | 10% | 2.14 |
| --- | --- | --- |

| HF | 1.5 % | 7.20 |
| --- | --- | --- |

| HNO3 | 31 % | 1.39 |
| --- | --- | --- |

In the Błąd! Nie można odnaleźć źródła odsyłacza.  there are rough values of  coefficient for KCl and NaCl depending on the temperature and concentration of the measured liquid.

Table 5.

| temp. `0C | coefficient |
| --- | --- |
|  | KCl solution | Saturated NaCl |
|  | 0,01M | 0,1M | 1,0M |  |
| 5 | 2,68 | 2,68 | 2,39 | 2,77 |
| 10 | 2,45 | 2,36 | 2,20 | 2,53 |
| 15 | 2,27 | 2,19 | 2,04 | 2,38 |
| 20 | 2,11 | 2,06 | 1,89 | 2,21 |
| 25 | 1,91 | 1,86 | 1,75 | 2,03 |
| 30 | 1,80 | 1,77 | - | 1,91 |

One should use the  coefficient determined for the reference temperature to which the meter makes the calculations. It may be assumed that the coefficient is constant in the range ±5 oC from the reference temperature.

For larger differences between the measurement and reference temperature the value of the  coefficient may be determined according to the below given description.

To determine the coefficient one should:

1.	The measured solution bring to the reference temperature TR and measure it’s conductivity (GTR).

2.	Change the solution temperature TX to value in which the measurement will be done.

3.	Turn the meter to manual temperature compensation by disconnecting the temperature probe from the meter.

4.	Enter with the keyboard the value of reference temperature TR .

- Again measure the conductivity of the solution.  This value will be different than in TR temperature (GTx).

- Determine the  coefficient using the  formula:

=

where:	TR	-	reference temperature value in 0C

Tx	-	value of the changed temperature in 0C

GTR	-	conductivity measured in ref. Temperature TR.

GTx	-	conductivity measured in temperature Tx

In this way the   coefficient was determined for the reference temperature TR and measurement temperature Tx.

During measurements in temperatures different than reference temperature it is necessary to introduce the calculated  coefficient value for the measurement temperature.

Caution: In case when the TR reference temperature is equal to 25°C the above mentioned formula will be:

=

where:	Tx	-	value of the changed temperature in 0C

G25	-	conductivity measured in 25 0C.

GTx	-	conductivity measured in temperature Tx

##### DISSLVED OXYGEN MEASUREMENT

# BASIC INFORMATION ABOUT THE DISSOLVED OXYGEN MEASUREMENT

The measurement of dissolved oxygen in water solutions is performed using an oxygen sensor. The basic element of the sensor is a Teflon semi-permeable membrane, which enables the penetration of oxygen contained in the measured solution, into the electrolyte – inside of the sensor. The sensor creates a cell, which voltage depends on the oxygen content in the electrolyte.

The meter enables measurement in % of oxygen saturation and mg/l. The calculation of the mg/l value is based on the saturation measurement in % and the temperature measurement. During mg/l measurements, the values of salinity and atmospheric pressure must be additionally introduced. The saturation measurement in % does not depend on those factors.

The quality of the oxygen sensor has a decisive effect on the measurement accuracy. The troubles arising during the measurements are caused mainly (98%) by the sensor, yet not by the device. In many cases the troubles result from lack of basic maintenance of  the sensor from the user’s side. It must be remembered, that during the measurement the sensor takes oxygen from the environment of the membrane.

The sensor’s manufacturers recommend in their instructions the minimal flow-rate of the tested water, assuring a stable result. If that requirement is not complied the result will regularly decrease. During measurements in stagnant solutions the flow can be partly simulated by keeping the sensor in motion with a suitable speed. In laboratory conditions, i.e. performing the measurements in a vessel, the flow can be forced with a magnetic stirrer. Although it must be kept in mind that when measuring low O2 saturation intensive stirring can cause an increase of oxygen content in the tested solution. Taking water samples and transporting them to the laboratory can alter the O2 concentration in the samples. The best results can be achieved only in conditions, which are recommended by the manufacturer of the sensor in the operation manual.

Long-lasting storing of the sensor without performing of measurements (above 2 months) requires removing of the electrolyte. After this period the container must be filled with a fresh electrolyte and the sensor stored in distilled water for about 24 hours. A correct measurement is determined by the good condition of the membrane. The membrane must be free of any cracks (appearing of electrolyte-drops or white spots when dry). Before the measurement the electrode should be activated by storing in distilled water for about 15 minutes.

Strongly polluted wastewater causes after some time clogging of the membrane. This is revealed by the impossibility of calibrating of the device at 100% oxygen content (the calibration range becomes too narrow). In both cases the container with membrane should be replaced according to the manufacturer’s instructions. When replacing the membrane and replenishing the electrolyte it is important to pay attention weather there are no air bubbles in the container beneath of the membrane, because otherwise the measurements will be false. In that case the cap with membrane should be unscrewed and gently knocked on the table, after that refilling the electrolyte and screwing back is necessary.

During measurement in order to obtain a stable result there must be some awaiting time before making the reading. According to the sensor’s manufacturers that is about 1-1,5 min, depending on the thickness of the membrane. The accuracy of the measurement is connected with the temperatures of calibration and measurement. The greater the difference of these temperatures, the greater the measurement error.

Depending on the  measured values the sensors require a one or two-point calibration. For measurements of concentration 30  80% one point calibration will be enough in 100% point. For measurements of very small concentration (few %) or higher accuracy requirements the sensor additionally should be calibrated in the zero O2 (0%) solution. Clean water contains about 60 ÷ 80% oxygen. Waste water and chemical solutions are in general less saturated with oxygen but liquids with forced aeration are much more saturated. When performing accurate measurements the sensor’s manufacturers recommend carrying out calibration just before the measurement since after some time the sensor’s parameters are changing. Even the best oxygen sensors have so called drift about ±1%/24 h. Wide measuring range in multifunction meter CX-505 enables measurements of waters which are permeated with oxygen. This situation occurs when in the water plants are blooming and growing, in this moment during the photosynthesis process large quantities of oxygen are produced.

# DISSOLVED OXYGEN SENSOR

The meters co-operates with the galvanic sensor made by ELMETRON with accuracy of ±1%, if the measurement is done in this same temperature as the calibration. The accuracy of measurement decreases together with growing of the difference between the temperature of calibration and temperature of measurement. It is <3% when the difference is ±5°C and 5% when the difference is ±10°C.  In case of using another sensor the device requires its adaptation by the manufacturer.

It must be remembered that the membrane of the sensor should be replaced if it is mechanically damaged or if the device can’t be calibrated. The necessary procedures are given in the instruction of the probe’s manufacturer.

Before starting the work please read carefully the information given in chapter 19.

# ENTERING THE OXYGEN METER PARAMETERS

## Change of the unit

The measurement result may be displayed in % of the oxygen saturation or in mg/l. To choose the unit one should:

-	in the oxygen measuring mode press the  button till displaying in the upper row of LCD  (unit) sign;

-	with  ,  buttons choose the symbol:

% - measurement in % of oxygen saturation

mg/l – measurement in mg/l (Pic. 28).

Pic. 28.

-	return to the measuring mode by pressing the  button.

## Compensation of the salinity influence

Salinity of the solution decreases the oxygen solubility in water and requires taking into consideration during measurements in mg/l. 1 g/l change of salinity changes the oxygen saturation for about 5%. The meter enables entering the salinity value in g/l and counts the change of oxygen saturation in mg/l.

### Automatic introduction of the salinity value

The meter enables automatic introduction of measured solution salinity. To use this feature one should:

-	by pressing the  button enter the conductivity measuring function;

-	measure the salinity of solution in NaCl (accurate description in chapter 17.1);

-	after stabilisation of the salinity result enter the oxygen measuring mode by pressing the  button;

-	press the  button till moment of displaying in the lower LCD row a  (salinity) symbol;

-	shortly press the  button in the upper LCD row automatically the value of measured salinity in g/l will be provided. If after pressing the button a  (error) sign will be displayed and next the old value of salinity, it informs that during the conductivity measurement the unit wasn’t changed for g/l of NaCl;

-	return to the measuring mode by pressing the  button.

## Manual entering of the salinity value

The salinity value may be determined on the basis of known conductivity of the measured solution. The Table 5 gives the real dependence between the salinity and conductivity counted in NaCl.

To enter the salinity value one should:

-	with any conductivity meter measure the conductivity of the solution and from the Table 5 read the salinity value;

-	in the oxygen measuring mode press the  button till displaying in the lower LCD row    (salinity) sign;

-	with  ,  buttons enter in the upper LCD row the salinity value red from the Table 5;

-	enter the oxygen measuring mode by pressing the  button.

Measurement in % of saturation doesn’t require entering the salinity value

| mS/cm | g/l | mS/cm | g/l | mS/cm | g/l |
| --- | --- | --- | --- | --- | --- |

| 1 | 0.49 | 28 | 16.87 | 55 | 34.34 |
| --- | --- | --- | --- | --- | --- |

| 2 | 1.00 | 29 | 17.52 | 56 | 34.99 |
| --- | --- | --- | --- | --- | --- |
| 3 | 1.52 | 30 | 18.17 | 57 | 35.64 |
| 4 | 2.08 | 31 | 18.82 | 58 | 36.28 |
| 5 | 2.63 | 32 | 19.46 | 59 | 36.93 |
| 6 | 3.19 | 33 | 20.11 | 60 | 37.58 |

| 7 | 3.74 | 34 | 20.76 | 61 | 38.23 |
| --- | --- | --- | --- | --- | --- |

| 8 | 4.29 | 35 | 21.41 | 62 | 38.87 |
| --- | --- | --- | --- | --- | --- |

| 9 | 4.85 | 36 | 22.05 | 63 | 39.52 |
| --- | --- | --- | --- | --- | --- |

| 10 | 5.40 | 37 | 22.70 | 64 | 40.17 |
| --- | --- | --- | --- | --- | --- |

| 11 | 6.00 | 38 | 23.35 | 65 | 40.81 |
| --- | --- | --- | --- | --- | --- |

| 12 | 6.61 | 39 | 23.99 | 66 | 41.46 |
| --- | --- | --- | --- | --- | --- |

| 13 | 7.21 | 40 | 24.64 | 67 | 42.11 |
| --- | --- | --- | --- | --- | --- |

| 14 | 7.83 | 41 | 25.29 | 68 | 42.75 |
| --- | --- | --- | --- | --- | --- |

| 15 | 8.45 | 42 | 25.93 | 69 | 43.40 |
| --- | --- | --- | --- | --- | --- |

| 16 | 9.07 | 43 | 26.58 | 70 | 44.05 |
| --- | --- | --- | --- | --- | --- |

| 17 | 9.70 | 44 | 27.23 | 71 | 44.70 |
| --- | --- | --- | --- | --- | --- |

| 18 | 10.35 | 45 | 27.87 | 72 | 45.34 |
| --- | --- | --- | --- | --- | --- |

| 19 | 11.01 | 46 | 28.52 | 73 | 45.99 |
| --- | --- | --- | --- | --- | --- |

| 20 | 11.66 | 47 | 29.17 | 74 | 46.64 |
| --- | --- | --- | --- | --- | --- |

| 21 | 12.31 | 48 | 29.82 | 75 | 47.28 |
| --- | --- | --- | --- | --- | --- |

| 22 | 12.96 | 49 | 30.46 | 76 | 47.93 |
| --- | --- | --- | --- | --- | --- |

| 23 | 13.61 | 50 | 31.11 | 77 | 48.58 |
| --- | --- | --- | --- | --- | --- |

| 24 | 14.26 | 51 | 31.76 | 78 | 49.22 |
| --- | --- | --- | --- | --- | --- |

| 25 | 14.91 | 52 | 32.40 | 79 | 49.87 |
| --- | --- | --- | --- | --- | --- |

| 26 | 15.56 | 53 | 33.05 | 80 | 50.63 |
| --- | --- | --- | --- | --- | --- |

| 27 | 16.22 | 54 | 33.70 |  |  |
| --- | --- | --- | --- | --- | --- |

Table 5.	Determining the salinity in g/l NaCl on the basis of conductivity in mS/cm (in temperature 25oC).

## Automatic compensation of the air pressure influence

The value of oxygen saturated in water determined in mg/l depends directly on the atmospheric pressure value, this means that 10% pressure change causes oxygen saturation change for 10%. The meter ensures automatic compensation thanks build in atmospheric pressure sensor. This influence is automatically counted during measurements in mg/l.

There is possibility of reading the atmospheric pressure value:

-	in the oxygen measuring mode press the  button till moment of appearing in the lower LCD row  (pressure) sign (Pic. 29);

-	in the upper row a value of atmospheric pressure will be displayed in hPa;

Pic. 29.

-	enter the oxygen measuring mode by pressing the  button.

During the oxygen measurement in % of saturation the atmospheric pressure has no influence on the result.

# CALIBRATION OF THE DISSOLVED OXYGEN SENSOR

In order to eliminate the measurement error arising from the individual characteristic of the sensor a calibration of the device should be carried out. This procedure should be performed always before operation with a new sensor, after replacing the membrane or for special requirements concerning the measurement accuracy. The sensors have a so-called “signal drift” associated with the interval between the calibration and measurement. A longer interval decreases the measurement accuracy. The calibration is also recommended if the temperature of the tested solution differs greatly from the temperature in which the probe was calibrated, because then an additional error arises. In this case calibration solutions should be used, which temperature is almost the same as the predicted temperature of the tested solutions.

If it is impossible to calibrate the device, the membrane of the sensor must be replaced according to the manufacturer’s instruction. This situation usually takes place if the membrane is strongly polluted or ruptured (sometimes almost invisible). After replacing the membrane the sensor should be conditioned in distilled water for 24 hours.

The used oxygen sensors require one or two point calibration in standard solutions. During two point calibration a solution with 0% oxygen saturation is used (solution of Na2SO3) and second  with 100% oxygen saturation. It can be prepared by aeration of water for more than ten minutes. When using this solution it should flow or be stirred. The method of preparing both solutions is described in detail in the users manual for the oxygen sensor. A simplified calibration for 100% O2 saturation can be carried out in the air, without immersing the sensor in water. Before that the membrane must be wetted with water for several minutes.

It is assumed that the O2 content in the air correspondents to 100%- saturation, what enables a simplified calibration to be carried out.

The one point calibration is done only in the solution with 100% oxygen saturation.

Starting the calibration under the chosen sensor number deletes the characteristic stored in memory under this number.

If after choosing the sensors number and entering the calibration mode one will exit this mode the stored characteristic will be deleted and standard characteristic will be applied.

To calibrate the meter one should:

-	place the sensor in the vessel with 0% saturation solution;

-	choose the saturation in % measurement according to chapter 21.1;

-	press and hold the  button till appearing in the left lower corner of display a CAL symbol, the meter will enter the calibration mode and automatically will change the unit to %;

-	after the stabilisation of result press the  button. The result will pulsate what informs about storing the value of calibration at this same time in the upper row of LCD a corrected value of the measurement (0%) will be displayed.

-	Take the sensor out, wash it accurately in distilled water and put it into 100% oxygen saturated solution ensuring the liquid flow or leave it on the air.

-	after the stabilisation of the result press the  button.

The result will pulsate, what informs about storing the calibration data at this same time in the upper LCD row corrected value of the measurement will be displayed (100%).

-	enter the measuring mode by pressing the  button.

If after pressing the  button the meter can’t detect the value of the sample solution (0% or 100%), in the place of result for a moment an    sign will be displayed (Pic. 30). In this case it is necessary to check the condition of the membrane and the used solutions.

Pic. 30.

# OXYGEN CONCENTRATION MEASUREMENT

Before starting the oxygen concentration measurement the meter should be prepared for work (chapter Błąd! Nie można odnaleźć źródła odsyłacza.) and the oxygen sensor calibrated (chapter 22). As mentioned the measurement in % saturation does not require additional measurements associated with the temperature, salinity and atmospheric pressure. Although the measurement in mg/l, used more frequently, depends on those factors. This influence is corrected automatically by the device, taking into the consideration the temperature value measured by the sensor or in case of manual compensation value entered by the user. The oxygen sensor is equipped with an additional system compensating the temperature influence on the membrane. Because of the limited accuracy of this compensation the highest accuracy can be achieved by calibrating the sensor at the same temperature at which the measurement will be carried out. The measurement error increases with the increase of the difference between the calibrating and measurement temperatures and results from characteristic features of the sensor but not from the device. For the applied ELMETRON sensor this error is about < 3% at a (50C temperature difference and increases to 5% at a (100C temperature difference.  If a higher accuracy is required the interval from the last calibration must be additionally taken into consideration (signal drift). If the salinity of the tested solution is very small, the measurement can be started without entering it’s value (value 0.00 g/l should be entered). However accurate measurements should be preceded by the determination of the salt content in the tested solution.  The salinity of the solution can be determined simply by measuring the electric conductivity using the conductivity meter. The easiest way to determine the salinity is conductivity measurement with conversion to NaCl. The salinity value is introduced according to chapter 21.2.

## Measurement with automatic temperature compensation

To make the measurement with automatic temperature compensation:

-	connect the temperature probe with the meter;

-	according to chapter Błąd! Nie można odnaleźć źródła odsyłacza. choose the resolution , and according to chapter Błąd! Nie można odnaleźć źródła odsyłacza. choose the number of oxygen probe;

-	insert the oxygen and temperature probes to the measured solution;

-	turn the meter on with  button;

-	with  button choose the oxygen measuring function;

-	choose the unit according to chapter 21.1;

-	by accurate measurements in mg/l enter the salinity value (chapter 21.2);

-	check or simulate the flow of the measured solution;

- wait till stabilisation of the result about 1  1.5 min (depends on the probe) and read the result.

## Measurement with manual temperature compensation

To make the measurement with manual temperature compensation one should:

-	disconnect the temperature probe from the meter;

turn the meter on by pressing the  button;

-	with  button choose the oxygen measuring mode;

-	choose the unit according to the chapter 21.1;

-	by accurate measurements in mg/l enter the salinity value (chapter 21.2);

-	insert the oxygen sensor to the measured solution;

-	using thermometer measure the temperature of the solution;

-	with  ,  buttons enter in the lower row of LCD value of temperature of the measured solution (simultaneous pressing of   buttons sets the temperature value to 20 0C;

-	check or simulate the flow of the measured solution;

-	after the stabilisation read the result. (Pic. 31);

During series of measurements it is recommended to check and correct the introduced value of solution temperature.

Pic. 31.

Caution:	in case of measurements in solutions with low salinity check according to the chapter 21.2, weather the introduced salinity value is 0.00 g/l.

##### ATMOSPHERIC PRESSURE MEASUREMENT

# ATMOSPHERIC PRESSURE MEASUREMENT

The meter enables atmospheric pressure measurement. To read it’s value one should:

-	turn the meter on by pressing the  button;

-	with  button choose the oxygen concentration mode (O2);

-	press the  button till moment of displaying in the lower row of LCD the  (pressure) sign;

-	in the upper row the value of the pressure in hPa will be displayed 
(Pic. 32);

Pic. 32.

-	return to the measuring mode by pressing the  button.

##### ORP (mV) AND TEMPERATURE
 MEASUREMENT

# ORP (mV) MEASUREMENT

Multifunction meter CX-505 is an accurate mV meter (Oxidation Reduction Potential). The measurement may be done with special redox electrode or during titration. The result may be checked after choosing the mV mode with  button (Pic. 33).

Pic. 33.

# TEMPERATURE MEASUREMENT

The temperature measurement is done as follows:

-	by pressing the  button switch the meter on;

-	to the Chinch connector connect the temperature probe, on the display a  symbol will be displayed;

-	put the temperature probe to the measured solution;

-	wait till the stabilisation of the value and read the result in the lower row of numbers.

The meter co-operates with the Pt-1000 probe. Depending on its class the accuracy of the measurement changes.

CAUTION: break in the circuit of the temperature probe switches the meter to the manual temperature compensation mode. It is signalised by changing the  symbol to  symbol. On the display the value of the temperature entered by the user is displayed.

Blinking -50°C value during measurement in positive temperatures informs about short circuit in the temperature probe.

##### OTHER

# CLOCK WITH DATE

After choosing the time mode with  button the meter will display the actual time. By pressing the  button one may display the date, backlight mode, auto switch off time and number of the software version.

## Clock

The time is displayed in two rows on the display (Pic. 34). In the upper one the hour and minutes are displayed and in the lower one seconds. The way of changing the hours is described in p. Błąd! Nie można odnaleźć źródła odsyłacza..

Pic. 34.

## Date

The date is displayed as follows: Month – Day – Year. (Pic. 35). In the upper row the month and day are displayed and in the lower the year.

Pic. 35.

## Setting the time and date

The setting mode of the actually displayed parameter (hour or date) one enters by pressing and holding the . button. The position which we are going to change starts blinking, the value is changed with  or . buttons. The pulsating position is chosen by short pressing of the   button. There is no way to set the seconds, they are deleted after the setting mode is left. Press the  button to leave the setting mode.

## The LCD backlight mode

In the time mode press the  button till displaying in the upper row of LCD a  sign (Pic. 36). In the lower row an ,  or  parameter will be displayed.

- the backlight is switched off;

- mode of automatic switching on of the backlight for 60 seconds    after each pressing of any button;

- the backlight is always on.

Pic. 36.

With  or  buttons one may choose the backlight mode.

Return to the time mode after pressing the  button.

## Brightness control

Controlling the brightness of the backlight is especially important during the work In different conditions of lighting. During the process of choosing the backlight mode (p. 27.4) it is possible to control the brightness by short pressing the  button. In the place of  or  parameter the value of brightness in % will be displayed. With  or  buttons it is possible to set the brightens value from 10 ÷ 100%, with 10% step.

Pic. 37.

Return to the backlight mode be pressing the  button, and to the measuring mode by pressing any function button.

## Readout of the software version number

In the  time function press the  button till displaying a screen, as on the picture below (Pic. 38). In the upper row the software version number is displayed

Pic. 38.

Return to the time mode after pressing the  button.

# STORING AND READOUT OF THE RESULTS

The meter enables storing of 950 results of the actually measured function. The results are stored in EEPROM memory, which is non-volatile, so the data isn’t lost even after complete lack of power. Before starting the work it is necessary to choose the parameters of storing or readout of the stored results.

## Parameters of storing and readout from the memory

The parameters are changed in the readout mode. It may be entered from every measuring function by pressing and holding the  button, until on the display the number of last stored result will be displayed. This number is displayed on turns with the stored result.

Before starting the storing one has to choose the kind of results collecting: on request or automatically as a series, and also the way of displaying the result.

Next pressing of the  button shows screens with following functions which may be changed:

a.	  - taking or printing series or single results.

Pic. 36.

With  or  button choose in lower row   or  symbols  
(Pic. 36).  Choosing  activates automatic storing of the results and  single, after every pressing of  button.

b.	  - time interval during series taking (Pic. 40).

Pic. 40.

Value of the time interval is displayed in the lower row of digits, and the informative symbol   in the upper row.

Buttons  and  are used to set the chosen time in minutes and seconds. The shortest time is 1 second and the longest 60 minutes. Holding the button makes changing quicker. (repetition).

In case of choosing the  parameter to  the position  is not displayed.

c		 - Way of displaying the stored results

- successively number of sample, result, time and date of storing the result.

- successively number of sample and result

Changing with  or  buttons.

Return to the results readout display after pressing  button.

Exit from the readout mode after pressing of the button of chosen measuring function.

## Storing the single measurements to the memory

If according to previous section storing of single results was chosen, every pressing of  button stores the measured result. The results are stored as the next ones after the last stored. If someone was checking the earlier stored results and didn’t return to the last one the results won’t be deleted and the value will be stored from the first empty position. If the user wants to store the result from the chosen number first he has to delete the results (as described in the point 28.5) and next start storing the results by pressing the button.

During storing of the results the number of result will be displayed for a moment.

If after pressing the  button instead of number an  sign will be displayed it informs that the maximal number of results was stored.

## Storing the measuring series

There is possibility to store series of measurements in the meter’s memory. To do so one has to :

-	choose the serial collecting the measurements (p 28.1.a);

- enter the time interval (p 28.1.b);

delete the stored results starting from the chosen one (p 28.5);

-	 by the proper button choose the function which results one wants to store;

-	with  button start taking the series. The measurements will be stored starting from the first free number.

The red LED in the  button will begin flashing informing that the meter is during storing of series of measurements. Before each  storing of the result the number of measurement will be displayed for a moment. The taking of the series may be stopped by pressing   button, any function key or by filling the memory. If the series results are not filling the memory totally one may start to take the next series till the end of memory.

## Reviewing the results

Reviewing of the stored results is started from the measuring mode, by pressing and holding the  button until the number of last stored result on turns with it’s value will be showed.

Every pressing of the  or  button shows the next or previous number and result with time and date if  function was chosen (p 28.1.c).

In this mode the buttons  ,  work with repetition and after longer holding the numbers change very quickly till stopping on the highest or lowest number.

The reviewing mode may be left by pressing any function button.

## Deleting the stored results

To delete the stored results one has to:

-	press and hold the  button;

-	with buttons  ,  set the number of measurement from which we want to start deleting the memory;

-	press and hold the  button what will delete the stored results from the chosen one till the end of the memory. On the display in the place of results there will be  sign displayed what confirms the deleting;

-	to exit the reviewing mode press any function button.

If one wants to clear all the memory the deleting should be started from the first number.

# CO-OPERATION WITH A PC

Connecting the meter with a PC enables storing data directly on the computer, what eliminates the limitation of the number of stored results. The PC should be equipped with USB port. For the transmission a special software prepared by our company may be ordered. The software is delivered on a CD. After inserting the CD to the drive the installation program will start automatically. It is necessary to follow the given instructions.

In the upper wall of the meter a USB port is placed, it enables connecting with a PC using a cable. After connecting the meter and the PC should be switched on and the transmission software should be run. We have two options available:

-	“Collect series” is used for collecting results of current measurement. After choosing this option we will see a window with the result of current measurement. Collected and stored will be only elements which are marked in the field “Send”. It is necessary to set the number of measurements which are to be stored and time breaks between storing it. On the basis of this data the software will count the time of collecting the whole series. The series is stored in temporary file . In case of lack of power the so fare collected data will be stored in a file “NoNamexx”. The collecting is started by pressing the “Collect” button;

-	“Download data from memory” enables sending the chosen part or whole of the data stored in the meter’s memory to a file. In option “Collect” we mark the data we want to be sent. The transfer is started by pressing the button “Download”

Caution:	the meter and PC should be switched on after connecting the cable.

# TECHNICAL DATA

pH MEASUREMENT:

| Range | Resolution | Accuracy (±1 digit) |
| --- | --- | --- |

| -2.000 16.000 pH | 0.001 / 0.01 pH | ±0.002 pH |
| --- | --- | --- |

INPUT IMPEDANCE:					1012

TEMPERATURE COMPENSATION:	manual/automatic

RANGE OF COMPENSATION:			-5.0 ÷ 110.0 oC

pH ELECTRODE CALIBRATION: 	 	automatic,

in 1 ( 5 points

RANGE OF RECOGNITION AND ENTERING OF THE pH BUFFER SOLUTIONS

| Calibration point | Range |
| --- | --- |
| 1 | 0,800 ( 2,100 |
| 2 | 3,900 ( 4,100 |
| 3 | 6,800 ( 7,100 |
| 4 | 8,900 ( 10,200 |
| 5 | 11,500( 14,000 |

AUTOMATIC CHANGE OF THE pH BUFFER VALUE TOGETHER WITH THE TEMPERATURE CHANGE, FOR SAMPLES CONSISTENT WITH NIST (table page 19) IN RANGE   			0 ÷ 60 oC

THERMAL STABILITY OF ZERO:		0.001 pH/ oC

mV MEASUREMENT:

| Ranges | Resolution | Accuracy (±1 digit) |
| --- | --- | --- |

| -1000 1000 mV | 0.1 mV | ±0.1 mV |
| --- | --- | --- |

INPUT IMPEDANCE:				1012

CONDUCTIVITY MEASUREMENT:

| Ranges | Resolution | Accuracy (±1 digit) | Frequency |
| --- | --- | --- | --- |

| 0.000 ÷ 19.999 S/cm | 0.001 / 0.01 S/cm | ±0.1 % | 100 Hz |
| --- | --- | --- | --- |

| 20.00 199.99 S/cm | 0.01 / 0.1 S/cm | ±0.1 % | 1 kHz |
| --- | --- | --- | --- |

| 200.0 1999.9 S/cm | 0.1 / 1 S/cm | ±0.1 % | 2 kHz |
| --- | --- | --- | --- |

| 2.000 19.999 mS/cm | 0.001 / 0.01 mS/cm | ±0.1 % | 5 kHz |
| --- | --- | --- | --- |

| 20.00 199.99 mS/cm | 0.01 / 0.1 mS/cm | ±0.25 % | 10 kHz |
| --- | --- | --- | --- |
| 200.0 1999.9 mS/cm | 0.1 / 1 mS/cm | ±0.25 % | 10 kHz |

*	Accuracy given for the end value of the range.

Ranges of frequency changes were given for constant  K = 1. For other values of the constant K the values will change proportionally to changes of this constant.

TEMPERATURE COMPENSATION:	manual/automatic

COMPENSATION RANGE:			-5.0 ÷ 70.0 oC

CONSTANT K RANGE:				0.010  19.999 cm-1

α COEFFICIENT RANGE:				0.00  10.00 %/ oC

TDS COEFFICIENT RANGE:			0.20  1.00

MEASURING RANGE KCl:				0  200 g/l

MEASURING RANGE NaCl:			0  250 g/l

PROBE CALIBRATION:				one point

1. by entering the constant K of the probe

2. using the calibration solution

OXYGEN MEASUREMENT:

| Range | Resolution | Accuracy |
| --- | --- | --- |

| 0 400.0 % | 0.1 % | Of probe* 1 digit |
| --- | --- | --- |

| 0 60.00 mg/l | 0.01 mg/l | Of probe* 1 digit |
| --- | --- | --- |

* Accuracy given in the “Oxygen Probe” section.

Temperature compensation range:		0.0 ÷ 40.0 oC

Salinity compensation range:			0.0  50.0 g/l

Atmospheric pressure compensation range:	800  1100 hPa

Probe calibration:

Two point							0% and 100% O2

One point							in 100%O2

Oxygen probe:						membrane, galvanic

# ATMOSPHERIC PRESSURE MEASUREMENT:

| Range | Resolution | Accuracy* (±1 digit) |
| --- | --- | --- |

| 800 ÷ 1100 hPa | 1 hPa | 2 hPa |
| --- | --- | --- |

Atmospheric pressure sensor	HP03S

# TEMPERATURE MEASUREMENT:

| Range | Resolution | Accuracy* (±1 digit) |
| --- | --- | --- |

| - 50.0 ÷ 199.9 oC | 0.1 oC | 0.1 oC |
| --- | --- | --- |

*	accuracy of the meter. Final accuracy of the measurement depends on the accuracy of the used PT-1000 probe

TEMPERATURE PROBE:				platinum resistor Pt-1000

ACCURACY OF THE PROBE IN RANGE 0 ÷ 100 0C:

FOR PT1000B RESISTOR:			±0.8 0C

FOR PT10001/3B RESISTOR:			±0.27 0C

# OTHER:

# MEMORY OF THE RESULTS:				950 results

# WORKING TEMPERATURE:		0 ÷ 40 oC

POWER:		power adapter 12V/100mA.

POWER CONSUMPTION:		60 mW

SCREEN:		LCD 55 x 45 mm

DIMENSIONS:		200 x 180 x 50 mm

WEIGHT:		620g

# EQUIPMENT

The standard equipment for the CX-505 meter is:

1.	Temperature probe  Pt-1000B (standard);

2.	Power adapter 12V/100mA.

3.	Software for collecting the data on the PC

4.	USB cable

5.	User's manual with warranty.

The additional equipment available for this meter is:

1.	Combined pH electrode (glass membrane);

2.	Conductivity cell chosen according to measuring range;

3.	Adapter for connecting measuring end reference electrode in place of combined one;

4.	Temperature probe Pt-1000 1/3B with higher accuracy;

5.	ORP measuring electrodes;

6.	Dissolved oxygen sensor.

W A R R A N T Y

The ELMETRON company ensures a 24 months warranty for the multifunction meter CX-505 number:

..................................................

In case of damage the producer will repair the meter within 14 days from the day of delivery.

The warranty doesn’t cover the damages caused by usage not in conformity with the user’s manual, using wrong power adapter, mechanical damages and damages caused by repairs made by unauthorised persons.

The pH electrode, conductivity cell, oxygen sensor have the warranty of the producer.

NOTICE: Before sending the meter to us please contact the firm by phone or email.

When sending the meter, the used electrode, conductivity cell, oxygen sensor, temperature probe and power adapter should be also included.

Date of production.........................................

Date of sale...................................................

Date of warranty expiry..................................

(  Sp. j.

41-814  Zabrze   Witosa 10   Poland

tel.  +48 32 2738106   fax  +48 32 2738114

www.elmetron.com.pl

e-mail: info@elmetron.com.pl